SOURCE CONTROL SPECIFICATION

Part Number: E.SLQ002

Only the item described in this source control to meet the specification listed herein is approved by SILQ.  A substitute item shall not be used without prior approval.

# Purpose

## This document provides the specification for the 750 gallon high purity water tank assembly.

## The High Purity Water Tank Assembly is to accommodate intermediate, temporary storage of high purity water.

# High Purity Tank Specification

## Manufacturer Part Name:  Norwesco 750 Gallon Vertical Tank

## Manufacturer Part Number:  VERT-750

## Tank Capacity: 750 gallons

## Dimensions (Diameter x Height): 48” x 101”

## Material: Polyethylene

## Connection: 2” Poly Bulkhead Tank Fitting (PBM Supply Item #TF200)

# Assembly

## Top Components

### Thick-Wall Through-Wall Pipe Fitting, UV-Resistant Polypropylene for Chemicals, Connector, 1 NPT Female

### PTFE Thread Sealant Tape

### Super-Flow Barbed Fitting for Food and Beverage, 90 Degree Elbow Barbed Adapter for 1" Tube ID x 1 NPT Male

## Bottom Components

### Thick-Wall Polypropylene Pipe Fitting for Chemicals, Straight Connector, 2 NPT Male, 6" Long

### Easy-Install Double-Union Polypropylene On/Off Valve for Water, T-Handle, Black, 2” NPT Female

### UV-Resistant Polypropylene Pipe Fitting for Chemicals, Bushing Reducing Adapter, 2 NPT Male x 1 NPT Female

### UV-Resistant Polypropylene Pipe Fitting for Chemicals, Bushing Reducing Adapter, 1 NPT Male x 1/2” NPT Male

### Super-Flow Barbed Fitting for Food and Beverage, 3/8" Tube ID x 1/2" NPT Female

## The following modifications and additions have been made to the waste tank

### Through Wall Pipe Fitting

#### Drill a hole suitable in diameter for attaching the through-wall pipe fitting.

#### Approximate location at the top of 750 GAL – High Purity Water Tank is demarked by a red asterisk in Attachment 1 – 750 Gallon Vertical Tank drawing.

#### Insert Through-Wall pipe fitting and secure to tank.

### High Purity Water Inlet Fittings

#### At top of tank, secure 90 Degree Elbow Barbed Adapter for 1" Tube ID x 1 NPT Male to through-wall fitting. 1” ID hose from high purity water filter assembly may now be installed.

### High Purity Water Outlet Fittings

#### At bottom of tank (cone bottom), secure 2” NPT male 6” straight connector. Secure Double-Union Polypropylene T-Handle On/Off Valve. Secure 2” NPTM x 1” NPTF Bushing reducing adapter to on/off valve. Secure 1” NPTM x 1/2” NPTM reducing adapter to reducing adapter. Secure 3/8" Tube ID x 1/2" NPT Female barbed fitting to reducing adapter. 3/8” ID Hose for high purity water outlet may now be connected for providing high purity water to the Repligen KTF-2000 system (E.SLQ006).

# Attachments

## Attachment 1 – 750 Gallon Vertical Tank drawing

## Attachment 2 – Purified Water Tank Fitting Assemblies drawing