SOURCE CONTROL SPECIFICATION

Part Number: E.SLQ006

Only the item described in this source control to meet the specification listed herein is approved by SILQ.  A substitute item shall not be used without prior approval.

# Purpose

## This document provides the specification for Diafiltration System (KrosFlo KTF-2000 System).

# Supplier

## Repligen Corporation, www.repligen.com, 18617 South Broadwick Street, Rancho Dominguez, CA 90220, Tel: (310) 885-4600, email: customerserviceUS@repligen.com.

# Manufacturer

## Repligen Corporation, www.repligen.com, 18617 South Broadwick Street, Rancho Dominguez, CA 90220, Tel: (310) 885-4600, email: customerserviceUS@repligen.com.

# Reference Part Number(s)

## Manufacturer Part Name:  KrosFlo® KTF-2000 System

## Manufacturer Model Number:  SYTF-2000

# Specification

## Description: Designed for high volume, high concentration applications, the KrosFlo® KTF-2000 System from Repligen is designed for robust downstream ultrafiltration and microfiltration.

## Volume: 200 L – 5000 L

## Flow rate: Pumps 1 and 2: 85 liters per minute

## Filtration area: 2 each up to 17.2 m2 (hollow fiber)

## Cabinet width: 30”

## Tubing Connections to Tank: 1 in. ID Outlet to Levitronix Pump from Tank, 0.75” ID return to Tank

## Hollow fiber filter used: KrosFlow Max 68cm 10kD MPES 0.5mm 6TC x 1.5TC (Repligen part number X06-E010-05-N)

## Software: iFIX 6.0 (process requirements), Dream Reports (batch reports), iHistorian, E-signature

## GMP: Enables 21CFR11

## Power: 200-240 VAC 1ph 15A

# System architecture

## The fully integrated KrosFlo KTF-2000 System features two magnetically driven centrifugal pump heads, custom ProConnex® Flow Paths, flow meters, peristaltic pumps, hollow fiber filters, and connections for a weight scale for additional process monitoring. Other components of the system include the retentate flow meters, permeate flow meter, pressure sensors, Repligen control panel controlled by an Allen-Bradley PLC, flexible tube flow path, and a stainless steel frame/cart that serves as the system support structure.

## Major components

### Retentate (recirculating) Pump (P-01): Levitronix LPM-600.3-10 low-shear centrifugal pump

### Product pump (P-02): MasterFlex® 77420-10 I/P peristaltic pump

### Buffer Pump (P-03): MasterFlex® 77420-10 I/P peristaltic pump

### Retentate flowmeter (FL-01): Sonotec® FS04.210 C0.55/230 V2.0

### Permeate flowmeter (FL-02): Levitronix LFS-06SU / LFS-10SU

### Pressure sensors (PE-01/02/03): Repligen ACPM-799-01N

### PLC: Allen-Bradley 1766-L32BWA Micrologix 1400

### Human Machine Interface: Allen-Bradley 6181P15C2MWX1AC Integrated Panel Display Computer

### Silicone Tubing 1” ID to meet FDA standards and from approved supplier list (eg. McMaster-Carr 5236K244 or equivalent)

### Silicone Tubing 0.75” ID to meet FDA standards and from approved supplier list (eg. McMaster-Carr 5113K51 or equivalent)

# Attachments

## KrosFlo® Systems User Guide