SOURCE CONTROL SPECIFICATION

Part Number: M.SLQ001

Only the item described in this source control to meet the specification listed herein is approved by SILQ.  A substitute item shall not be used without prior approval.

# Purpose

## This document provides the specification for the unpurified polymeric dispersion solution.

# Supplier

## TCI America, 9211 N. Harborgate Street, Portland, OR 97203, Contact name: Corey Smith, Contact Phone number: 503-935-5128, Contact email: corey.smith.@tcichemicals.com, website: tcichemicals.com

# Manufacturer

## TCI America, 9211 N. Harborgate Street, Portland, OR 97203, Contact name: Corey Smith, Contact Phone number: 503-935-5128, Contact email: corey.smith.@tcichemicals.com, website: tcichemicals.com

# Reference Part Number(s)

## Manufacturer part number: Z5503

# General Information

## Handling:  Wear goggles and gloves when handling the material. Keep in a closed container until ready to use.

## Storage: Keep the chemical solution is dry location away from direct sunlight until ready to use.

## Consistent with dispersions, the solids may settle to the bottom of the container during storage.

# Specification

## Description: The unpurified polymeric dispersion solution contains a three-component methacrylate copolymer, polymerized via free radical polymerization from polysulfobetaine methacrylate, perfluorophenylazide methacrylate, and methacrylic acid dissolved in 2,2,2-Trifluoroethanol. The copolymer structure and mol ratios of each component of the polymer is shown below:

## The material is dispersed in 2,2,2-Trifluoroethanol.

## Materials Specification:

### Appearance (color)*: Yellow

### Appearance (form)*: Viscous liquid

### Residue on Evaporation: 10-20%

### Specific gravity: 1.2-1.5

### pH: 5-8

## Packaging: The material is packaged in HDPE, stainless-steel, or glass bottles or drums.

# Certificates

## A Certificate of Analysis (CoA) from the manufacturer is required for each batch/lot number of product.

# Receiving Inspection

## Ensure there is no visible signs of damage to the product primary container.

## Review the manufacturer CoA and packing slip and verify order documentation for quantity and part number match.

## Review the manufacturer CoA and verify the product properties listed above with an asterisk (*) match material specifications or are within required ranges.