Installation Qualification,

Diafiltration System, ST-005

Test Report

# Executive Summary

| ## Status: | ## PASS |
| --- | --- |

## Testing was conducted to ensure the Repligen KTF-2000 diafiltration system was installed appropriately in accordance with manufacturer’s specification. This report documents the results of the equipment installation qualification for the diafiltration system.

## Installation qualification performed met the respective acceptance criteria for the operation of the diafiltration system.

# Deviations

## No deviations occurred during the execution of the test protocol (Document #VV.011).

# Results

## Installation qualification on the Repligen KTF-2000 diafiltration system (Equipment ID #ST-005, S/N 190603-001) was successfully completed in accordance with the protocol VV.SLQ011.

## The required documentation was found in the equipment file:

### User manual in English (See Attachment 2)

### Valid preventive maintenance and calibration report (See Attachment 3)

### Calibration and preventive maintenance schedule (See Attachment 4)

## The required utilities were verified to meet acceptance criteria:

### Mains power available is 240VAC 1ph on 20A breaker which meets the required power input specification.

### No damage to power supply cable/wiring was observed.

### The temperature of the work environment was measured to be 18.5℃ and within the acceptable operating temperature specified by manufacturer of 0 – 50℃.

## The KTF-2000 system was powered on and verified to be operational.

# Conclusions

## All required documentation and utilities for the KTF-2000 diafiltration system have been verified to meet manufacturer’s specification, and the system has been verified to be operational.

## The KTF-2000 diafiltration system (Equipment ID # ST-005) has been installed properly in accordance with manufacturer’s specification.

# Attachments

## Attachment 1: Installation Qualification datasheet

## Attachment 2: Users Manual

## Attachment 3: PM Calibration Report

## Attachment 4: Calibration and PM Requirements