Installation Qualification,

Peristaltic Pump, ST-008

Test Report

# Executive Summary

| ## Status: | ## PASS |
| --- | --- |

## Testing was conducted to ensure the U.S Solid M6-12L peristaltic pump was installed appropriately in accordance with manufacturer’s specification. This report documents the results of the equipment installation qualification for the peristaltic pump.

## Installation qualification performed met the respective acceptance criteria for the operation of the peristaltic pump.

# Deviations

## No deviations occurred during the execution of the test protocol (Document #VV.012).

# Results

## Installation qualification on the U.S Solid M6-12L peristaltic pump (Equipment ID #ST-008, S/N 11006080100) was successfully completed in accordance with the protocol VV.SLQ012.

## The required documentation was found in the equipment file:

### User manual in English (See Attachment 2)

### Preventive maintenance schedule (See Attachment 3)

## The required utilities were verified to meet acceptance criteria:

### Mains power available is 110VAC which meets the required power input specification.

### No damage to power supply cable was observed.

### The temperature and humidity of the work environment were measured to be 18.5℃ and 40% RH, respectively, and were within the acceptable operating temperature and humidity ranges specified by manufacturer of 0 – 40℃ and <80% RH, respectively.

## The peristaltic pump was powered on and verified to be operational.

# Conclusions

## All required documentation and utilities for the M6-12L peristaltic pump have been verified to meet manufacturer’s specification, and the pump has been verified to be operational.

## The M6-12L peristaltic pump (Equipment ID # ST-008) has been installed properly in accordance with manufacturer’s specification.

# Attachments

## Attachment 1: Installation Qualification datasheet

## Attachment 2: Users Manual

## Attachment 3: PM Requirements