Installation Qualification,

Waste Diaphragm Pump, ST-004

Test Report

# Executive Summary

| ## Status: | ## PASS |
| --- | --- |

## Testing was conducted to ensure the ARO 6661AJ-344-C waste diaphragm pump was installed appropriately in accordance with manufacturer’s specification. This report documents the results of the equipment installation qualification for the waste diaphragm pump.

## Installation qualification performed met the respective acceptance criteria for the operation of the waste diaphragm pump.

# Deviations

## No deviations occurred during the execution of the test protocol (Document #VV.014).

# Results

## Installation qualification on the ARO 6661AJ-344-C waste diaphragm pump (Equipment ID #ST-004, Serial Number SPL2281133) was successfully completed in accordance with the protocol VV.SLQ014.

## The required documentation was found in the equipment file:

### User manual in English (See Attachment 2)

### Preventive maintenance schedule (See Attachment 3)

## The required utilities were verified to meet acceptance criteria:

### Air is available to power the waste diaphragm pump which meets the required power input specification.

### No damage to power source cable was observed.

### The temperature of the work environment was measured to be 18.5℃ and was within the acceptable operating temperature range specified by manufacturer of 2 – 79℃.

## The waste diaphragm pump was powered on and verified to be operational.

# Conclusions

## All required documentation and utilities for the 6661AJ-344-C waste diaphragm pump have been verified to meet manufacturer’s specification, and the pump has been verified to be operational.

## The 6661AJ-344-C waste diaphragm pump (Equipment ID # ST-004) has been installed properly in accordance with manufacturer’s specification.

# Attachments

## Attachment 1: Installation Qualification datasheet

## Attachment 2: Users Manual

## Attachment 3: PM Requirements