Silq 2026 Quality Plan

Prepared by:

Ethan Rao

Director of R&D, QA, RA

Silq Technologies Corporation

Signature: 			Date: 19 Apr 2026

# Purpose

## This quality plan defines the quality practices, resources, and activities to be executed by Silq in 2026. It consolidates action items from the 2025 Management Review Meeting, reconciliation actions from a QMS Gap Analysis (FDA QMSR and ISO 13485:2016), and ongoing regulatory compliance activities.

## Per QM.SLQ025 A, this plan addresses changes to the quality system, identification of training requirements, review and updates to risk management processes, preparation for internal and external assessments, and identification of feedback mechanisms for continuous quality improvement.

# Background

## The FDA Quality Management System Regulation (QMSR) took effect February 2, 2026, replacing the legacy 21 CFR Part 820 framework with ISO 13485:2016 as the foundational QMS requirement. QMS Gap Analysis identified 2 significant gaps and several document deficiencies. The 2025 Management Review concluded the quality system is overall effective, suitable, and adequate with opportunities for improvement identified.

## The FDA conducted a Level-2 Surveillance Inspection on October 14–16, 2025 (FEI: 3014650073), resulting in two Form 483 observations. CAPA 2025-001 (Observation 1, MDR reporting) and CAPA 2025-003 (Observation 2, design control for unauthorized valve modification) are in process. Design control project DC.SLQ001 was established as Corrective Action #2 of CAPA 2025-003 to retrospectively apply design controls per QM.SLQ004 B Section 16. A second design control project, DC.SLQ002, manages the validation and transition of the SilqQMS EDMS platform from FileHold per QM.SLQ032.

## This plan outlines action items concerning QMS procedure revisions, CAPA closure activities, training initiatives, and quality objective enhancements to maintain compliance through the transition period.

# ISO 13485:2016 Section 4 — Quality Management System

## QMSR Supplementary Provision Mapping (Owner: Ethan Rao, Q2 Completion)

### Problem: QM.SLQ027 (Quality Manual) does not explicitly integrate the QMSR supplementary provisions (§820.10(b), §820.35, §820.45) as "applicable regulatory requirements."

### Action Item: Add a QMSR supplementary provision mapping section to QM.SLQ027 that maps §820.10(b) requirements (UDI, MDR, corrections/removals, combination products) to their implementing procedures, and include a formal combination-product applicability determination statement.

## Medical Device File Framework (Owner: Ethan Rao, Q2 2026)

### Problem: ISO 13485 clause 4.2.3 requires a "medical device file" as a single defined construct that contains or references all documentation needed to demonstrate device conformity. This concept was not in the old Part 820. Silq maintains the underlying components -- DMR (QM.SLQ048), DHF (design control SOPs), DHR (QM.SLQ029) -- but there is no explicit framework tying them together under the ISO construct. An auditor asking for the "medical device file" would not find a single defined reference.

### Action Item: Revise QM.SLQ027 or QM.SLQ048 to include a "Medical Device File" section. Define the construct and provide a cross-reference mapping existing DMR, DHF, DHR, risk management file, labeling documentation, and product specifications to clause 4.2.3.

## QMS Platform Transition, DC.SLQ002 (Owner: Ethan Rao, Per DC.SLQ002 Project Schedule)

### Problem: Silq's current EDMS (FileHold) has functional limitations that prompted the transition to SilqQMS, a custom-developed web application. The SilqQMS platform must be validated per QM.SLQ032 and ISO 13485 clause 4.1.6 before production use, and all FileHold-specific procedure references must be updated.

### Action Item: Design control project DC.SLQ002 manages validation of SilqQMS and the controlled platform transition. The project plan (DC.SLQ002 A) defines a phased approach: software validation (SW.SLQ007-012 deliverable series), then procedure-by-procedure migration from FileHold to SilqQMS. The 2026 QMS Transition Plan will be distributed to the management team.

## Systemic Legacy 820 References (Owner: Ethan Rao, Q3 2026)

### Problem: 19 QM documents contain explicit references to old Part 820 section numbers (e.g., section 820.40, section 820.30, section 820.198).

### Action Item: Execute a controlled batch update through the DCO process to replace old Part 820 citations with the corresponding ISO 13485 clause and QMSR supplementary provision references. Controlled batch update will coincide with additional document revisions required in DC.SLQ002 described above

# ISO 13485:2016 Section 5 — Management Responsibility

## Quality Objectives: Revise Existing Objectives (Owner: Ethan Rao, Q2 2026)

### Problem: The two existing quality objectives (Incoming Quality: lot acceptance rate greater than 80%; Finished Product Quality: complaint rate less than 5%) were determined to be insufficiently detailed with thresholds that are too lenient.

### Action Items:

#### Assemble historical data for both existing objectives and propose revised, more rigorous thresholds for review and agreement by the management team. Upon alignment, prepare and execute a DCO for QM.SLQ037.

#### Propose specific, quantifiable goals for three new objectives: (1) Active Post-Market Surveillance, (2) Dynamic Employee Training Program, and (3) Quarterly Execution of Quality Plan Objectives. Present to management team for review and alignment, then execute DCO for QM.SLQ037.

# ISO 13485:2016 Section 6 — Resource Management

## New Hire Training, Haley Shomo (Owner: Ethan Rao, 01 May 2026)

### One new hire (Haley Shomo) joined in 2026 and requires completion of QMS training per QM.SLQ003.

## QMSR Transition Training (Owner: Ethan Rao, June 2026)

### The QMSR transition changes the regulatory framework under which Silq operates. All employees need to understand the shift from legacy 21 CFR Part 820 to ISO 13485:2016 incorporation by reference and the QMSR supplementary provisions.

## Design Control Retraining, CAPA 2025-003 (Owner: Ethan Rao, 01 Jul 2026)

### Problem: CAPA 2025-003 (FDA 483 Observation 2) requires retraining of all applicable Silq employees on design control policies as a corrective action to prevent recurrence of unauthorized design changes proceeding without required regulatory processes.

### Action Item: Retrain all applicable employees on QM.SLQ004 through QM.SLQ010 (design control procedure series). Successful completion constitutes the effectiveness confirmation for CAPA 2025-003.

## Regulatory Reference Materials (Owner: Verne Sharma, Q2 2026)

### Problem: Silq does not currently have the AAMI/ISO 13485:2016 Practical Guide or certain regulatory standard documents needed to support the QMSR transition and ongoing compliance activities.

### Action Item: Approve purchase of AAMI/ISO 13485:2016 Practical Guide (PDF) and other regulatory standard documents

## Dynamic Employee Training Program (Owner: Ethan Rao, Q3 2026):

### Problem: The current training system documents read-and-understand acknowledgments but does not assess comprehension of quality procedures or effectiveness of training delivery.

### Action Item: Establish a dynamic employee training program that assesses comprehension of existing quality procedures and procedure revisions, rather than relying solely on read-and-sign acknowledgment.

# ISO 13485:2016 Section 7 — Product Realization

## Complaint/Advisory/MDR Escalation Pathway (Owner: Ethan Rao, Q2 2026)

### Problem: Ownership boundaries between complaint handling (QM.SLQ021), advisory notice/recall (QM.SLQ030), and MDR reporting (QM.SLQ022) are not consolidated in a single reference. An auditor cannot quickly trace the escalation pathway from complaint intake to regulatory reporting decision. Additionally, Risk Management reviews have historically not adequately accounted for product complaints received by Silq.

### Action Item: Add a cross-reference section to QM.SLQ021 or QM.SLQ030 that maps event types (complaint, advisory notice, MDR, correction/removal) to the responsible procedure and decision owner.

## Design Control: Valve Modification Assessment, DC.SLQ001 (Owner: Ethan Rao, CAPA 2025-003 Corrective Actions by 01 Jul 2026)

### Problem: A design change was implemented by Ningbo Yingmed prior to November 6, 2024. Silq manufactured and distributed devices incorporating the modified valve without completing required design control activities. This was cited as FDA 483 Observation 2, resulting in CAPA 2025-003.

### Action Item: Design control project DC.SLQ001 was established as Corrective Action #2 to retrospectively apply design controls per QM.SLQ004 B Section 16 (Design Changes). The project includes: documentation of affected design inputs/outputs, review of existing risk management files (RM-0018 through RM-0021) to confirm coverage, compilation of design verification evidence from CAPA 001-2025 containment inspection data, finalization of the Letter-to-File regulatory determination, formal design review, and project closure. The Risk Management Review Confirmation Memo has been drafted confirming no revision to existing RM files is required.

## Design Control Procedure Revisions, CAPA 2025-003 (Owner: Ethan Rao, 01 Jul 2026)

### Problem: The root cause of the 483 Observation 2 finding was that QM.SLQ004, QM.SLQ005, and QM.SLQ006 did not mandate that supplier modification notifications trigger a mandatory design review, risk assessment, and regulatory evaluation.

### Action Item: Revise QM.SLQ004, QM.SLQ005, and QM.SLQ006 to add an explicit requirement that any notification of a supplier-initiated component or process modification triggers mandatory design review, risk assessment, and regulatory evaluation prior to continued manufacturing.

## Supplier Audit: Pathway MedTech (Owner: Ethan Rao, Q4 2026)

### Problem: Pathway MedTech is Silq's Category I contract manufacturer. The management team agreed that an on-site supplier audit is appropriate given the 483 findings and the contract manufacturing relationship.

### Action Item: Update the 2026 Supplier Audit schedule to include an on-site audit of Pathway MedTech.

# ISO 13485:2016 Section 8 -- Measurement, Analysis, and Improvement

## Feedback Governance Procedure (Owner: Ethan Rao, Q2 2026)

### Problem: ISO 13485 clause 8.2.1 requires a documented feedback procedure that provides early warning of quality problems and feeds into risk management. Existing processes, complaint handling (QM.SLQ021), post-market surveillance (QM.SLQ033), management review (QM.SLQ018), and CAPA (QM.SLQ016), lack a single governance statement tying them together as a unified feedback system with defined escalation triggers.

### Action Item: Add a governance section to QM.SLQ033 (Post-Market Surveillance SOP) that explicitly defines how complaint data (QM.SLQ021), CAPA trend data (QM.SLQ016), and management review outputs (QM.SLQ018) feed into the feedback loop, and define what constitutes an escalation trigger.

## Active Post-Market Surveillance Program (Owner: Chuck Greiner, Q2 2026; Ethan Rao, Q3 2026)

### Problem: The significant decrease in product complaint quantity correlates with the reduction in post-market clinical study activities, suggesting a gap in complaint capture rather than a genuine improvement. The management team agreed that relying on passive complaint receipt is insufficient and that an Active Post-Market Surveillance procedure must be established.

### Action Item: (1) Establish Active Post-Market Surveillance objectives and procedures. (2) Present to management team for review and alignment. (3) After alignment, open a CAPA to detail revisions to relevant quality management procedures.

## Complaint Non-Investigation Rationale (Owner: Ethan Rao, Q2 2026)

### Problem: ISO 13485 clause 8.2.2 and QMSR section 820.35 require that when a complaint is not investigated, the documented rationale must be retained. QM.SLQ021 does not currently contain an explicit requirement for this.

### Action Item: Add a requirement statement to QM.SLQ021 mandating documented rationale for every non-investigation decision. Single-sentence addition to existing procedure.

## Regulatory Reporting QMSR Framing (Owner: Ethan Rao, Q2 2026)

### Problem: Reporting channels exist for MDR (QM.SLQ022) and corrections/removals (QM.SLQ030), but neither procedure includes an explicit crosswalk to the QMSR section 820.10(b) "applicable regulatory requirements" framework.

### Action Item: Add a brief "applicable regulatory requirements" statement to QM.SLQ022 or QM.SLQ030 referencing the QMSR section 820.10(b) overlay and confirming Silq's reporting procedures satisfy this provision.

## CAPA 2025-001 Effectiveness Confirmation (Owner: Ethan Rao, 17 Nov 2026)

### CAPA 2025-001 addressed valve malfunction complaints (PCF 2025001). Preventative action (100% visual/dimensional inspection of SLQ-211610SPT finished goods) was performed 17 Nov 2025. Effectiveness must be confirmed by monitoring product complaints for inspected devices.

### Action Item: Monitor all product complaints regarding inspected devices. CAPA will be closed if no reoccurrences of valve malfunction are observed by the confirmation date.

## CAPA 2025-002 Effectiveness Confirmation (Owner: Ethan Rao, 01 Dec 2026)

### CAPA 2025-002 addressed FDA 483 Observation 1 (failure to submit MDR reports within 30 days for PCF 2023001 and PCF 2023004). Corrective action included detailed complaint review, retroactive MDR submissions (15 Dec 2025), and revision of QM.SLQ022 Rev B to align with the FDA definition of "Serious Injury."

### Action Item: Evaluate all product complaints 1 year from QM.SLQ022 Rev B effective date (01 Dec 2025). CAPA will be closed if all events meeting reportability criteria have been reported appropriately.

## CAPA 2025-003 Corrective Action Completion (Owner: Ethan Rao, 01 Jul 2026)

### Problem: CAPA 2025-003 addresses FDA 483 Observation 2 (manufacturing/distributing devices with unauthorized design change). Corrective actions include: design review of the check valve modification (in process under DC.SLQ001), Letter-to-File finalization, procedure revisions to QM.SLQ004/005/006 (supplier modification triggers), and employee retraining.

### Action Item: Complete the full corrective action process: finalize DC.SLQ001 deliverables, execute the Letter-to-File, revise QM.SLQ004/005/006, and retrain all applicable employees on design control policies. Effectiveness confirmation = successful completion of retraining.

# Risk Management (ISO 14971)

## Post-Production Risk File Update Triggers (Owner: Ethan Rao, Q3 2026)

### Problem: QM.SLQ012 Section 14 requires evaluation of post-production information for impact on risk management but does not define explicit trigger criteria for when such information requires a formal risk file revision versus monitoring until the next scheduled review.

### Action Item: Add a short decision framework to QM.SLQ012 Section 14 (or to QM.SLQ033) that defines specific trigger criteria for risk file revision

## Failure Mode and Probability Rating Review (Owner: Ethan Rao, Q2 2026)

### Problem: The 2024 and 2025 Management Reviews both identified a need to evaluate whether existing failure modes accurately reflect all observed product complaint root causes, and whether existing probability ratings should increase based on field complaint experience. Existing hazard #10.1 covers three failure modes with probability currently rated at "1" (Improbable) despite five valve-related complaints over 3 years.

### Action Item: Review RM-0018, RM-0019, RM-0020, RM-0021, RM-0094, RM-0141, and RM-0155 against historical product complaint files. Evaluate if failure modes are comprehensive and if probability ratings are appropriately calibrated to field experience.

## ASTM F623-25 Risk Management Alignment (Owner: Ethan Rao, Q3 2026)

### Problem: ASTM F623-25 (replacing ASTM F623-19) now states it does not comprehensively cover biocompatibility of Foley catheters with surface treatments or those used in suprapubic insertions. This scope change may require updates to risk management policies to maintain alignment with ISO 10993-1:2018 and ISO 14971.

### Action Item: Evaluate and update risk management policies to address the ASTM F623-25 scope change and align with ISO 10993-1:2018 and ISO 14971.

## Packaging Inspection for ASTM F1886 Compliance (Owner: Ethan Rao, Q2 2026)

### Problem: ASTM F1886/F1886M-25 (replacing ASTM F1886/F1886M-16) added a section for visual inspection of packaging material surfaces for defects that could compromise integrity.

### Action Item: Perform inspection of existing finished goods to verify compliance with the updated standard.

## Gage R&R Study, UV Light Absorbance and Rinse Test (Owner: Ethan Rao, Upon Next C.SLQ001 Production)

### Problem: Postponed action item from the 2024 Management Review. Gage R&R studies on in-process quality control test methods UV Light Absorbance (TM-HDX-001) and Rinse Test (TM-HDX-002) have not been completed.

### Action Item: Complete Gage R&R or suitable alternate study upon next production of C.SLQ001.

## UV Spectroscopy Test Protocol Development (Owner: Na He, Q3 2026)

### Problem: There is no established test protocol for analysis of UV spectroscopy of incoming and finished catheters. This was recommended as an opportunity for improvement at the 2025 Management Review.

### Action Item: Develop a test protocol for UV spectroscopy analysis of incoming and finished catheters.

# Action Item Tracking

## This plan will be tracked quarterly per the management review recommendation for "Quarterly Execution of Quality Plan Objectives." Status updates will be provided to the management team at quarterly intervals. The plan owner (Director, R&D, QA, RA) is responsible for maintaining status and escalating delays to the management team.