Design Project Plan

Design Assessment, Valve Component Modification, ClearTract Catheter

# Purpose

## The purpose of this design control project plan is to define the activities, reviews, deliverables, and resources necessary to complete a retrospective design assessment of a supplier-initiated valve geometry modification to the check valve assembly of the SILQ ClearTract Catheter product family. This project is initiated as Corrective Action #2 of CAPA 003-2025 in response to FDA 483 Observation 2 (16 Oct 2025), which cited Silq for manufacturing and distributing devices incorporating a modified component without completing required design control activities. The design assessment is conducted per QM.SLQ004 B Section 16 (Design Changes) and planned per QM.SLQ005 B (Design Project Planning SOP). This plan ensures the retrospective design assessment is conducted in compliance with 21 CFR 820.30 (Design Controls) and ISO 13485:2016 clause 7.3 (Design and Development).

# Scope

## This document is the design project plan for a limited, retrospective design assessment of a supplier-initiated component modification to the SILQ ClearTract Catheter check valve assembly. The project scope is defined in the approved Design Project Scope Form (FM1-QM.SLQ004 DC.SLQ001) and consists of:

### Project planning and scope definition (this document and the scope form)

### Documentation of the affected design input (valve dimensional specification) and design output (revised drawings/specifications)

### Review of existing risk management files (RM-0018, RM-0019, RM-0020) to confirm coverage of the valve geometry change

### Compilation of design verification evidence from containment inspection data (CAPA 001-2025) and receiving inspection results

### Finalization of the Letter-to-File regulatory determination (no new 510(k) required)

### Formal design review (phase review per QM.SLQ008) and project closure

## This is a retrospective design assessment. The valve geometry modification was implemented by the supplier (Ningbo Yingmed) prior to November 6, 2024. Devices incorporating the modified valve have been manufactured and distributed. Containment activities were completed under CAPA 001-2025. This project formalizes the design control documentation after the fact.

# Documents

## Reference Documents

### QM.SLQ001	Document Control SOP

### QM.SLQ004	Design Control Program SOP

### QM.SLQ005	Design Project Planning SOP

### QM.SLQ006	Design Input SOP

### QM.SLQ007	Design Output SOP

### QM.SLQ008	Design Review SOP

### QM.SLQ009	Design Verification and Validation SOP

### QM.SLQ012	Risk Management SOP

### QM.SLQ013	Risk Analysis SOP

### 21 CFR Part 820	Design Controls

### ISO 13485:2016, Medical Devices — Quality Management Systems — clause 7.3 (Design and Development)

### ISO 14971:2019, Medical Devices — Applications of Risk Management to Medical Devices

# Appendices / Associated Forms

### FM1-QM.SLQ004 A Design Project Scope Form DC.SLQ001

### Letter-to-File, Regulatory Assessment: Valve Component Modification

### CAPA 003-2025 (Source CAPA, FDA 483 Observation 2, investigation, root cause, and corrective action plan)

### CAPA 001-2025 | Containment activities — NCR, 100% inspection data, valve inspection criteria

# Definitions

## Abbreviations / Acronyms

### DCO: Document Change Order

### CAPA: Corrective Action / Preventative Action

### DHF: Design History File

### DFMEA: Design Failure Mode and Effects Analysis

### DMR: Device Master Record

### QMS: Quality Management System

### LTF: Letter to File

### NCR: Nonconformance Report

### PHA: Preliminary Hazard Analysis

### RA: Regulatory Affairs

### QA: Quality Assurance

### SOP: Standard Operating Procedure

### V&V: Verification and Validation

## General Definitions

### Check Valve Assembly: A non-patient-contacting component of the ClearTract Catheter (Item No. 5 per device drawing) that prevents backflow of inflation fluid and supports balloon inflation/retention during indwelling use. Located at the proximal balloon inflation port.

### Valve Geometry Modification: A dimensional/geometry refinement to the check valve assembly implemented by the component supplier (Ningbo Yingmed) to improve seating consistency and reduce dislodgement probability. No material, functional, or performance changes were made.

### Retrospective Design Assessment: design control activity performed after a design change has already been implemented, to formally document and evaluate the change under the organization's design control program. Conducted per QM.SLQ004 B Section 16 (Design Changes).

### Containment: Actions taken to control nonconforming product and prevent further distribution, including quarantine, 100% inspection, segregation, and disposition of affected units.

# Product Overview

## The SILQ ClearTract Catheter is a sterile, single-use, 100% silicone urological retention catheter intended for drainage (and irrigation, where applicable) of the urinary tract. The device is available in multiple French sizes.

## The catheter consists of a silicone shaft with drainage lumen(s), an inflatable retention balloon near the distal tip, a balloon inflation lumen, a proximal balloon inflation port incorporating a check valve assembly, and a drainage funnel connector. The inflatable balloon is filled via the proximal inflation port using a saline syringe. The check valve assembly prevents backflow of inflation fluid and supports balloon inflation/retention during indwelling use.

## The device operates via passive fluid drainage through the main lumen(s) and mechanical retention through balloon inflation. It contains no active energy source, electronic components, or software.

## The check valve assembly (Item No. 5 per device drawing) underwent a dimensional/geometry refinement by the component supplier (Ningbo Yingmed). The modification improves valve seating consistency and reduces dislodgement probability. No material or functional changes were made. The existing ≤26 mm dimensional acceptance criterion was specified prior to the modification and remains unchanged.

# Product Quality

## This project will be implemented to provide deliverables in compliance with the design control requirements set forth in 21 CFR 820.30 (Design Controls) and ISO 13485:2016 clause 7.3 (Design and Development).

## The design assessment is conducted per QM.SLQ004 B (Design Control Program SOP), specifically Section 16 (Design Changes), which requires a Design Project Scope Form and appropriate design control activities when a change to a component or process is made by a contract manufacturing organization. The project plan follows the requirements of QM.SLQ005 B (Design Project Planning SOP).

## Risk management activities are conducted per QM.SLQ012 (Risk Management SOP) and QM.SLQ013 (Risk Analysis SOP), in accordance with ISO 14971:2019.

# Key Milestones

## This project consists of five phases for the retrospective design assessment of the valve geometry modification:

### Complete and approve the Design Project Scope Form (FM1-QM.SLQ004 DC.SLQ001) and this Design Project Plan (DC.SLQ001 A). Review the existing Risk Management Plan (RM-0018 Rev D) for applicability to the valve geometry change. These documents establish the scope, activities, and resources for the design assessment.

### Document the affected design input: the valve dimensional specification (≤26 mm from sleeve to valve tip acceptance criterion). Document the design output: the revised device drawings and specifications reflecting the dimensional refinement. These drawing revisions were completed under CAPA 001-2025 Corrective Action #2; this phase formally documents them as design outputs of the change assessment. Confirm traceability between the affected design input and corresponding design output.

### Review existing risk management files — RM-0019 Rev D (Hazard Analysis Risk Management Report) and RM-0020 Rev D (DFMEA Risk Management Report) — to confirm they adequately cover the valve geometry change. Per CAPA 003-2025 Section V, the risk management report was reviewed and determined not to require update (no new failure modes, no increase in incidence, no new hazards). Document the risk assessment conclusion in a confirmation memo or addendum.

### Compile verification evidence demonstrating that design outputs meet design inputs for the affected component. Evidence sources include: containment inspection data from CAPA 001-2025 (100% visual/dimensional inspection of affected lots) and receiving inspection results for lots manufactured with the modified valve. Document a verification summary confirming conformance to the ≤26 mm acceptance criterion and related dimensional specifications.

### Finalize the Letter-to-File regulatory assessment confirming no new 510(k) is required per FDA guidance ("Deciding When to Submit a 510(k) for a Change to an Existing Device," Oct 2017). Conduct a formal design review (phase review per QM.SLQ008) with representatives from QA, R&D/Engineering, and Manufacturing to review all deliverables and confirm completeness. Record meeting minutes per FM1-QM.SLQ008. Submit all project deliverables under DCO and close CAPA 003-2025 Corrective Action #2 with reference to the completed design assessment.

# Key Resources

## Ethan Rao (QA/RA/R&D) Responsible for all documentation, risk assessment review, verification evidence compilation, and regulatory determination (Letter-to-File).

## Brian McVerry (R&D/Engineering) Design review participant. Provides engineering review of design input/output documentation and risk assessment conclusions. May consult with Pathway MedTech (contract manufacturer) for manufacturing-related data or confirmation as needed.

## Chris Turner (Manufacturing) Design review participant.

## Pathway MedTech (External Contract Manufacturer) Consulted as needed for manufacturing process information, component supply chain data, and receiving inspection records. Not performing design control activities.

## Ningbo Yingmed (External Component Supplier) Source of the valve geometry modification. Not a direct participant in design control activities; information provided through Pathway MedTech.

# Project Activities, Reviews and Deliverables

## The DC.SLQ001 Design Assessment, Valve Component Modification, ClearTract Catheter is a combination of the following activities: Project Planning, Design Input/Output Documentation, Risk Management Review, Design Verification, and Regulatory Determination and Design Review.

## The following table details the activities, reviews, and deliverables required for each phase of the design control project. Activities and deliverables are not limited to only those listed in the table. The status of each item listed in the table will be updated as the project unfolds.

Table 1: Design Control Project Activities, Reviews, and Deliverables for Design Project Plan, Design Assessment, Valve Component Modification, ClearTract Catheter

| Project Phase | Activity or Deliverable | Required Review | Estimated Date | Responsible Party | Status | Comments / Additional Information |
| --- | --- | --- | --- | --- | --- | --- |
| Phase 1: Project Planning | Approve the following documents through DCO: FM1-QM.SLQ004 A Design Project Scope Form DC.SLQ001, DC.SLQ001 A Design Project Plan (this document). Review Risk Management Plan (RM-0018 Rev D) for applicability to valve geometry change | Yes | 17 Mar 2026 | Ethan Rao | Complete | DCO089 |
| Phase 2: Design Input/Output Documentation | Document the affected design input: the valve dimensional specification (≤26 mm from sleeve to valve tip acceptance criterion). Document the design output: the revised device drawings and specifications reflecting the dimensional refinement. Confirm design input/output traceability for the affected component | No | 31 Mar 2026 | Ethan Rao | Complete | Revised device drawings included in Design Control Folder. |
| Phase 3: Risk Management Review | Review RM-0019 Rev D (Hazard Analysis Risk Management Report) for valve geometry change coverage, Review RM-0020 Rev D (DFMEA Risk Management Report) for valve geometry change coverage. Document risk assessment conclusion — confirmation memo or addendum | No | 07 Apr 2026 | Ethan Rao | Complete | Risk Management Review Confirmation Memo — DC.SLQ001 |
| Phase 4: Design Verification | Compile verification evidence: containment inspection data from CAPA 001-2025 (100% visual/dimensional inspection of Lots SLQ-11202024 and SLQ-05012025), Compile verification evidence: receiving inspection results for lots manufactured with modified valve, Document design verification summary confirming design outputs meet design inputs | No | 16 Apr 2026 | Ethan Rao | Complete |  |
| Phase 5: Regulatory Determination and Design Review | Finalize Letter-to-File: Regulatory Assessment — Valve Geometry Modification, conduct Design Review (Phase Review per QM.SLQ008), Submit all DC.SLQ001 deliverables under DCO 089, Close CAPA 003-2025 Corrective Action #2 with reference to completed design assessment. | Yes | 29 Apr 2026 | Ethan Rao | Initiate upon completion of Phase 4 |  |