Design Assessment, Valve Component Modification, ClearTract Catheter

Design Input / Output (I/O) Matrix

# Purpose

## The purpose of this document is to provide traceability between the affected design inputs, design outputs, and design verification activities for DC.SLQ001, a retrospective design assessment of a supplier-initiated valve geometry modification to the check valve assembly (Item No. 5 per device drawing) of the SILQ ClearTract Catheter product family.

## This I/O matrix traces each affected design input to its corresponding design output and to the verification evidence demonstrating that the design output meets the design input. This document is prepared per QM.SLQ007 A (Design Output SOP) Section 7.7 and QM.SLQ006 A (Design Input SOP) Section 8.

# Scope

## The scope of this document is limited to the design inputs, design outputs, and verification activities affected by the valve geometry modification to the check valve assembly of the SILQ ClearTract Catheter. This is a retrospective design assessment of a minor dimensional refinement to a single non-patient-contacting component. No changes were made to materials, valve type, functional mechanism, intended use, or performance claims. The modification improves valve seating consistency and reduces dislodgement probability. This matrix does not replicate the full device-level design I/O traceability for the ClearTract Catheter product family. Only the design inputs directly affected by the valve geometry change, and the design outputs and verification evidence corresponding to those inputs, are included.

# Documents

## Reference Documents

### QM.SLQ006	Design Input SOP

### QM.SLQ007	Design Output SOP

### QM.SLQ008	Design Review SOP

### QM.SLQ009	Design Verification and Validation SOP

### QM.SLQ012	Risk Management SOP

### DC.SLQ001	A Design Project Plan

### FM1-QM.SLQ004 A Design Project Scope Form DC.SLQ001

### CAPA 001-2025

### CAPA 003-2025

## Appendices / Associated Forms

### SLQ-6000 Rev B -- Catheter, 14 Fr x 10 mL, Silicone, NY Foley, 2-Way

### SLQ-6002 Rev B -- Catheter, 18 Fr x 10 mL, Silicone, NY Foley, 2-Way

### SLQ-6004 Rev B -- Catheter, 16 Fr x 10 mL, Silicone, NY Foley, 2-Way

### NCR-25-079 (Pathway MedTech NCR, 100% inspection of Lots SLQ-11202024 and SLQ-05012025)

### CHR-SLQ-6002 Rev A, LN 20250324, QTY 5500 (Receiving inspection record)

### Risk Management Review Confirmation Memo -- DC.SLQ001

# Definitions

## Abbreviations / Acronyms

### CHR: Component History Record (receiving inspection record)

### CAPA: Corrective Action / Preventive Action

### DHF: Design History File

### DFMEA: Design Failure Mode and Effects Analysis

### DMR: Device Master Record

### I/O: Input / Output

### LTF: Letter-to-File

### NCR: Nonconformance Report

### V&V: Verification and Validation

## General Definitions

### Check Valve Assembly: A non-patient-contacting component of the ClearTract Catheter (Item No. 5 per device drawing) that prevents backflow of inflation fluid and supports balloon inflation/retention during indwelling use. Located at the proximal balloon inflation port.

### Valve Geometry Modification: A dimensional/geometry refinement to the check valve assembly implemented by the component supplier (Ningbo Yingmed) to improve seating consistency and reduce dislodgement probability. No material, functional, or performance changes were made.

# Input / Output Matrix

## The following I/O matrix traces only the design inputs affected by the valve geometry modification assessed under DC.SLQ001. Three attributes of the check valve assembly are traced: (1) the dimensional acceptance criterion, (2) the mechanical retention requirement, and (3) balloon inflation integrity as a confirmation that the modification did not adversely affect an interfacing function. All three attributes were verified through the inspection activities documented in CAPA 001-2025 and subsequent receiving inspections.

Design I/O Matrix Table:

| Item |  | Attribute / Characteristic | User Needs | Verifiable Product Requirement (Design Input) | Traceability Information |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  | Design Output | Verification / Validation | Test Method |
| Finished Device – <Insert Device Name> |  |  |  |  |  |  |  |
| 1. | Check valve assembly dimensional specification -- maximum length from sleeve (Item 6) to valve tip |  | The catheter retention balloon shall inflate and deflate reliably. The check valve shall remain properly seated within the catheter neck connector throughout the intended use period to maintain balloon function. | The total length from the sleeve (Item 6) to the end of the check valve (Item 5) shall not exceed 26 mm. (Drawing Note 5.5, applicable to all ClearTract configurations.) | SLQ-6000 Rev B, Note 5.5; SLQ-6002 Rev B, Note 5.5; SLQ-6004 Rev B, Note 5.5. (Note 5.5 was added in Rev B to formalize this acceptance criterion on the device drawing. The 26 mm maximum was an existing receiving inspection criterion prior to the modification; Rev B codified it as a drawing note.) | (a) NCR-25-079: 100% dimensional inspection of existing finished goods Lot SLQ-11202024 (585 of 655 units accepted) and Lot SLQ-05012025 (870 of 912 units accepted). Nonconforming units were scrapped. (b) CHR-SLQ-6002 Rev A, LN 20250324, QTY 5500: Receiving inspection of post-modification lot at Pathway MedTech per drawing SLQ-6002 Rev A acceptance criteria. All 22 sample units passed dimensional inspection. | 100% visual/dimensional inspection (NCR-25-079, CAPA 001-2025 containment). AQL 4.0 sampling per drawing general note 1.3 for routine receiving inspection (CHR). Dimensional measurement using calibrated ruler |
| 2 | Check valve mechanical retention -- valve properly secured to catheter neck connector |  | The catheter retention balloon shall inflate and deflate reliably. The check valve shall not dislodge from the catheter neck connector during normal handling, use, or removal. | The check valve (Item 5) shall be properly secured to the connector. (Drawing Note 5.5, applicable to all ClearTract configurations. | SLQ-6000 Rev B, Note 5.5; SLQ-6002 Rev B, Note 5.5; SLQ-6004 Rev B, Note 5.5. (Note 5.5 was added in Rev B. The valve geometry modification was specifically intended to improve seating consistency and reduce dislodgement probability for this attribute.) | (a) NCR-25-079: 100% visual inspection of existing finished goods -- visual confirmation that the check valve is properly seated and secured in each inspected unit. Units with visible dislodgement or improper seating were rejected. (b) CHR-SLQ-6002 Rev A, LN 20250324, QTY 5500: Visual inspection of post-modification lot at Pathway MedTech. All 22 sample units passed visual inspection for valve seating. | 100% visual inspection (NCR-25-079, CAPA 001-2025 containment). AQL 4.0 sampling per drawing general note 1.3 for routine receiving inspection (CHR). Visual assessment per drawing Note 5.5 criteria. |
| 3 | Balloon inflation integrity -- balloon free of defects affecting inflation and retention function |  | The catheter retention balloon shall inflate to the specified volume without leakage, maintain inflation during the indwelling period, and deflate completely upon syringe aspiration. | The catheter shall undergo a leak test to verify the balloon is free of defects. (Drawing Note 5.4, per TM-0013 Foley Catheter Leak Test, applicable to all ClearTract configurations.) | SLQ-6000 Rev B, Note 5.4; SLQ-6002 Rev B, Note 5.4; SLQ-6004 Rev B, Note 5.4. (Note 5.4 is unchanged from Rev A. This row confirms the valve geometry modification did not adversely affect the balloon inflation function, which interfaces with the check valve assembly.) | CHR-SLQ-6002 Rev A, LN 20250324, QTY 5500: Leak test performed per TM-0013 on 22 sample units from the post-modification lot. All 22 units passed. | AQL 4.0 sampling per drawing general note 1.3. Leak test per TM-0013 (Foley Catheter Leak Test). This is a standard receiving inspection test confirming balloon integrity; it is included here to demonstrate that the valve geometry modification did not introduce adverse effects on the interfacing balloon inflation function. |