FileHold Document Management Software System

Product Requirements Specification

# Purpose

## The purpose of this design inputs document is to define the design requirements for the software document management system FileHold.  This document specifies the product requirements (design inputs) for the functional, security, and interface features that have been identified for FileHold.

# Scope

## This design input document applies to FileHold which is used by SILQ as the document management software system to control and maintain Quality Management System documentation in accordance with the document control process.

# Documents

## Reference Documents

### SW.SLQ001		Software Validation Plan, FileHold

### ISO 13485:2016	Medical Devices – Quality Management Systems – Requirements for Regulatory Purposes (Section 4.2 Documentation requirements)

### 21 CFR Part 820.40	FDA Quality System Regulation, Document Controls

### 21 CFR Part 11 	Electronic Records and Electronic Signatures

## Associated Forms

### N/A

# Definitions

## Abbreviations / Acronyms

### CPU: Central Processing Unit

### DCO: Document Change Order

### RAM: Random Access Memory

### URL: Uniform Resource Locator (web address)

### QMS: Quality Management System

## General Definitions

### FileHold:  Software based document management system used to electronically store controlled documents.

### Harm: Physical injury and/or damage to health of people, damage to property, or damage to the environment.

### Metadata:  Provides key details of a document file such as document type or code, schema, document category, document control number, document title, document revision, effective date, DCO number; used to summarize basic information about data which can make tracking and working with specific data easier.

### Risk: The combined likelihood and severity of harm.

### Risk Analysis: The systemic use of available information on a device or process to identify hazards and estimate risk.

### Risk Assessment: The overall process of risk analysis and risk evaluation.

### Risk Control:  An action taken to reduce the likelihood of a hazard or to reduce the effect from the hazard, or both.

### Schema: Schemas are the highest level of the metadata structure. Schema is a way to group like files together even when they are filed in disparate places within the FileHold library. Schemas manage how files are added to the library and what information is collected about them via the metadata.

### The schema assigned to a document will determine what metadata fields will be required and the format that the metadata will take.

# Product Overview

## FileHold is a third party, off-the-shelf software document management system used to control and maintain documentation in SILQ’s document control system as part of the Quality Management System. Documents are reviewed and approved through FileHold and released for use within SILQ’s QMS. FileHold is used to retain all QMS documents and records. Access to documents is determined by relevant job function and may be limited accordingly.

## FileHold is used by SILQ’s employees and subcontractors. Each SILQ employee and subcontractor is granted secure access with unique user identification (user name and password).

# Assumptions and Dependencies

## It is assumed that FileHold software is installed in accordance to manufacturer’s specifications.

# Intended Use and Process Definition

## The intended use of FileHold is to provide a secure means to control, approve and maintain electronic documentation, including records, as part of SILQ’s document control process in accordance with ISO 13485, 21 CFR 820, and 21 CFR Part 11.

## The following is the process definition for which FileHold is used:

### Submit document change orders (DCOs) in FileHold for review and approval of controlled documents.

### As a DCO approver, review documents and approve using electronic signature.

## Maintain all controlled documentation, including records, in electronic form on a secure, cloud-based server.

# Risk Analysis

## Risk Assessment

|  | Risk assessment question | Yes / No |
| --- | --- | --- |
| 1.1 Product safety (harm) | Is there a potential risk to product safety if the software malfunctions? - Patient harm - Operator harm - Bystander harm - Environmental harm | No |
| 1.2 Product safety (harm) | Is there a potential risk to product safety if the user of the software makes a mistake? - Patient harm - Operator harm - Bystander harm - Environmental harm | No |
| 2.1 Product quality | Is there a potential risk to product quality (other than safety) if the software malfunctions? | No |
| 2.2 Product quality | Is there a potential risk to product quality (other than safety) if the user makes a mistake? | No |
| 3.1 Record integrity | Is there a potential risk to record integrity in a system that is a record repository? - Record loss - Record corruption | Yes |
| 4.1 Demonstration of compliance to an FDA regulation or ISO standard | Is there a potential risk regarding the ability to demonstrate regulatory compliance? - Record loss - Record corruption | Yes |

## Analysis of process risk

### Failure of the process, or FileHold, could result in the following:

### Unauthorized modification of controlled documentation

### Loss/corruption of manufacturing documentation (Device Master Record)

### Loss/corruption of device history records

### Loss/corruption of Quality Management System documentation and records

### Risk control measures in place to mitigate the identified risks include:

### Employees are assigned to User Groups that prohibit inadvertent or unauthorized modification or deletion of controlled documents and records

### Nightly backups of data

### Antivirus and monitoring of all components of the FileHold installation including URL availability, CPU, Disk and RAM utilization, system uptime, reboots and hardware status of system server

### Continuous vulnerability scanning of system server

### Redundant Cisco firewall maintenance and patching of system server

## A failure of the process would not result in direct or indirect harm to the patient. Also, the risk controls in place would prevent any loss of data in the unlikely event that multiple failures occur simultaneously. The risk category for use of FileHold software is Minor.

# Product Requirements

| ID # | FileHold Requirement Specification (FRS) |
| --- | --- |
| 1 | Library Structure, User Roles and Access Requirements |
| FRS-1.1 | FileHold shall be accessible through the web client via standard web browsers such as Chrome, Safari, Internet Explorer, Edge and Firefox. |
| FRS-1.2 | Electronic documentation shall be stored in appropriate Cabinet > Drawer > Folder. Documents can only be stored in a folder. |
| FRS-1.3 | User access to documentation shall be assigned by the system administrator based on their role and responsibility within the company. |
| FRS-1.4 | User shall be able to access cabinets of which they are a member. System Administrator assigns access to cabinets by assigning user group. |
| FRS-1.5 | If added as a member of a cabinet, users shall be able to: - View the cabinet; - View drawers of the cabinet; - View folders located in the drawers if they are a member of the folder; and - View documents of which they are a member of the schema. |
| FRS-1.6 | System Administrator shall assign members to each schema. This allows members access to documents of that schema. |
| FRS-1.7 | In order to access a document, the user must be a member of the schema. |
| FRS-1.8 | Every document maintained in FileHold shall have a schema assigned to it. The schemas include: Controlled Documents, Document Change Order, Draft Documents, Log, Quality Record and Standards. |
| FRS-1.9 | “Controlled Documents” schema shall include the following metadata fields: Document Control Number, Document Title, Document Revision, Document Category, Document Type or Code, Training Required?, DCO Number, Effective Date, and Release Date. |
| FRS-1.10 | FileHold shall require the following metadata fields to be completed when adding a document to “Controlled Documents” schema: Document Control Number, Document Title, Document Revision, Document Category, Document Type or Code, Training Required?, and DCO Number. |
| FRS-1.11 | “Document Change Order” document schema shall include the following metadata fields: DCO Number, Originator, Risk Assessment?, Verification or Validation Required?, Material Disposition Required?, Potential Regulatory Impact?, Release Date, Effective Date, and DCO Notes. |
| FRS-1.12 | FileHold shall require the following metadata fields to be completed when adding a document to “Document Change Order” schema: DCO Number, Originator, Risk Assessment?, Verification or Validation Required?, Material Disposition Required?, Potential Regulatory Impact? |
| FRS-1.13 | “Quality Record” schema shall include the following metadata field: Document Title. |
| FRS-1.14 | FileHold shall not require metadata field to be completed when adding a document to “Quality Record” schema. |
| FRS-1.15 | “Draft Documents” schema shall include the following metadata field: Document Control Number and Document Title. |
| FRS-1.16 | FileHold shall not require metadata fields to be completed when adding a document to “Draft Documents” schema. |
| FRS-1.17 | Documents requiring approval per the document control process shall be reviewed and approved through the DCO Approval workflow; documents requiring an e-signature shall be e-signed through the Training workflow and Trainer Signature workflow. |
| FRS-1.18 | Depending on the role and responsibilities of the FileHold user within the company, users shall be assigned to one of the following user groups: Read Only, Document Publisher, Publisher, and System Administrator. Each user group has a defined set of privileges and restrictions. |
| FRS-1.19 | Read-Only user role shall have the ability to download (or open) documents from FileHold. Read-Only users shall be restricted from downloading documents or viewing documents in FileHold Viewer if this functionality is disabled by System Administrators. |
| FRS-1.20 | Read-Only user role shall have the ability to participate in workflows as a reviewer/approver. |
| FRS-1.21 | Read-Only user role shall not have the ability to edit metadata, add documents or delete documents from FileHold, or initiate workflows. |
| FRS-1.22 | Document Publisher user role shall have the ability to: - Download (or open) documents from FileHold; - Add documents to FileHold; - Check-in documents to FileHold/check-out documents from FileHold; - Edit metadata; - Move documents that are owned by them to different locations; and - Initiate workflows. |
| FRS-1.23 | Document Publisher user roles shall not be able to delete any documents including those which they have added to the system. |
| FRS-1.24 | Publisher user role shall have the ability to: - Download (or open) documents from FileHold; - Add documents to FileHold; - Check-in documents to FileHold/check-out documents from FileHold; - Edit metadata; - Move documents that are owned by them to different locations; - Initiate workflows; - Create new folders or rename folders that they own within an existing folder group; - Copy or move folders that they have already created; and - Clone folders and folder groups created by other users and become the owners of the folders. |
| FRS-1.25 | Publisher user roles shall not be able to delete any documents, folders or folder groups including those which they have added to the system. |
| FRS-1.26 | System Administrator user role shall have the ability to: - Download (or open) documents from FileHold; - Add documents to FileHold; - Check-in documents to FileHold/check-out documents from FileHold; - Edit metadata; - Move documents to different locations; - Initiate workflows; - Create new folders or rename folders; - Copy or move folders; - Clone folders and folder groups created by other users and become the owners of the folders; - Delete folders and files; - Add users to the system (including assigning the initial password and setting requirements for all new passwords); - Add users to user groups |

| ID # | FileHold Requirement Specification (FRS) |
| --- | --- |
| 2 | Security Requirements |
| FRS-2.1 | User authentication (unique username and password) shall be required to log into FileHold. |
| FRS-2.2 | Users shall be logged off of their secure session after a configurable period of time. The default period of time is 60 minutes. |
| FRS-2.3 | After a configuration number of unsuccessful login attempts, the system shall disable the user account and an email alert is sent to all system administrators. The system administrator will need to enable the account in the Users area and if the user is a local user, reset their password. The default number of unsuccessful login attempts is 8. |
| FRS-2.4 | All actions performed by a user in the system shall be recorded in an audit log which includes the date and time the action was performed. Recorded actions include: check out, check in, downloaded, emailed, linked, copied, viewed, and deleted. |
| FRS-2.5 | The Library Audit Log shall log the following information: - Deleting document versions - Deleting library objects such as cabinets, drawers, folder groups and folders - Deleting schemas - Deleting metadata fields and removing metadata fields from schemas - Adding, removing, and modifying cabinet and folder memberships - Added, changed or deleted in a drop-down metadata field - Changing the owner of a cabinet, folder, or document version - Creating, modifying, deleting, and cloning a workflow template - Creating library objects such as cabinets, drawers, folder groups, and folders. |
| FRS-2.6 | The System Audit Log shall log the following information: - Adding users - Deleting users - Modifying user accounts and FileHold groups - Adding and deleting FileHold groups - Resetting passwords - Adding and removing users to and from FileHold groups - Change in general system settings - Change in email settings - Change security settings - Change search settings |