In-Vitro Bacterial Adhesion Comparative Study

Test Report

# Purpose

## The purpose of this report is to summarize the comparative study used to demonstrate the reduced rate of bacterial adhesion on Foley catheters due to the Silq’s Surface Treatment relative to commercial counterparts. This report outlines the test method and results of the comparative study.

# Scope

## This comparative study applies to Silq’s Surface Treatment process.

## Testing was conducted on 16 Fr Foley catheters treated with Silq’s coating and compared against commercial counterparts by measuring the level of bacterial adhesion on the catheters after being incubated in a broth of E. coli cells that have been clinically isolated from CAUTI patients.

# Definitions

## Abbreviations / Acronyms

### ATP: Adenosine Triphosphate

### CAUTI:  Catheter-associated Urinary Tract Infections

### mL: milliliter

### mM: milli Molar

### OD600: Optical Density at 600nm

### PBS: Phosphate Buffer Solution

### RLU: Relative Light Units

### RPM: Revolutions per Minute

### uL: microliter

# Background

## Bacterial biofilms are responsible for clinical failures of many medical devices such as urinary catheters and are associated with many infectious complications1.2. Catheter coatings have been introduced into the market to prevent these infections, however, preclinical and clinical evaluations need to accurately quantify the bacterial load in the biofilm in vitro and ex vivo. There is no uniform gold standard for biofilm quantification for different surfaces and established biofilms. In order to evaluate the effectiveness of a coating, it is necessary to determine the amount of active biomass present on the substrate. Because ATP is present in all active microorganisms, monitoring ATP levels can be useful for predicting the active biomass content. The assay used in this comparative study combines two published methods to determine microbial activity and associated biofilm presence on the surface of catheters.

## In order to test the ATP levels, the adhered cells needed to be extracted from the catheter surfaces. This is achieved by the ‘Vortex-Sonicate-Vortex’ method3,4. Vortexing before sonication helps in dislodging any loosely attached layers of biofilm and enhances the effect of sonication on deep-seated layers that are strongly attached to the catheter surfaces. Further vortexing after limited sonication additionally helps in breaking down bacterial communities into individual cells and improves bacterial isolation thereby enabling a more accurate and uniform quantification without destroying bacterial cells excessively. This procedure has been published in various peer reviewed journals and has been optimized by our team and a 3rd party testing facility (Oak Crest Institute of Science) to ensure complete removal of cellular material from foley catheter surfaces after incubation in clinically isolated bacterial solutions.

## After extraction of the surface-attached bacteria, the sample is subjected to the BacTiter-Glo™ Microbial Cell Viability Assay to measure bacteria present in the sample. The BacTiter-Glo™ reagent solution primarily contains two proprietary active ingredients. The first is a detergent mixture which induces cell lysis. The second ingredient is a molecule which, when in the presence of Adenosine Triphosphate (ATP), expresses a luminescent signal which can be detected by common laboratory luminometers to provide a relative quantification of cellular activity5-9. Since bacterial cells maintain a relatively constant intracellular ATP concentration, this assay provides a reproducible relative measure of attached cells.

## Trautner BW, Darouiche RO. Role of biofilm in catheter-associated urinary tract infection. Am J Infect Control. 2004;32(3):177-183. doi:10.1016/j.ajic.2003.08.005

## Nargis Sabir, Aamer Ikram, Gohar Zaman, Luqman Satti, Adeel Gardezi, Abeera Ahmed, Parvez Ahmed. Bacterial biofilm-based catheter-associated urinary tract infections: Causative pathogens and antibiotic resistance American Journal of Infection Control.2017; 45(10):1101-1105. doi: 10.1016/j.ajic.2017.05.009

## Mandakhalikar, K.D., Rahmat, J.N., Chiong, E. et al. Extraction and quantification of biofilm bacteria: Method optimized for urinary catheters. Sci Rep 8, 8069 (2018). https://doi.org/10.1038/s41598-018-26342-3

## 4) Rosa, Luigi & Lepanto, Mariastefania & Cutone, Antimo & Berlutti, Francesca & de Angelis, Massimiliano & Vullo, Vincenzo & Mastroianni, Claudio & Valenti, Piera & Oliva, Alessandra. (2019). BioTimer assay as complementary method to vortex-sonication-vortex technique for the microbiological diagnosis of implant associated infections. Scientific Reports. 9. 10.1038/s41598-019-44045-1.

## https://www.promega.com/resources/pubhub/promega-notes-2008/use-of-the-bactiter-glo-microbial-cell-viability-assay-to-study-bacterial-attachment-in-biofilm/

## https://www.promega.com/resources/pubhub/promega-notes-2007/atp-measurement-as-a-means-for-directly-estimating-active-biomass/

## Berney M, Weilenmann HU, Egli T. Flow-cytometric study of vital cellular functions in Escherichia coli during solar disinfection (SODIS). Microbiology (Reading). 2006 Jun;152(Pt 6):1719-1729. doi: 10.1099/mic.0.28617-0. PMID: 16735735.

## Fan F, Wood KV. Bioluminescent assays for high-throughput screening. Assay Drug Dev Technol. 2007 Feb;5(1):127-36. doi: 10.1089/adt.2006.053. PMID: 17355205.

## Necchi F, Saul A, Rondini S. Setup of luminescence-based serum bactericidal assay against Salmonella Paratyphi A. J Immunol Methods. 2018 Oct;461:117-121. doi: 10.1016/j.jim.2018.06.025. Epub 2018 Jun 30. PMID: 29969587.

# Test Method

## Bacterial Solution Preparation: Biomass is introduced to the foley catheters by exposing them to bacterial solution. Bacterial solution was prepared as follows.

### A clinical isolate of E. coli cells was obtained from the UCLA Department of Pathology and Laboratory Medicine.

### Cells were incubated for 2 days at 37oC in sterilized LB Broth.

### An aliquot of the cellular suspension was removed and diluted with LB broth until the OD600 reading was 0.1. This is a measure of the concentration of E. coli in solution. The threshold was established as a benchmark across different experiments.

## Catheter Segment Preparation

### The following types of 16 Fr Foley catheters were evaluated in this study:

#### Latex catheter (no surface or coating treatment)

#### Silver coated catheter

#### Silq treated catheter

### Two 1.5” pieces were cut from each type of 16 Fr Foley catheter.

### Each catheter segment was placed in an individual scintillation vial which was then filled with 10mL of the previously prepared bacterial solution.

### All vials were placed into an incubator set to 37oC spinning at 120 RPM  for 24 hours.

### After 24 hours, all catheter segments were removed and placed into 15mL falcon tubes containing 10mL of DI water.

### Falcon tubes were gently agitated by inverting each tube 3 times to remove all loosely adhered cells.

### Catheter segments were then removed from the falcon tubes and gently tapped onto lint-free wipes to remove any excess liquid remaining in the inner lumens. The catheter segments were then placed into a new 15mL falcon tube containing 8mL of sterile 1X PBS buffer.

## Measuring the presence of bacterial biomass using Adenosine Triphosphate Measurement Assay

### The BacTiter-Glo™ reagent (Purchased from Promega Inc) was removed from the freezer and allowed to thaw to room temperature.

### In order to fully dislodge the adhered bacterial cells from the catheter segments, each falcon tube was vortexed for 3 minutes (Scientific Industries Vortex Genie-2), a tip probe sonicator (Virsonic) was then inserted into the falcon tubes and turned on at 10mA for 2 minutes, followed by another 3 minutes of vortexing.

### Catheter segments were then removed from the falcon tubes.

### 3 aliquots of 100uL were taken from each sample falcon tube and placed into individual wells in a 96 well plate.

### To all wells containing solution, a 100ul aliquot of the BacTiter Glo™ reagent was added and allowed to sit for 1 minute before the luminescence of each well was measured using a luminometer (Molecular Devices SpectraMAX, GEMINI XS).

# Results

## Oak Crest Institute of Science performed the ATP measurement assay on the three types of catheters and the raw data can be found in Attachment 1 Bacterial Adhesion Study Data. Rows denoted H1 and H2 refer to Silq surface treated catheters, S1 and S2 refer to silver coated catheters, and L1 and L2 refer to latex catheters.

## A standard curve was established pairing known concentrations of ATP and their luminescence readings. After establishing this curve, the bacterial biomass of a sample was determined based on the luminescence reading of that sample.  The ATP Standard Curve was obtained as follows:

### Solutions of ATP at known concentrations (0mM, 1 mM, 10mM, 50mM, 100mM, 250mM, and 1000mM) were prepared in PBS buffer solution.

### Three separate 100 uL aliquots of each standard solution were taken and placed into wells with 100uL of the BacTiter-Glo™ reagent solution.

### A luminescence reading of each well was taken and the results are shown in Table 7.1.

### After subtracting out the average background luminescence (wells with 0mM ATP) from the average RLU at each concentration, a standard curve (RLUs vs mM ATP) was made with a trendline of y= 0.6558x – 2.14 (R2=1) as shown in Graph 7.1.

### Table 7.1

| ### Concentration ATP (mM) | ### Luminescence ### Replicate 1 (RLU) | ### Luminescence ### Replicate 2 (RLU) | ### Luminescence ### Replicate 3 (RLU) | ### Avg Relative Light Units (RLUs) |
| --- | --- | --- | --- | --- |
| ### 0 | ### 53.083 | ### 48.54 | ### 46.879 | ### 49.5 |
| ### 1 | ### 54.741 | ### 57.961 | ### 50.809 | ### 54.5 |
| ### 10 | ### 62.959 | ### 75.216 | ### 74.647 | ### 70.94 |
| ### 50 | ### 132.362 | ### 139.539 | ### 125.055 | ### 132.32 |
| ### 100 | ### 202.773 | ### 202.817 | ### 206.817 | ### 204.14 |
| ### 250 | ### 427.471 | ### 418.138 | ### 435.675 | ### 427.09 |
| ### 1000 | ### 1547.694 | ### 1511.154 | ### 1679.093 | ### 1579.31 |

### Graph 7.1

## Three different types of Foley catheter segments were prepared by Silq in accordance with Sections 6.1 and 6.2. Catheter types included a latex catheter, silver coated catheter, and Silq treated catheter. The luminescence of each well containing sample solution was then measured. Results are summarized in Table 7.2.

### Table 7.2

| ### Sample | ### Luminescence Replicate 1 (RLU) | ### Luminescence Replicate 2 (RLU) | ### Luminescence Replicate 3 (RLU) | ### Avg RLUs | ### STDEV |
| --- | --- | --- | --- | --- | --- |
| ### Silq Coated Catheter A | ### 70.136 | ### 74.273 | ### 67.391 | ### 70.6 | ### 3.46 |
| ### Silq Coated Catheter B | ### 89.129 | ### 76.698 | ### 74.67 | ### 80.17 | ### 7.83 |
| ### Silver Coated Catheter A | ### 127.371 | ### 118.515 | ### 121.133 | ### 122.34 | ### 4.55 |
| ### Silver Coated Catheter B | ### 125.068 | ### 122.476 | ### 134.468 | ### 127.34 | ### 6.31 |
| ### Latex Catheter A | ### 263.495 | ### 262.838 | ### 259.803 | ### 262.05 | ### 1.97 |
| ### Latex Catheter B | ### 304.831 | ### 300.087 | ### 311.137 | ### 305.35 | ### 5.54 |

## The average RLUs was then calculated for each sample, the background luminescence was subtracted from that value, and the concentration of ATP present in each sample tube was calculated from the standard curve. Results are summarized in Table 7.3.

### Table 7.3

| ### Sample | ### Avg RLUs | ### Avg RLUs-BG | ### ATP Concentration (mM) |
| --- | --- | --- | --- |
| ### Silq Coated Catheter A | ### 70.6 | ### 21.1 | ### 11.69 |
| ### Silq Coated Catheter B | ### 80.17 | ### 30.67 | ### 17.97 |
| ### Silver Coated Catheter A | ### 122.34 | ### 72.84 | ### 45.63 |
| ### Silver Coated Catheter B | ### 127.34 | ### 77.84 | ### 48.90 |
| ### Latex Catheter A | ### 262.05 | ### 212.55 | ### 137.25 |
| ### Latex Catheter B | ### 305.35 | ### 255.85 | ### 165.64 |

## The average ATP present on each catheter type was then calculated as shown in Table 7.4. Based on the results, catheters treated with Silq’s coating surface treatment process demonstrate the ability to limit the amount of bacteria by a factor of 3x when compared to silver coated catheter which are commonly used by clinicians to mitigate infections in their patients.

### Table 7.4

| ### Catheter Type | ### Avg ATP Concentration (mM) |
| --- | --- |
| ### Silq Coated Catheters | ### 14.83 |
| ### Silver Coated Catheter | ### 47.27 |
| ### Latex Catheter | ### 151.45 |

# Analysis

## The presence of bacteria of each catheter type was compared against each other using a two sample t-test (assuming equal variance).

## For test results which were found to be "statistically significant", the data had to show that the measured luminescence (RLU) in the test catheter was greater than the measured luminescence observed in the control catheter.  For results significant at P < 0.05, there was a 95% certainty that there was a difference between test catheter and control catheter.  For those results significant at P < 0.01, there was a 99% certainty that there was a difference between test catheter and control catheter.

## Comparison between the silver coated catheter (test) and the latex catheter (control)

### The results of the two sample t-test calculated a P value of 2.1 x 10-8 for silver coated catheter, demonstrating a statistically significant difference (with 99% certainty) in the presence of bacteria that remained on the latex catheter compared to the silver coated catheter.

### The results verify the capability of the silver coated catheter to retain statistically significant less bacteria than the latex catheter.

## Comparison between the Silq treated catheter (test) and the latex catheter (control)

### The results of the two sample t-test calculated a P value of 1.9 x 10-9 for Silq treated catheter, demonstrating a statistically significant difference (with 99% certainty) in the presence of bacteria that remained on the latex catheter compared to the Silq treated catheter.

### The results verify the capability of the Silq treated catheter to retain statistically significant less bacteria than the latex catheter.

## Comparison between the Silq treated catheter (test) and the silver coated catheter (control)

### The results of the two sample t-test calculated a P value of 1.5 x 10-7 for Silq treated catheter, demonstrating a statistically significant difference (with 99% certainty) in the presence of bacteria that remained on the silver coated catheter compared to the Silq treated catheter.

### The results verify the capability of the Silq treated catheter to retain statistically significant less bacteria than the silver coated catheter.

## The ATP concentrations of each catheter type were averaged and the results showed that the Silq treated catheter had 3.2x less bacteria than the silver coated catheter and 10.2x less bacteria than the latex catheter.

# Conclusion

## Infections are common in medical devices such as urinary catheters. Although a more expensive option than non-coated catheters, silver coated urinary catheters are used by clinicians in order to prevent urinary tract infections or other bacteria-induced infections. The results of this study demonstrated that the Silq treated catheter retained significantly less bacteria than the silver coated catheter and also confirmed the silver coated catheter retained less bacteria than the non-coated catheter.