Crystal Violet Absorbance,

Silq Surface Treatment

Test Protocol

# Purpose

## The purpose of this report is to detail the testing protocol used to verify the presence of Silq’s Surface Treatment coating using a crystal violet stain.

# Scope

## This test protocol applies to the Silq Surface Treatment process.

## The Silq Surface Treatment process may be applied to a variety of polymeric substrates. Testing in this procedure was conducted on 16 Fr Foley catheters that have been coated using the Silq Surface Treatment process.

# Definitions

## Abbreviations / Acronyms

### cm: Centimeter

### DI: Deionized

### in: Inches

### mg: Milligram

### mL: Milliliter

### nm: Nanometer

# Background

## Crystal violet is a triarylmethane dye commonly used as a histological stain to determine the presence of proteins and DNA in biological research. It has been found to have a strong interaction with the polymers present in the Silq Surface Treatment. When surface treated materials are immersed in a crystal violet solution, the dye intercalates into the polymer network present on the surface. Subsequently, the stained substrate is immersed in an acetic acid solution which fully dissolves the crystal violet molecules allowing them to diffuse out of the network and into the solution. Through measurement of the absorbance of the resulting solution at 585nm, a relative measure of the amount of zwitterionic polymer present on the surface of the substrate can be obtained.

# Test Method

## The following equipment and supplies were required for the execution of this test.

### ES-0092-01 – Vis Spectrophotometer calibrated using Neutral Density Calibration Filters (FireflySci serial No. 886)

### Crystal Violet Dye Solution (4 mg/mL)

### 2:1 DI Water: Acetic Acid Solution

### 50 mL Falcon Tubes

### Glass test Tubes (1cm inner diameter, 45cm length)

### X-Acto Knife or Equivalent

### Inch Ruler, Calibrated

### Timer, Calibrated

### Non-Latex Gloves

### Metal Tray

## The spectrophotometer should be powered up and allowed to warm up for at least 15 minutes with the wavelength set to 585nm

## Test procedure

### Tester wears gloves when performing this procedure.

### Secure 3 glass tubes with an inner diameter of 1cm and length of 45cm.

### Fill one tube with a well stirred 4mg/mL crystal violet solution to the demarcated line.

### Fill the other two tubes with DI Water to the demarcated line.

### Take a dried, surface treated Foley catheter sample and immerse it in the tube containing the crystal violet solution. Allow the sample to soak for 1 minute.

### Slowly remove the catheter sample from the tube allowing for all crystal violet solution to drain. Immediately dip the catheter sample into a tube containing DI water and again slowly remove the catheter sample allowing for all residual liquid to drain. Repeat with the 2nd tube of DI water.

### Place the catheter sample on a clean, dry, metal surface.

### Using an X-Acto knife or razor blade, cut off the tip of the catheter sample just above the islet at the bottom.

### Using a ruler, measure 10” from the end of the catheter sample and make another cut.

### Cut the 10” catheter sample into five 2”  segments. Use tweezers to place all segments into a 50mL falcon tube containing 25 mL of the 2:1 DI water: Acetic acid solution.

### Shake the falcon tube for 30 seconds until all dye has been removed from the catheter segments.

### Pour 2 mL of the solution into a clean cuvette.

### Separately, prepare a cuvette containing 2 mL of the 2:1 DI water:acetic acid solution. Use this cuvette to blank the spectrophotometer at 585nm.

### Measure the absorbance value of the sample cuvette and document results.

## Sample size

### A sufficient number of test samples to satisfy a 90%/90% confidence/reliability interval using variable data analysis.

## Acceptance Criteria

### The absorbance value shall be a minimum of 0.400.

### Rationale: This acceptance criteria threshold was derived from previous testing where the amount of coating resulting in an absorbance of 0.300 was sufficient to produce the desired effect of reduced microbial adhesion on device surfaces. Testing also demonstrated that higher levels of coating will result in the same or better efficacy in reducing bacterial adhesion. Therefore, an additional safety factor of 33% was added to the absorbance of 0.300.

# Conclusion

## Crystal violet absorbance is an appropriate quality control test method for verifying the presence of the Silq Surface Treatment on treated substrates due to the intercalation of the crystal violet dye with the polymer network of the surface treatment.