Coating Adherence Verification,

Silq Surface Treatment

Test Protocol

# Purpose

## The purpose of this report is to summarize the testing protocol to evaluate if the coating using the Silq Surface Treatment process remains adhered to the substrate. The coated substrate is rinsed with deionized water and the rinse is examined for the presence of coating using spectroscopy.

# Scope

## This test protocl applies to the Silq Surface Treatment process.

## The Silq Surface Treatment process may be applied to a variety of polymeric substrates. Testing in this procedure was conducted on 16 Fr Foley catheters coated with the Silq Surface Treatment process.

# Definitions

## Abbreviations / Acronyms

### cm: Centimeter

### DI: Deionized

### in: Inches

### mL: Milliliter

### nm: Nanometer

# Background

## Rinse testing is used to verify if coating on a substrate continues to adhere to the substrate. There is no concern of toxicity or harm to the patient from the Silq Surface Treatment based on the results of biocompatibility studies. However, coating material that insufficiently adheres to the substrate may result in devices that do not meet the desired level of efficacy in reducing microbial activity on the device surface.

## Samples are briefly immersed in water repeatedly so that any loosely bound material would be dislodged from the surface of the device. The water is then tested for the presence of coating material imparted by the Silq Surface Treatment process via absorbance at 440nm. Absorbance values of greater than 0.02 would indicate that device samples were insufficiently rinsed.

# Test Method

## The following equipment and supplies were required for the execution of this study.

### ES-0092-01 – Vis Spectrophotometer calibrated using Neutral Density Calibration Filters (FireflySci serial No. 886)

### DI Water

### Glass test tubes (inner diameter of 1cm, length 40cm)

### Non-Latex Gloves

### Cuvettes

## The spectrophotometer was powered up and allowed to warm up for at least 15 minutes. The wavelength was set to 440 nm.

### Note: Machine must warm up for at least 15 minutes

## Test procedure

### Tester wear gloves when performing this procedure.

### Secure a glass tube with an inner diameter of 1cm and length of 45 cm and fill with DI water to the demarcated line.

### Take a dried, surface treated Foley catheter and immerse it in the tube. Gently lift the catheter out of the tube and solution and then resubmerge it for a total of five times.

### Remove the catheter from the tube.

### Using a gloved finger to plug the top of the tube, invert the tube at least 3 times vigorously to ensure that any materials present in solution will be homogenously distributed.

### Pour out 2mL of the solution into a clean cuvette.

### Separately, pour 2mL of pure DI water into a cuvette and use this cuvette to blank the spectrophotometer at 440 nm.

### Measure the absorbance value of the sample cuvette.

## Sample size

### A sufficient number of test samples were tested to satisfy a 90%/90% confidence/reliability interval using variable data analysis.

## Acceptance Criteria

### The absorbance value shall be a less than or equal to 0.02.

### Rationale: This value was arrived at due to measurement variability at absorbances below 0.05, and a reasonable acceptable tolerance of extremely low levels of loosely bound material due to the Silq Surface Treatment that would not result in any measurable loss in performance.

# Conclusion

## Absorbance measurements of solution resulting from a rinsing of samples subject to the Silq Surface Treatment process is an appropriate quality control test method for verification that the material deposited to the foley catheter surface is sufficiently bound and that only negligible levels of loosely bound material remains.