Crystal Violet Absorbance,

Silq Surface Treatment

Test Report

# Purpose

## The purpose of this report is to summarize the testing performed to verify the presence of Silq’s Surface Treatment coating using a crystal violet stain.

# Scope

## This test report applies to the Silq Surface Treatment process.

## The Silq Surface Treatment process may be applied to a variety of polymeric substrates. Testing in this procedure was conducted on 16 Fr Foley catheters that have been coated using the Silq Surface Treatment process.

# Definitions

## Abbreviations / Acronyms

### cm: Centimeter

### DI: Deionized

### in: Inches

### mg: Milligram

### mL: Milliliter

### nm: Nanometer

# Background

## Crystal violet is a triarylmethane dye commonly used as a histological stain to determine the presence of proteins and DNA in biological research. It has been found to have a strong interaction with the polymers present in the Silq Surface Treatment. When surface treated materials are immersed in a crystal violet solution, the dye intercalates into the polymer network present on the surface. Subsequently, the stained substrate is immersed in an acetic acid solution which fully dissolves the crystal violet molecules allowing them to diffuse out of the network and into the solution. Through measurement of the absorbance of the resulting solution at 585nm, a relative measure of the amount of zwitterionic polymer present on the surface of the substrate can be obtained.

# Test Method

## The test method for this report was adopted directly from the protocol detailed in document: VV.SLQ002 A Test Protocol, Crystal Violet Absorbance, Silq Surface Treatment

## Test samples

### Segments from twenty-one (21) 16 Fr Foley catheters treated with the Silq Surface Treatment process were tested satisfying a 90%/90% confidence/reliability interval using variable data analysis.

## Acceptance Criteria

### The absorbance value shall be a minimum of 0.400.

### Rationale: This acceptance criteria threshold was derived from previous testing where the amount of coating resulting in an absorbance of 0.300 was sufficient to produce the desired effect of reduced microbial adhesion on device surfaces. Testing also demonstrated that higher levels of coating will result in the same or better efficacy in reducing bacterial adhesion. Therefore, an additional safety factor of 33% was added to the absorbance of 0.300.

# Results

## Twenty-one (21) Foley catheters were treated with the Silq Surface Treatment. Preparation of the catheter samples and absorbance measurement was performed in accordance with Section 5.3.

## The absorbance results of the catheter samples are shown in Table 6.1.

## Table 6.1

| ## Catheter Sample # | ## Absorbance |
| --- | --- |
| ## 1 | ## 0.55 |
| ## 2 | ## 0.597 |
| ## 3 | ## 0.569 |
| ## 4 | ## 0.445 |
| ## 5 | ## 0.441 |
| ## 6 | ## 0.581 |
| ## 7 | ## 0.616 |
| ## 8 | ## 0.597 |
| ## 9 | ## 0.56 |
| ## 10 | ## 0.588 |
| ## 11 | ## 0.565 |
| ## 12 | ## 0.532 |
| ## 13 | ## 0.525 |
| ## 14 | ## 0.585 |
| ## 15 | ## 0.496 |
| ## 16 | ## 0.565 |
| ## 17 | ## 0.509 |
| ## 18 | ## 0.544 |
| ## 19 | ## 0.541 |
| ## 20 | ## 0.585 |
| ## 21 | ## 0.51 |
| ## Sample Mean | ## 0.548 |
| ## Standard deviation | ## 0.047 |

# Analysis

## Using variable data analysis with a single-sided specification, we assume a normal distribution.

¯x   - ks > minimum specification

## The sample mean of the absorbance is 0.548, the standard deviation is 0.047, and the minimum specification is 0.400. Given a sample size of 21, the tolerance interval factor ‘k’ of 2.994 demonstrates that there is 90% confidence that 99% of the population meets specification.

# Conclusion

## Crystal violet absorbance is an appropriate quality control test method for verifying the presence of the Silq Surface Treatment on treated substrates due to the intercalation of the crystal violet dye with the polymer network of the surface treatment.

## Based on the results of the crystal violet absorbance test, the presence of coating on catheter samples that underwent the Silq Surface Treatment process has been verified to meet acceptance criteria.