Coating Adherence Verification,

Silq Surface Treatment

Test Report

# Purpose

## The purpose of this report is to summarize the testing performed to evaluate if the coating using the Silq Surface Treatment process remains adhered to the substrate. The coated substrate is rinsed with deionized water and the rinse is examined for the presence of coating using spectroscopy.

# Scope

## This test report applies to the Silq Surface Treatment process.

## The Silq Surface Treatment process may be applied to a variety of polymeric substrates. Testing in this procedure was conducted on 16 Fr Foley catheters coated with the Silq Surface Treatment process.

# Definitions

## Abbreviations / Acronyms

### cm: Centimeter

### DI: Deionized

### in: Inches

### mL: Milliliter

### nm: Nanometer

# Background

## Rinse testing is used to verify if coating on a substrate continues to adhere to the substrate. There is no concern of toxicity or harm to the patient from the Silq Surface Treatment based on the results of biocompatibility studies. However, coating material that insufficiently adheres to the substrate may result in devices that do not meet the desired level of efficacy in reducing microbial activity on the device surface.

## Samples are briefly immersed in water repeatedly so that any loosely bound material would be dislodged from the surface of the device. The water is then tested for the presence of coating material imparted by the Silq Surface Treatment process via absorbance at 440nm. Absorbance values of greater than 0.02 would indicate that device samples were insufficiently rinsed.

# Test Method

## The test method for this report was adopted directly from the protocol detailed in document: VV.SLQ003 A Test Protocol, Coating Adherence Verification, Silq Surface Treatment

## Test samples

### Twenty-one (21) 16 Fr Foley catheters treated with the Silq Surface Treatment process were tested satisfying a 90%/90% confidence/reliability interval using variable data analysis.

## Acceptance Criteria

### The absorbance value shall be a less than or equal to 0.02.

### Rationale: This value was arrived at due to measurement variability at absorbances below 0.05, and a reasonable acceptable tolerance of extremely low levels of loosely bound material due to the Silq Surface Treatment that would not result in any measurable loss in performance.

# Results

## Twenty-one (21) Foley catheters were treated with the Silq Surface Treatment. Preparation of the catheter samples, rinse, and absorbance measurement were performed in accordance with Section 5.

## The absorbance results of the catheter sample rinses are shown in Table 6.1

## Table 6.1

| ## Catheter Sample # | ## Absorbance |
| --- | --- |
| ## 1 | ## 0.01 |
| ## 2 | ## 0.001 |
| ## 3 | ## 0.009 |
| ## 4 | ## 0.012 |
| ## 5 | ## 0.003 |
| ## 6 | ## 0.009 |
| ## 7 | ## 0.012 |
| ## 8 | ## 0.015 |
| ## 9 | ## 0.007 |
| ## 10 | ## 0.009 |
| ## 11 | ## 0.016 |
| ## 12 | ## 0.008 |
| ## 13 | ## 0.004 |
| ## 14 | ## 0.011 |
| ## 15 | ## 0.014 |
| ## 16 | ## 0.002 |
| ## 17 | ## 0.001 |
| ## 18 | ## 0.006 |
| ## 19 | ## 0.006 |
| ## 20 | ## 0.012 |
| ## 21 | ## 0.011 |
| ## Sample Mean | ## 0.0085 |
| ## Standard deviation | ## 0.0045 |

# Analysis

## Using variable data analysis with a single-sided specification, we assume a normal distribution.

¯x   + ks ≤ maximum specification

## The sample mean of the absorbance is 0.0085, the standard deviation is 0.0047, and the minimum specification is 0.02. Given a sample size of 21, the tolerance interval factor ‘k’ of 2.165 demonstrates that there is 90% confidence that 95% of the population meets specification.

# Conclusion

## Based on the results of the rinse testing, the coating material from the Silq Surface Treatment process remains on the Foley catheter samples meeting the acceptance criteria.