Lubricity Durability Verification,

Silq Surface Treatment

Test Report

# Purpose

## The purpose of this report is to summarize the results of the comparative study used to demonstrate the enhanced lubricity of flat silicone sheets treated with the Silq Surface Treatment relative to non-treated silicon sheets and the durability of the Silq Surface Treatment.

# Scope

## This comparative study applies to Silq’s Surface Treatment process.

## Testing was conducted to determine the drag force of a flat plastic sheet sliding over silicone substrate in accordance with ASTM D1894-14.

# Documents

## Reference Documents

### ASTM D1894-14, Standard Test Method for Static and Kinetic Coefficients of Friction of Plastic Film and Sheeting

### VV.SLQ004 A Test Protocol Lubricity Durability Study, Silq Surface Treatment

# Definitions

## Abbreviations / Acronyms

### ASTM: American Society for Testing and Materials

# Background

## The coefficient of friction is determined by dividing the drag force required to pull one material across another. In this measurement, the two materials are analogous samples of silicone: Silq surface treated silicone or bare silicone. Bare silicone is known to have an especially high coefficient of friction due to its inherent tackiness. This report is meant to demonstrate the drastically reduced coefficient of friction of silicone as a result of applying the Silq Surface Treatment. The results of this study also demonstrate the durability of the Silq Surface Treatment by showing both the static and kinetic coefficient of frictions remain constant after 100 cycles of abrasion. For this study, a 216 gram sled (4.5 in2) was pulled across the full length of a 10” by 5” substrate at a rate of 5.9 in/min.

## The static coefficient of friction was calculated from the following equation:

## µs = As/B

## where, µs is the static coefficient of friction,

## As is the average static drag force, and

## B is the sled weight.

## The kinetic coefficient of friction was calculated from the following equation:

## µk = Ak/B

## where, µk  is the kinetic coefficient of friction,

## Ak is the average kinetic drag force, and

## B is the sled weight.

## Both the drag forces and sled weight are measured in grams.

# Test Method

## The test method was adopted from ASTM D1894-14 and followed directly from the test method detailed in document: VV.SLQ004 A Test Protocol, Lubricity Durability Study, Silq Surface Treatment

# Results

## Five samples of silicone sheets coated with the Silq Surface Treatment underwent the drag force test method described above for 5 cycles. For each test cycle, the static drag force and kinetic drag force were measured. The average static drag force and average kinetic drag forces were then used to calculate the coefficient of friction for each test sample. The results are shown in Table 1 below.

## Table 1: Results of Silq Surface Treated Silicone Samples (n = 5)

|  | ## Sample 1 |  | ## Sample 2 |  | ## Sample 3 |  | ## Sample 4 |  | ## Sample 5 |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| ## Test Cycle | ## Static Drag Force (g) | ## Kinetic Drag Force (g) | ## Static Drag Force (g) | ## Kinetic Drag Force (g) | ## Static Drag Force (g) | ## Kinetic Drag Force (g) | ## Static Drag Force (g) | ## Kinetic Drag Force (g) | ## Static Drag Force (g) | ## Kinetic Drag Force (g) |
| ## 1 | ## 109 | ## 86 | ## 103 | ## 81 | ## 101 | ## 78 | ## 98 | ## 71 | ## 100 | ## 73 |
| ## 2 | ## 104 | ## 84 | ## 104 | ## 84 | ## 108 | ## 86 | ## 96 | ## 69 | ## 94 | ## 69 |
| ## 3 | ## 99 | ## 80 | ## 98 | ## 80 | ## 112 | ## 88 | ## 101 | ## 73 | ## 99 | ## 73 |
| ## 4 | ## 102 | ## 81 | ## 104 | ## 84 | ## 110 | ## 91 | ## 97 | ## 68 | ## 95 | ## 68 |
| ## 5 | ## 98 | ## 82 | ## 108 | ## 86 | ## 108 | ## 84 | ## 99 | ## 71 | ## 98 | ## 67 |
| ## Average Drag Force (g) | ## 102.4 | ## 82.6 | ## 103.4 | ## 83 | ## 107.8 | ## 85.4 | ## 98.2 | ## 70.4 | ## 97.2 | ## 70 |
| ## Average Coefficient of Friction | ## 0.474 (µs) | ## 0.382 (µk) | ## 0.479 (µs) | ## 0.384 (µk) | ## 0.499 (µs) | ## 0.395 (µk) | ## 0.455 (µs) | ## 0.326 (µk) | ## 0.450 (µs) | ## 0.324 (µk) |
| ## Standard Deviation | ## 0.020 | ## 0.011 | ## 0.017 | ## 0.011 | ## 0.019 | ## 0.023 | ## 0.009 | ## 0.009 | ## 0.012 | ## 0.013 |

## One test sample of bare silicone sheet was subjected to the same test. Due to the especially high coefficient of friction and tackiness of bare silicone, accurate and precise readings were difficult to obtain due to ‘chattering’ of the sled across the substrate. The motion of the sled was not constant and would consistently lift off the surface of the substrate resulting in highly varied readings on the force gauge. This is evidenced by the kinetic drag force being nearly equivalent or greater than the static drag force. For this reason, only one specimen of bare silicone was measured. The results are shown in Table 2 below.

## Table 2. Results of Bare Silicone Sample (n = 1)

|  | ## Sample 1 (Bare silicone) |  |
| --- | --- | --- |
| ## Test Cycle | ## Static Drag Force (g) | ## Kinetic Drag Force (g) |
| ## 1 | ## 680 | ## 645 |
| ## 2 | ## 786 | ## 792 |
| ## 3 | ## 724 | ## 708 |
| ## 4 | ## 845 | ## 812 |
| ## 5 | ## 910 | ## 900 |
| ## Average Drag Force (g) | ## 789 | ## 771.4 |
| ## Average Coefficient of Friction | ## 3.65 (µs) | ## 3.57 (µk) |
| ## Standard Deviation | ## 0.426 | ## 0.455 |

## The static and kinetic coefficients of friction for Silq treated silicone samples and the bare silicone sample were averaged and are summarized in Table 3. The average coefficient of frictions for both static and kinetic drag forces are lower in Silq treated samples.

## Table 3: Coefficient of Friction (Static and Kinetic drag forces) Comparison

|  | ## Silq Surface Treated Silicone | ## Bare Silicone |
| --- | --- | --- |
| ## Average Static Coefficient of Friction | ## 0.471 (µs) (SD 0.015) | ## 3.65 (µs) (SD 0.426) |
| ## Average Kinetic Coefficient of Friction | ## 0.362 (µk) (SD 0.013) | ## 3.57 (µk) (SD 0.455) |

## To evaluate the durability of the Silq Surface Treatment, one treated sample was subjected to 100 additional test cycles where the static drag force and kinetic drag force were measured each test cycle. The results of each test cycle can be found in Attachment 1. Table 4 summarizes the average static and kinetic drag forces including the average coefficient of frictions of the first 10 cycles, last 10 cycles, and of all 100 cycles. and standard deviation of the 100 cycles. A graph of the static drag force and kinetic drag force is presented in Figure 1 below.

## Table 4. Silq Surface Treatment Durability

| ## Test Cycle | ## Static Drag Force (g) | ## Kinetic Drag Force (g) |
| --- | --- | --- |
| ## Average Drag Force (First 10 cycles) | ## 103.1 | ## 81.7 |
| ## Average Coefficient of Friction (First 10 cycles) | ## 0.477 (µs) (SD 0.013) | ## 0.388 (µk) (SD 0.011) |
| ## Average Drag Force (Last 10 cycles) | ## 100.3 | ## 80.1 |
| ## Average Coefficient of Friction (Last 10 cycles) | ## 0.464 (µs) (SD 0.007) | ## 0.371 (µk) (SD 0.007) |
| ## Average Drag Force (100 cycles) | ## 100.5 | ## 81.7 |
| ## Average Coefficient of Friction (100 cycles) | ## 0.465 (µs) (SD 0.012) | ## 0.378 (µk) (SD 0.013) |

## Figure 1. Static and Kinetic Drag Force over 100 Cycles

# Analysis

## Comparing bare silicone sheet to Silq surface treated silicone sheet

### The static and kinetic coefficient of friction of each catheter type was compared against each other using a two sample t-test (assuming equal variance).

### For test results which were found to be "statistically significant", the data had to show that the coefficient of friction in the test sample (Silq treated silicone sample) was greater than the coefficient of friction observed in the control sample (bare silicone sample).  For results significant at P < 0.05, there was a 95% certainty that there was a difference between test sample and control sample.  For those results significant at P < 0.01, there was a 99% certainty that there was a difference between test catheter and control catheter.

### Comparison of the static coefficient of frictions between the Silq treated silicone samples (test) and the bare silicone sample (control)

#### The results of the two sample t-test calculated a P value of 3.06 x 10-26, demonstrating a statistically significant difference (with 99% certainty) in the static coefficient of friction of the bare silicone sample compared to the Silq treated silicone samples.

#### The results verify the capability of the Silq treated silicone sample to reduce statistically significant less static drag force than the bare silicone sample.

### Comparison of the kinetic coefficient of frictions between the Silq treated silicone sample (test) and the bare silicone sample (control)

#### The results of the two sample t-test calculated a P value of 1.79 x 10-25, demonstrating a statistically significant difference (with 99% certainty) in the kinetic coefficient of friction of the bare silicone sample compared to the Silq treated silicone samples.

#### The results verify the capability of the Silq treated silicone sample to reduce statistically significant less kinetic drag force than the bare silicone sample.

### Comparing the average static and kinetic coefficients of friction of the Silq treated silicone samples to the bare silicone sample, the samples treated with Silq’s Surface Treatment demonstrate a significantly more lubricious surface as shown in Table 3. The average static coefficient of friction was 7.7x less in the Silq treated silicone sample compared to the bare silicone sample, 0.471 and 3.65, respectively. The average kinetic coefficient of friction was 9.9x less in the Silq treated silicone sample compared to the bare silicone sample, 0.362 and 3.57, respectively.

## The durability of the Silq Surface Treatment was demonstrated by measuring the static and drag forces of a weighted sled for 100 runs. The static and drag forces of the Silq treated silicone sample stayed equivalent over 100 cycles of abrasion as shown in Figure 1. Performing a simple linear regression on the static and kinetic forces show that the lubricity of the Silq treated sample improves slightly over time as seen with a negative slope of -0.0164 and -0.0475, respectively. After 100 cycles of traversing a sled across its surface, it was shown that the coating did not degrade.

# Conclusions

## The results of this study demonstrate the enhanced lubricity imparted on silicone substrates as a result of the Silq Surface Treatment.

## The results of this study demonstrate that the Silq Surface Treatment is durable with the lubricity of the Silq treated sample remaining equivalent over time. The durability of the Silq surface treatment is demonstrated through maintaining equivalent static and kinetic drag forces over 100 cycles of testing a weighted sled across the treated silicone substrate.

# Attachments

## Attachment 1 Test Report, Lubricity Durability Study, Silq Surface Treatment.xlsx