Saline Soak Stability Experiment

Silq Surface Treatment

Test Report

# Purpose

## The purpose of this report is to summarize the testing performed to verify the presence of Silq’s Surface Treatment coating using a crystal violet stain after a simulated use condition.

## Sterilized samples were subjected to a 30 day period of immersion in a 0.9% Saline solution at 37oC to simulate physiological conditions of implanted devices.

# Scope

## This test report applies to the Silq Surface Treatment process.

## The Silq Surface Treatment process may be applied to a variety of polymeric substrates. Testing in this procedure was conducted on 16 Fr Foley catheters that have been coated using the Silq Surface Treatment process.

# Definitions

## Abbreviations / Acronyms

### mL: Milliliter

### nm: Nanometer

# Background

## Crystal violet is a triarylmethane dye commonly used as a histological stain to determine the presence of proteins and DNA in biological research. It has been found to have a strong interaction with the polymers present in the Silq Surface Treatment. When surface treated materials are immersed in a crystal violet solution, the dye intercalates into the polymer network present on the surface. Subsequently, the stained substrate is immersed in an acetic acid solution which fully dissolves the crystal violet molecules allowing them to diffuse out of the network and into the solution. Through measurement of the absorbance of the resulting solution at 585nm, a relative measure of the amount of zwitterionic polymer present on the surface of the substrate can be obtained.

# Test Method

## The test method for this report was adopted from the protocol detailed in document: VV.SLQ008 A Test Protocol, Saline Soak Stability Experiment, Silq Surface Treatment.

## Test samples and sample size

### A control and a test group were evaluated for the presence of coating using crystal violet stain.

### Twenty-one (21) 16 Fr Foley catheters were treated with the Silq Surface Treatment process in accordance with established procedures and underwent the validated ethylene oxide sterilization process. Afterwards, two segments from each catheter were removed for the control group and test group.

### A sample size of twenty-one was chosen to demonstrate 90%/90% confidence/reliability using variable data analysis.

## Acceptance Criteria

### The absorbance value shall be a minimum of 0.400.

### Rationale: This acceptance criteria threshold was derived from previous testing where the amount of coating resulting in an absorbance of 0.300 was sufficient to produce the desired effect of reduced microbial adhesion on device surfaces. Testing also demonstrated that higher levels of coating will result in the same or better efficacy in reducing bacterial adhesion. Therefore, an additional safety factor of 33% was added to the absorbance of 0.300.

# Deviations

## Test samples were immersed for a period of 30 days instead of the 28 days as specified in the test protocol. Immersion in the saline solution for an additional two days challenges the ability of the Silq Surface Treatment coating to adhere to the catheter for a prolonged duration. This deviation has no impact on the validity of the test results.

# Results

## Twenty-one (21) Foley catheters were treated with the Silq Surface Treatment and sterilized using ethylene oxide sterilization. Samples in the control group were stained with crystal violet dye and absorbance was measured. The results of the control group are shown in Table 6.1. The test group was soaked in a 0.9% Saline solution at an average temperature of 37.4oC for 30 days. Subsequently, samples from the test group were stained with crystal violet dye and absorbance was measured. The results of the test group are shown in Table 6.2.

## Table 6.1 – Control group absorbance results

| ## Catheter Sample # | ## Absorbance |
| --- | --- |
| ## 1 | ## 0.55 |
| ## 2 | ## 0.597 |
| ## 3 | ## 0.569 |
| ## 4 | ## 0.445 |
| ## 5 | ## 0.441 |
| ## 6 | ## 0.581 |
| ## 7 | ## 0.616 |
| ## 8 | ## 0.597 |
| ## 9 | ## 0.56 |
| ## 10 | ## 0.588 |
| ## 11 | ## 0.565 |
| ## 12 | ## 0.532 |
| ## 13 | ## 0.525 |
| ## 14 | ## 0.585 |
| ## 15 | ## 0.496 |
| ## 16 | ## 0.565 |
| ## 17 | ## 0.509 |
| ## 18 | ## 0.544 |
| ## 19 | ## 0.541 |
| ## 20 | ## 0.585 |
| ## 21 | ## 0.51 |
| ## Sample Mean | ## 0.548 |
| ## Standard deviation | ## 0.047 |

## Table 6.2 – Test group absorbance results

| ## Catheter Sample # | ## Absorbance |
| --- | --- |
| ## 1 | ## 0.552 |
| ## 2 | ## 0.476 |
| ## 3 | ## 0.435 |
| ## 4 | ## 0.475 |
| ## 5 | ## 0.512 |
| ## 6 | ## 0.495 |
| ## 7 | ## 0.5 |
| ## 8 | ## 0.445 |
| ## 9 | ## 0.536 |
| ## 10 | ## 0.61 |
| ## 11 | ## 0.441 |
| ## 12 | ## 0.516 |
| ## 13 | ## 0.496 |
| ## 14 | ## 0.565 |
| ## 15 | ## 0.496 |
| ## 16 | ## 0.565 |
| ## 17 | ## 0.436 |
| ## 18 | ## 0.456 |
| ## 19 | ## 0.495 |
| ## 20 | ## 0.58 |
| ## 21 | ## 0.465 |
| ## Sample Mean | ## 0.502 |
| ## Standard deviation | ## 0.050 |

# Analysis

## Using variable data analysis with a single-sided specification, we assume a normal distribution.

¯x   - ks > minimum specification

### For the test group, the sample mean of the absorbance is 0.502, the standard deviation is 0.0502, and the minimum specification is 0.400. Given a sample size of 21, the tolerance interval factor ‘k’ of 2.994 demonstrates that there is 90% confidence that 99% of the population meets specification.

# Conclusion

## Based on the results of the crystal violet absorbance test, the coating remained present on all catheter samples after immersion in a 0.9% Saline solution at 37oC for 30 days.

# Attachments

## Attachment 1 – Absorbance Results, Saline Soak Stability Experiment.xlsx