Suspension

Master Validation Plan

# Purpose

## The purpose of this document is to define the activities required to validate the applicable manufacturing processes of the Silq Suspension. Where the output of a process is not verifiable, process validation is designed to establish by objective evidence that a process consistently produces a result or product that meets its predetermined specifications.

## This Master Validation Plan (MVP) identifies the proposed process validation activities that are required for the manufacturing of the Suspension.

# Scope

## This validation plan covers the manufacturing processes for Silq’s Suspension at 311 Otterson Drive facility in Chico, CA.  The individual manufacturing processes are described throughout this document

## The sterilization process and associated sterile barrier qualification for the Silq finished devices have been validated separately and are not covered in this plan.

## This validation is being performed to fulfil the requirements of the US quality system regulations, 21 CFR 820, and ISO 13485.

# Documents

## Reference Documents

### SP-C.SLQ001		Specification, Suspension

### MP-C.SLQ001		Manufacturing Procedure, Suspension

# Definitions

## Abbreviations / Acronyms

### V&V:  Verification and Validation

## General Definitions

### Installation Qualification (IQ):  a validation document that verifies that the system or equipment under test is installed per manufacturer’s and company’s specifications.

### Operational Qualification (OQ): a validation document that demonstrates that process operating at the process control limits within its intended ranges produces a product that meets product of acceptable quality.

### Process Qualification (PQ): a validation document that demonstrates that the process will consistently produce acceptable results under normal operating conditions.

### Product Performance Qualification (PPQ): a validation document that verifies that product built with the established process meets product specifications.

### Test Method Validation (TMV): a process of establishing, by objective evidence, that a test method consistently generates the appropriate result needed to meet the intended usage

# Responsibilities

## It is the responsibility of the R&D/Engineering, Quality Assurance, and Manufacturing personnel to implement this Master Validation Plan and to ensure the MVP status is updated as testing is completed.

# Master Validation Plan

## The manufacture of the Silq Suspension is performed by Silq at its Chico manufacturing facility as outlined in the Suspension Bill of Reference (Document #BOR-C.SLQ001). Silq performs all manufacturing activities of the Suspension including processing the unpurified polymeric dispersion solution, performing final inspection testing, labeling, and packaging.

## A review of the processes identified the following equipment is involved:

### Diafiltration System

### Floor Scale (Concentrate Reactor)

### Floor Scale (Suspension)

### Pneumatic Drum Stirrer

### Analytical Balance

## Installation Qualification (IQ)

### The IQ shall consist of the following:

#### Make a hardware inventory including identifying equipment manufacturer, model, location, etc.

#### Record specifications of installation and safety parameters of the equipment (e.g., voltage range, current rating, emergency stop).

#### Verify that the installation and safety parameters are met.

#### Identify software and/or operating system, if applicable.

#### Calibration

##### Identify instruments needing calibration.

##### Verify calibration is defined and current or, if needed, add instruments to the calibration program.

#### Maintenance of the equipment

##### Verify user’s manual from the manufacturer of the equipment, if applicable.

##### Obtain spare parts list, if applicable.

##### Verify maintenance is defined and current or, if needed, add instrument to the calibration program.

#### Training and personnel qualification

##### Identify training needs.

##### Verify or perform the necessary training.

Table 1 - Installation Qualification Requirement

| Process | Equipment Description | Equipment ID | Installation Qualification Required | Qualification Protocol |
| --- | --- | --- | --- | --- |
| Diafiltration | Diafiltration system (Repligen KTF-2000) | ST-005 | Yes | VV.SLQ011 |
|  | Peristaltic Pump, large flow | ST-008 | Yes | VV.SLQ012 |
|  | Pneumatic Drum Stirrer | ST-009 | Yes | VV.SLQ013 |
|  | Waste Diaphragm Pump | ST-004 | Yes | VV.SLQ014 |
|  | Floor Scale (Concentrate Reactor) | ST-006 | Yes | VV.SLQ015 |
| Packaging | Floor Scale (Suspension) | ST-007 | Yes | VV.SLQ016 |
| QC Analysis | Analytical Balance | ST-012 | Yes | VV.SLQ017 |

## Operation Qualification (OQ)

### OQ is not required for the Suspension manufacturing process. There are no operating ranges (process control limits) that are controlled during the manufacturing process.

## Process Qualification (PQ)

### PQ is not required for the Suspension manufacturing process. There are no operating ranges (process control limits) that are controlled during the manufacturing process. In lieu of a PQ, a product performance qualification (PPQ) will be performed to ensure product built with the established process and equipment meets specification.

## Product Performance Qualification (PPQ)

### The PPQ shall consist of producing four (4) consecutive batches of Suspension that meet product specifications outlined in SP-C.SLQ001.

Table 2 – Product Performance Qualification Requirement

| Manufacturing procedure | Equipment | PPQ Sampling Plan | Acceptance Criteria* | Qualification Protocol |
| --- | --- | --- | --- | --- |
| MP-C.SLQ001 | - Diafiltration system (Repligen KTF-2000), - Peristaltic Pump, large flow - Pneumatic Drum Stirrer - Waste Diaphragm Pump - Floor Scale (Concentrate Reactor) | Four (4) consecutive lots of Suspension | 1) Appearance (color): Opaque white suspension 2) Appearance (form): Inviscid liquid 3) Residue on Evaporation: 0.5% - 1.5% 4) pH: 5 - 8 | VV.SLQ018 |

### *Acceptance criteria derived from SP-C.SLQ001

## Test Method Validation

### Test method validation of the analytical assay method for final product verification is required to confirm, by objective evidence, that the test method consistently generates the appropriate result needed to meet the intended usage.

## Software Considerations

### The Repligen KrosFlo KTF-2000 Diafiltration System comes with custom software that operates the diafiltration system (e.g., controls, monitors). The custom software and electromechanical/fluidic system will be validated as a whole and not separately.

## Criteria for revalidation

### The following items will be reviewed and requirements for revalidation will be identified:

### Key process parameters

### Equipment changes

### Product changes

### New product/batch sizes

### Move of equipment

### Major repair-maintenance, including software upgrades

### New environmental conditions

### Change of requirements (standards, legal requirements, product)

## Validation testing is to be documented using a series of test protocols and summary reports, when applicable.

## Completed test protocol and summary reports are to be retained in accordance with Document Control procedure (QM.SLQ001).

# Output documents to be delivered

## IQ

### IQ protocol(s)

### IQ report(s)

## PPQ

### PPQ protocol

### PPQ report

## TMV

### TMV protocol

### TMV report

## Final validation summary report

### Summary of activities (results of IQ, PPQ, and TMV) and conclusions of validation including any follow-up actions