Installation Qualification (IQ) Protocol, Diafiltration System

Process Validation Test Protocol

# Purpose

## The purpose of this protocol is to define the Installation Qualification activities required to validate the KrosFlo KTF-2000 Diafiltration System intended for the diafiltration processing of Silq Suspension.

## The purpose of the IQ is to verify the conformity of the installation of the equipment meets manufacturer’s specifications.

# Scope

## This process validation protocol applies to Repligen’s KrosFlo KTF-2000 Diafiltration System located at Silq manufacturing facility in Chico, CA used in the diafiltration processing of Silq Suspension.

# Documents

## Reference Documents

### SP-E.SLQ006		Source Control Specification, Diafiltration System, KrosFlo KTF-2000 System

### MP-C.SLQ001		Manufacturing Procedure, Suspension

### 21 CFR Part 820.75	FDA Quality System Regulation – Process Validation

### ISO 13485:2016	Medical devices – Quality Management Systems – Requirements for Regulatory Purposes (7.5.6 Validation of processes for production and service provision)

# Definitions

## General Definitions

### Installation Qualification (IQ):  a validation document that verifies that the system or equipment under test is installed per manufacturer’s and company’s specifications.

# Test Equipment and Materials / Supplies

## The KTF-2000 diafiltration system is required for the execution of the protocol. Additional equipment and supplies may be used as needed. All equipment utilized for performing the testing must be recorded on the applicable data sheet.

# Test Method

## The installation qualification consists of the steps outlined below to ensure the installation of the equipment meets manufacturer’s requirements.

## The equipment to qualify is a diafiltration system:

| ## Brand | ## Model | ## Location |
| --- | --- | --- |
| ## Repligen | ## KrosFlo KTF-2000 | ## Manufacturing – Controlled Environment Area |

## Description of the diafiltration operation

### The operation consists of the removal or separation of components of the unpurified polymeric dispersion solution with the use of permeable membrane filters. The operation is performed in accordance with manufacturing procedure MP-C.SLQ001 “Manufacturing Procedure: Suspension.” It describes the different steps, from setting up the diafiltration system to performing the diafiltration process.

## Methodology

### An installation qualification must be conducted for each diafiltration system to ensure the manufacturer’s requirements to operate the equipment are met.

## Sample Size

### Production of test samples are not required for the installation qualification of the diafiltration system.

## Acceptance Criteria

### See Section 9 below.

## Management of non-conformities / deviations

### Any non-conformity / deviation must be documented. To accept the non-conformity / deviation, a written justification must be provided documenting why the anomaly does not affect the outcome of the test.

## Steps in the Qualification Process

### Installation Qualification

| ### The tests to be conducted in IQ | ### Test Sheet |
| --- | --- |
| ### Verification of the installation | ### Attachment 1 |
| ### Verification of the documentation | ### Attachment 2 |
| ### Verification of utilities | ### Attachment 3 |

# Installation Qualification

## Verification of the installation

### Objective

#### The purpose of this test is to verify the compliance of the installation with the documentation.

### Operating mode

#### Check the references of the equipment.

#### Check the conditions for proper use and installation.

### Acceptance criteria

#### The equipment is installed according to the conditions recommended in the manufacturer’s documentation.

### Test sheet

#### See protocol Attachment 1 for data sheet

## Verification of the documentation

### Objective

#### The purpose of this test is to verify the presence of the equipment documentation.

### Operating mode

#### Check the presence of the equipment user’s manual.

#### Check for the presence of a valid calibration and preventive maintenance of the equipment.

#### Check that the frequency of calibration and preventive maintenance has been defined.

### Acceptance criteria

#### The equipment user manual is present.

#### The equipment calibration and preventive maintenance report is present.

#### The frequency of calibration and preventive maintenance is documented in the equipment requirements form.

### Test sheet

#### See protocol Attachment 2 for data sheet.

## Verification of utilities

### Objective

#### The purpose of this test is to verify that the equipment is correctly installed.

### Operating mode

#### Identify the characteristics of the power supply required to operate the equipment.

#### Check the absence of damage to the power supply cable.

#### Verify operating temperature condition is within range.

#### Verify the equipment is operational when powered on.

### Acceptance criteria

#### The equipment is properly connected to mains power.

#### No damage to power supply cable.

#### Operating temperature: 0 – 50℃

#### Equipment can be powered on and is operational.

### Test sheet

#### See protocol Attachment 3 for data sheet.

# Analysis of Data

## Attribute data analysis techniques will be used to establish pass/fail results.

# Acceptance Criteria

## Installation qualification

### The diafiltration system must be installed according to its specifications. All technical documents must be included in the equipment file.

# Attachments

## Attachments 1 – 3: Data Sheets of Data Collection Forms

## Attachment 4: Calibration Log

## Attachment 5: Signature Log

## PROTOCOL ATTACHMENT 1 – IQ: Verification of the installation

| ### Verification | ### Acceptance criteria | ### Results observed | ### Conformity C / NC |
| --- | --- | --- | --- |
| ### Brand | ### Repligen |  |  |
| ### Model number | ### KrosFlo KTF-2000 |  |  |

| ### Comments: |  |
| --- | --- |

Completed By (Sign and Date):________________________________

Reviewed By (Sign and Date):________________________________

## PROTOCOL ATTACHMENT 2 – IQ: Verification of the documentation

| ### Documentation | ### Acceptance criteria | ### Results observed | ### Conformity C / NC |
| --- | --- | --- | --- |
| ### Equipment user manual | ### Presence of the user manual in English |  |  |
| ### Equipment calibration certificate | ### Presence of equipment calibration certificate |  |  |
| ### Calibration requirements | ### Schedule / frequency of calibration determined |  |  |
| ### Preventive maintenance requirements | ### Schedule / frequency of PM determined |  |  |

| ### Comments: |  |
| --- | --- |

Completed By (Sign and Date):________________________________

Reviewed By (Sign and Date):________________________________

## PROTOCOL ATTACHMENT 3 – IQ: Verification of utilities

| ### Verification | ### Acceptance criteria | ### Results observed | ### Conformity C / NC |
| --- | --- | --- | --- |
| ### Power supply | ### Compliant if the system is powered by 200 – 240 VAC 1ph 15A |  |  |
| ### Wiring | ### No damage to power supply cable |  |  |
| ### Operating temperature | ### 0 – 50℃ |  |  |
| ### Operation | ### Equipment can be powered on and is operational |  |  |

| ### Comments: |  |
| --- | --- |

Completed By (Sign and Date):________________________________

Reviewed By (Sign and Date):________________________________

## PROTOCOL ATTACHMENT 4 – Calibration Log

In the table below, record the calibration information for all equipment used during the test. If the equipment is calibrate before use then include or refer to the procedure used for calibration, record any equipment required and its calibration info, and record the results of the calibration (data).

| Instrument ID | Description | Use | Calibration Date | Due Date | Verified By/Date |
| --- | --- | --- | --- | --- | --- |

## PROTOCOL ATTACHMENT 5 – Signature Log

In the table below, all personnel who have completed any section of this protocol must include the listed information.

| Printed Name | Title | Initials | Signature | Company | Date |
| --- | --- | --- | --- | --- | --- |