Installation Qualification,

Floor Scale, Suspension, ST-007

Test Report

# Executive Summary

| ## Status: | ## PASS |
| --- | --- |

## Testing was conducted to ensure the Global Industrial 318506 floor scale was installed appropriately in accordance with manufacturer’s specification. This report documents the results of the equipment installation qualification for the suspension floor scale.

## Installation qualification performed met the respective acceptance criteria for the operation of the suspension floor scale.

# Deviations

## No deviations occurred during the execution of the test protocol (Document #VV.016).

# Results

## Installation qualification on the Global Industrial 318506 floor scale (Equipment ID #ST-007, S/N 318506-2108) was successfully completed in accordance with the protocol VV.SLQ016.

## The required documentation was found in the equipment file:

### User manual in English (See Attachment 2)

### Valid calibration certificate (See Attachment 3)

### Calibration schedule (See Attachment 4)

## The required utilities were verified to meet acceptance criteria:

### Floor scale is powered by AC adapter cable which meets the required power input specification.

### No damage to AC adapter cable was observed.

### The temperature and humidity of the work environment were measured to be 18.5℃ and 40% RH, respectively, and were within the acceptable operating temperature and humidity ranges specified by manufacturer of -10 – 40℃ and <90% RH, respectively.

## The 318506 floor scale was powered on and verified to be operational.

# Conclusions

## All required documentation and utilities for the 318506 floor scale have been verified to meet manufacturer’s specification, and the floor scale has been verified to be operational.

## The 318506 floor scale (Equipment ID # ST-007) has been installed properly in accordance with manufacturer’s specification.

# Attachments

## Attachment 1: Installation Qualification datasheet

## Attachment 2: Users Manual

## Attachment 3: Calibration Certificate

## Attachment 4: Calibration Requirements