Installation Qualification,

Weighing Scale, ST-012

Test Report

# Executive Summary

| ## Status: | ## PASS |
| --- | --- |

## Testing was conducted to ensure the Mettler Toledo AB54-S weighing scale was installed appropriately in accordance with manufacturer’s specification. This report documents the results of the equipment installation qualification for the weighing scale.

## Installation qualification performed met the respective acceptance criteria for the operation of the weighing scale.

# Deviations

## No deviations occurred during the execution of the test protocol (Document #VV.017).

# Results

## Installation qualification on the Mettler Toledo AB54-S weighing scale (Equipment ID #ST-012, S/N 1125311515) was successfully completed in accordance with the protocol VV.SLQ017.

## The required documentation was found in the equipment file:

### User manual in English (See Attachment 2)

### Valid calibration certificate (See Attachment 3)

### Calibration schedule (See Attachment 4)

## The required utilities were verified to meet acceptance criteria:

### Weighing scale is powered by AC adapter cable which meets the required power input specification.

### No damage to AC adapter cable was observed.

### The temperature and humidity of the work environment were measured to be 22.9℃ and 50% RH, respectively, and were within the acceptable operating temperature and humidity ranges specified by manufacturer of -10 – 30℃ and 15-80% RH, respectively.

## The AB54-S weighing scale was powered on and verified to be operational.

# Conclusions

## All required documentation and utilities for the AB54-S weighing scale have been verified to meet manufacturer’s specification, and the weighing scale has been verified to be operational.

## The AB54-S weighing scale (Equipment ID # ST-012) has been installed properly in accordance with manufacturer’s specification.

# Attachments

## Attachment 1: Installation Qualification datasheet

## Attachment 2: Users Manual

## Attachment 3: Calibration Certificate

## Attachment 4: Calibration Requirements