Installation Qualification (IQ) Protocol, UV/ Visible Spectrophotometer

Process Validation Test Protocol

# Purpose

## The purpose of this protocol is to define the Installation Qualification activities required to validate the UV/Visible Spectrophotometer used for the Crystal Violet Absorbance test and the Coating Adherence Verification test for Silq Suspension.

## The purpose of the IQ is to verify the conformity of the installation of the equipment meets manufacturer’s specifications.

# Scope

## This process validation protocol applies to the UV/Visible Spectrophotometer located at Silq manufacturing facility in Los Angeles, CA used in the quality control analysis of Silq Suspension.

# Documents

## Reference Documents

### SP-E.SLQ015	Specification, Spectrophotometer, Jenway 7205 UV-Visible Diode Array

### SP-TF.SLQ001	Specification, VIS Photometric Accuracy Calibration Standard, 400-700nm, HF-0.2

### 21 CFR Part 820.75 FDA Quality System Regulation – Process Validation

### ISO 13485:2016 Medical devices – Quality Management Systems – Requirements for Regulatory Purposes (7.5.6 Validation of processes for production and service provision)

# Definitions

## General Definitions / Acronyms

### Installation Qualification (IQ):  a validation document that verifies that the system or equipment under test is installed per manufacturer’s and company’s specifications.

### Ultraviolet (UV)

# Test Equipment and Materials / Supplies

## The UV/Visible Spectrophotometer is required for the execution of this protocol. Additional equipment and supplies may be used as needed. All equipment utilized for performing the testing must be recorded on the applicable data sheet.

# Test Method

## The installation qualification consists of the steps outlined below to ensure the installation of the equipment meets manufacturer’s requirements.

## The equipment to qualify is a UV/Visible Spectrophotometer:

| ## Brand | ## Model | ## Location |
| --- | --- | --- |
| ## Jenway | ## 7205 | ## Silq Quality Control Lab (Los Angeles, CA) |

## Description of operation

### The UV/Visible Spectrophotometer is used to determine the absorbance value at specified wavelengths of test solutions. The spectrophotometer is used to verify attributes of the Silq Suspension in quality control testing.

## Methodology

### An installation qualification must be conducted for the spectrophotometer to ensure the manufacturer’s requirements to operate the equipment are met.

## Sample Size

### Production of test samples is not required for the installation qualification of the spectrophotometer.

## Acceptance Criteria

### See Section 9 below.

## Management of non-conformities / deviations

### Any non-conformity / deviation must be documented. To accept the non-conformity / deviation, a written justification must be provided documenting why the anomaly does not affect the outcome of the test.

## Steps in the Qualification Process

### Installation Qualification

| ### The tests to be conducted in IQ | ### Test Sheet |
| --- | --- |
| ### Verification of the installation | ### Attachment 1 |
| ### Verification of the documentation | ### Attachment 2 |
| ### Verification of utilities | ### Attachment 3 |

# Installation Qualification

## Verification of the installation

### Objective

#### The purpose of this test is to verify the compliance of the installation with the documentation.

### Operating mode

#### Check the references of the equipment.

#### Check the conditions for proper use and installation.

### Acceptance criteria

#### The equipment is installed according to the conditions recommended in the manufacturer’s documentation.

### Test sheet

#### See protocol Attachment 1 for data sheet

## Verification of the documentation

### Objective

#### The purpose of this test is to verify the presence of the equipment documentation.

### Operating mode

#### Check the presence of the equipment user’s manual.

### Acceptance criteria

#### The equipment user manual is present.

### Test sheet

#### See protocol Attachment 2 for data sheet.

## Verification of utilities

### Objective

#### The purpose of this test is to verify that the equipment is correctly installed.

### Operating mode

#### Identify the characteristics of the power supply required to operate the equipment.

#### Check the absence of damage to the power supply cable.

#### Verify operating temperature and relative humidity conditions are within range.

#### Verify the equipment is operational when powered on.

#### Utilizing the VIS Photometric Accuracy Calibration Standard (TF.SLQ001), verify the absorbance values of 0.405 at 400nm, 0.240 at 440nm, 0.250 at 585nm, and 0.220 at 700nm (+/- 0.01).

### Acceptance criteria

#### The equipment is properly connected to mains power (100 – 240 VAC at 60-50Hz) via an AC adapter cable.

#### No damage to AC adapter cable.

#### Operating temperature: 5 – 40℃; Operating relative humidity: <80% for temperatures up to 31℃

#### Equipment can be powered on and is operational.

#### Absorbance values: 0.405 +/- 0.01 (at 400nm); 0.240 +/- 0.01 (at 440nm); 0.250 +/- 0.01 (at 585nm); and 0.220 +/- 0.01 (at 700nm).

### Test sheet

#### See protocol Attachment 3 for data sheet.

# Analysis of Data

## Attribute data analysis techniques will be used to establish pass/fail results.

# Acceptance Criteria

## Installation qualification

### The Jenway 7205 Spectrophotometer must be installed according to its specifications. Acceptance criteria is defined in Sections 7.1.3, 7.2.3, and 7.3.3.

# Attachments

## Attachments 1 – 3: Data Sheets of Data Collection Forms

## Attachment 4: Calibration Log

## Attachment 5: Signature Log

## PROTOCOL ATTACHMENT 1 – IQ: Verification of the installation

| ### Verification | ### Acceptance criteria | ### Results observed | ### Conformity C / NC |
| --- | --- | --- | --- |
| ### Brand | ### Jenway |  |  |
| ### Model number | ### 7205 |  |  |

| ### Comments: |  |
| --- | --- |

Completed By (Sign and Date):________________________________

Reviewed By (Sign and Date):________________________________

## PROTOCOL ATTACHMENT 2 – IQ: Verification of the documentation

| ### Documentation | ### Acceptance criteria | ### Results observed | ### Conformity C / NC |
| --- | --- | --- | --- |
| ### Equipment user manual | ### Presence of the user manual in English |  |  |

| ### Comments: |  |
| --- | --- |

Completed By (Sign and Date):________________________________

Reviewed By (Sign and Date):________________________________

## PROTOCOL ATTACHMENT 3 – IQ: Verification of utilities

| ### Verification | ### Acceptance criteria | ### Results observed | ### Conformity C / NC |
| --- | --- | --- | --- |
| ### Power supply | ### Compliant if the spectrophotometer is powered by AC adapter cable |  |  |
| ### Wiring | ### No damage to AC adapter cable |  |  |
| ### Operating temperature | ### 5 – 40℃ |  |  |
| ### Operating relative humidity | ### <80% RH for temperatures up to 31℃ |  |  |
| ### Operation | ### Equipment can be powered on and is operational |  |  |
| ### Performance | ### Absorbance values: |  |  |
|  | ### 0.405 +/- 0.01 (at 400nm) |  |  |
|  | ### 0.240 +/- 0.01 (at 440nm) |  |  |
|  | ### 0.250 +/- 0.01 (at 585nm) |  |  |
|  | ### 0.220 +/- 0.01 (at 700nm) |  |  |

| ### Comments: |  |
| --- | --- |

Completed By (Sign and Date):________________________________

Reviewed By (Sign and Date):________________________________

## PROTOCOL ATTACHMENT 4 – Calibration Log

In the table below, record the calibration information for all equipment used during the test. If the equipment is calibrate before use then include or refer to the procedure used for calibration, record any equipment required and its calibration info, and record the results of the calibration (data).

| Instrument ID | Description | Use | Calibration Date | Due Date | Verified By/Date |
| --- | --- | --- | --- | --- | --- |

## PROTOCOL ATTACHMENT 5 – Signature Log

In the table below, all personnel who have completed any section of this protocol must include the listed information.

| Printed Name | Title | Initials | Signature | Company | Date |
| --- | --- | --- | --- | --- | --- |