Suspension Shelf Life Evaluation,

Silq Surface Treatment

Test Protocol

# Purpose

## The purpose of this protocol is to verify the Silq Suspension, part number C.SLQ001, meets specified performance requirements for a defined shelf life (e.g., 2 years).

## Performance testing consists of:

### Verifying the quality attributes of the Suspension (i.e., appearance, residue on evaporation, pH, 1H NMR, and 19F NMR)

### Confirming the presence and adherence of the Suspension to the catheter

# Scope

## This test protocol applies to the Silq Suspension, part number C.SLQ001

## The Silq Surface Treatment process may be applied to a variety of polymeric substrates. A representative catheter (16 Fr Foley catheter) will be chosen to evaluate the performance of aged Suspension solution.

# Documents

## Reference Documents

### QC-C.SLQ001 Quality Control Procedure, Suspension Analysis

### SP-C.SLQ001 Specification, Suspension

### OM.SLQ003 Quality Control Procedure, Crystal Violet Absorbance

### OM.SLQ004 Quality Control Procedure, Coating Adherence Verification

# Background

## The C.SLQ001 Suspension solution is a water-based coating solution for polymeric and elastomeric surfaces. The solution contains a three-component methacrylate copolymer, polymerized via free radical polymerization from polysulfobetaine methacrylate, perfluorophenylazide methacrylate, and methacrylic acid. The material is dispersed at 1% w/v% in high purity water.

# Training

## Applicable personnel will be trained to this protocol (Training Record Form in Appendix 1).

# Test Configuration

## Aged Suspension will be tested to confirm its attributes meets specification.

## Finished 16 Fr Foley catheters built to the approved manufacturing procedure by Pathway will be used for the evaluation of the aged Suspension on catheters. 16 Fr Foley catheters are representative samples of the ClearTract Foley catheter product family.

## Manufacturing lot numbers will be recorded for traceability.

## All test samples of the finished 16 Fr Foley catheters will be packaged and subjected to the validated ethylene oxide sterilization cycle.

## Sample size

### Shelf life evaluation will consist of the following tests with the listed sample sizes:

| Test Group | Test Description | Sample Size |
| --- | --- | --- |
| Quality attributes (Suspension) | Verify the attributes of the Suspension (i.e., appearance, residue on evaporation, pH, 1H NMR, and 19F NMR) | 100% |
| Crystal Violet Absorbance (Finished catheter) | Verify the presence of Suspension using a crystal violet dye solution | A sufficient number of test samples to achieve 90%/90% confidence/reliability interval using attribute data analysis |
| Coating Adherence Verification (Finished catheter) | Verify the Suspension remains adhered to the catheter | A sufficient number of test samples to achieve 90%/90% confidence/reliability interval using attribute data analysis |

### For quality attribute testing of the Suspension, each lot of Suspension will be tested for a sample size of 100%.

### For crystal absorbance and coating adherence verification testing of finished catheters, a sufficient number of test samples will be selected to perform attribute data analysis at a 90/90% confidence/reliability interval. These statistical intervals were determined using a risk-based approach in compliance with the statistical techniques work instruction QM.SLQ011.

# Test Equipment and Materials / Supplies

## See the equipment, materials, and supplies required for the execution of this test protocol defined in documents QC-C.SLQ001 Quality Control Procedure, Suspension Analysis, OM.SLQ003 Quality Control Procedure, Crystal Violet Absorbance, and OM.SLQ004 Quality Control Procedure, Coating Adherence Verification.

# Test Procedure: Suspension

### Prepare and test the aged C.SLQ001 Suspension in accordance with QC-C.SLQ001 Quality Control Procedure, Suspension Analysis.

### Document the test results on the Suspension Quality Control Report template (TMP1-QC-C.SLQ001).

### Record the summary of the quality control test in Appendix 2. Attach a copy of the Suspension Quality Control Report to the final test report.

# Test Procedure: Finished Catheter

### Using the aged Suspension, Pathway shall manufacture 16 Fr ClearTract Foley catheter test samples in accordance with established manufacturing procedures. Following sterilization, the test samples will be delivered to Silq for testing.

### The presence of Suspension on the test samples shall be verified in accordance with OM.SLQ003 Quality Control Procedure, Crystal Violet Absorbance. Record the results of the crystal violet absorbance test in Appendix 2.

### Adherence of the Suspension on the test samples shall be verified in accordance with OM.SLQ004 Quality Control Procedure, Coating Adherence Verification. Record the results of the coating adherence verification test in Appendix 2.

# Analysis of Data

## Testing will collect attribute data.  The sample size corresponds to a 90% confidence / 90% reliability interval.

## Results will be compared to acceptance criteria and determined to pass or fail.  No additional data analysis is required.

# Acceptance Criteria

## Suspension

### Appearance – Color: Opaque white

### Appearance – Form: Inviscid liquid

### Residue on Evaporation: 0.5% – 1.5%

### pH: 5 – 8

### 1H NMR: A successful result is one that includes broad peaks within 0.2ppm of 0.99, 2.2, 2.9, 3.1, 3.5, 3.7, 4.4, and 4.75 ppm indicating successful incorporation of the zwitterion moiety into the polymer. The peak at 4.75 ppm is the ‘solvent peak’ resulting from the deuterium oxide solvent. All peaks should be easily identifiable and have an amplitude of at least 5% of the solvent peak.

### 19F NMR: A successful result is one that includes broad peaks within 5ppm of -140 and -150ppm indicating successful incorporation of the PFPA moiety into the polymer. These peaks should be broad with a small amplitude but still easily identifiable.

## Crystal Violet Absorbance: The absorbance value shall be a minimum of 0.400.

## Coating Adherence Verification: The absorbance value shall be a less than or equal to 0.02.

# Attachments

## Attachment 1: Protocol Training Record Form

## Attachment 2: Shelf Life Evaluation Data Sheet

## Attachment 1 – Protocol Training Record Form

## By signing this form, I understand the requirements and responsibilities to executing this protocol.

| Trainee Name | Initials | Signature | Date | Department / Company |
| --- | --- | --- | --- | --- |

## Attachment 2 – Suspension Shelf Life Evaluation Data Sheet

## Aged Suspension Information

| Part Number: | C.SLQ001 | Manufacture & Expiration Date: | Mfg date: Exp date: |
| --- | --- | --- | --- |
| Part Revision: |  | Batch / Lot #: |  |

## Suspension Quality Control Test Results

| ## Attribute | ## Result | ## Meets Criteria? |
| --- | --- | --- |
| ## Appearance (color) | ## See attached test report* | ☐ PASS ## ☐ FAIL |
| ## Appearance (form) | ## See attached test report* | ☐ PASS ## ☐ FAIL |
| ## Residue on Evaporation | ## See attached test report* | ☐ PASS ## ☐ FAIL |
| ## pH | ## See attached test report* | ☐ PASS ## ☐ FAIL |
| ## 1H NMR | ## See attached test report* | ☐ PASS ## ☐ FAIL |
| ## 19F NMR | ## See attached test report* | ☐ PASS ## ☐ FAIL |

## * Attach the Suspension Quality Control test report to this data sheet

Finished Catheter Testing

| ## Attribute | ## Result | ## Meets Criteria? |
| --- | --- | --- |
| ## Crystal violet absorbance | ## See attached test report** | ☐ PASS ## ☐ FAIL |
| ## Coating adherence | ## See attached test report*** | ☐ PASS ## ☐ FAIL |

## ** Attach the Crystal Violet Absorbance Quality Control test results to this data sheet

## *** Attach the Coating Adherence Quality Control test results to this data sheet

## OVERALL TEST RESULTS (circle one):  PASS  /  FAIL

Completed By (Print name, Sign and Date):_________________________________________________________________

Reviewed By (Print name, Sign and Date):__________________________________________________________________