Suspension Shelf Life Evaluation,

2 Year Aged, Silq Surface Treatment

Test Report

# Executive Summary

| ## Status: | ## PASS |
| --- | --- |

## Testing was conducted in accordance with test protocol VV.SLQ028 Rev A to verify the Silq Suspension (part number C.SLQ001) meets the established quality attributes and performance requirements after a shelf life of two (2) years.

## The C.SLQ001 Suspension solution is a water-based coating solution for polymeric and elastomeric surfaces.  The solution contains a three-component methacrylate copolymer, polymerized via free radical polymerization from polysulfobetaine methacrylate, perfluorophenylazide methacrylate, and methacrylic acid.  The material is dispersed at 1% w/v% in high purity water.  The Silq Surface Treatment process may be applied to a variety of polymeric substrates.  A representative catheter (16 Fr Foley catheter) was chosen to evaluate the performance of aged Suspension solution after aging for a period of greater than 2 years.

## Testing consisted of verification of the quality attributes of the aged Suspension (i.e., appearance, residue on evaporation, pH, 1H NMR, and 19F NMR) and the confirmation of the presence and adherence of the Suspension to a sample set of devices.

## Test samples (i.e., coated catheters) were manufactured using the aged Suspension in accordance with established manufacturing procedures including packaging and sterilization.

## Two deviations occurred during the execution of the test protocol. The deviations had no impact on the outcome of the test. See Section 2 for a detailed summary of the deviations.

## Testing performed met the respective acceptance criteria for quality attributes, crystal violet absorbance, and coating adherence testing.  Testing demonstrated the Suspension (C.SLQ001) performs as required after real-time aging for two (2) years.

# Deviations

## Deviation #1: Equipment ST-012 (Weighing scale) was used past its calibration due date to execute quality control procedure QC-C.SLQ001.  The weighing scale ST-012 was due for calibration on 04 Oct 2023.  The QC procedure was executed on 18 July 2024 at risk with the plan to calibrate the equipment after testing. The weighing scale was sent out for calibration on 02 Aug 2024 and was found to be in tolerance when received by Microprecision Calibration, Inc. Additionally, no adjustments were made to the weighing scale during calibration. Therefore, the data obtained using the scale for procedure QC-C.SLQ001 is deemed acceptable. A copy of the calibration report is provided in Attachment 6.

## Deviation #2: The sample size and analysis of data was conducted using attribute data instead of variable data as called out in the protocol.  For a 90% confidence / 90% reliability interval, the sample size is n = 22 for attribute data with zero failures (see Table 4 of QM.SLQ011).  This deviation has no impact on the results of this protocol as the sample size for attribute data achieved the required 90% confidence / 90% reliability interval.  The protocol will be updated to reflect attribute data analysis.

# Data Analysis

## For quality attribute testing of the Suspension, each lot of Suspension was be tested for a sample size of 100%.

## Attribute data (Pass/Fail) was collected for the performance tests (crystal violet absorbance and coating verification). The sample size of twenty-two (22) columns is based on attribute data demonstrating 90% confidence / 90% reliability interval with zero failures.

# Results

## Attribute testing was performed on Suspension lot number HYX-825021 on 18 Jun 2024 in accordance with QC-C.SLQ001.  Lot number HYX-825021 was manufactured on 25 Aug 2021 by Silq in accordance with established manufacturing procedures and full lot traceability documented in the corresponding Device History Records (DHRs).  A copy of the Suspension DHR is provided as Attachment 1.  The Suspension was stored at ambient conditions in Pathway’s manufacturing facility in Santee, CA.  All attributes (i.e., appearance, residue on evaporation, pH, 1H NMR, and 19F NMR) met established acceptance criteria summarized in Table 1 below.  The completed quality control test datasheet is provided in Attachment 3.

## Table 1. Suspension Analysis results

| ## Attribute | ## Result | ## Acceptance criteria | ## Pass / Fail |
| --- | --- | --- | --- |
| ## Appearance (color) | ## Opaque White | ## Opaque white | ## Pass |
| ## Appearance (form) | ## Inviscid Liquid | ## Inviscid liquid | ## Pass |
| ## Residue on Evaporation | ## 1.014% | ## 0.5% – 1.5% | ## Pass |
| ## pH | ## 6 | ## 5 – 8 | ## Pass |
| ## 1H NMR | ## Broad Peaks are present at approximately 0.99, 2.2, 2.9, 3.1, 3.5, 3.7, 4.4, and 4.75 ppm indicating successful incorporation of the zwitterion moiety into the polymer. See 1H NMR spectra in Attachment 3. | ## Identify Peaks with Amplitude >5% of Solvent Peak (4.75ppm) | ## Pass |
| ## 19F NMR | ## Two peaks at -140 and -150 ppm identified. No other peak was present at -76ppm indicating there is no TFE remaining in the suspension See 19F NMR spectra in Attachment 3. | ## Identify Peaks, if there is a peak at -76ppm determine approximate amplitude relative to peaks at -140ppm and -150ppm | ## Pass |

## The aged Suspension was then used to manufacture test samples using a representative catheter (16 Fr Foley catheter).  All test samples were manufactured, packaged, and sterilized by Pathway in accordance with established manufacturing procedures and full lot traceability documented in the corresponding DHRs.  A copy of the finished device DHR is provided in Attachment 2.

## Test samples were subjected to the Crystal Violet absorbance test outlined in QC test procedure OM.SLQ003. After staining each test sample with crystal violet dye and rinsed, the absorbance of each sample was measured.  All 22 test samples passed the acceptance criteria of a minimum absorbance value of 0.400.  Full results are provided in Attachment 4.

## Test samples were subjected to the Coating Adherence verification test outlined in QC test procedure OM.SLQ004.  Each test sample is rinsed, and the absorbance value of the effluent was measured.  All 22 test samples passed the acceptance criteria of a maximum absorbance value of 0.02. Full results are provided in Attachment 5.

# Conclusion

## The quality attributes of the two (2) year real-time aged Suspension PASSED the specified acceptance criteria for Suspension C.SLQ001).

## The performance testing of finished catheters using two (2) years real-time aged Suspension PASSED the specified acceptance criteria for crystal violet absorbance and coating adherence verification testing.

## The results of this testing support the Suspension (C.SLQ001) shelf life to be extended to two (2) years from the date of manufacture.

# Attachments

## Attachment 1 –  Suspension Certificate of Analysis, Lot No. HXY-825021

## Attachment 2 – Coated Catheter Device History Record, Lot No SLQ-80040411241E

## Attachment 3 – Suspension Quality Control Report, Aged Suspension

## Attachment 4 – Crystal Violet absorbance test datasheet, Lot No. SLQ-80040411241E

## Attachment 5 – Coating Adherence verification test datasheet, Lot No. SLQ-80040411241E

## Attachment 6 – Weighing Scale ST-012 Calibration Report

## Attachment 7 – Protocol Training Record