# Document Change Order (DCO)

| DCO #: | 087 | Change Priority: |  | High Medium Low |  |
| --- | --- | --- | --- | --- | --- |
| DCO Type: | Permanent DCO Permanent DCO - Administrative Temporary Deviation Order – Duration: |  |  |  |  |
| Document Category: | Policy Standard Operating Procedure (SOP) Design Document Work Instruction (WI) Technical Spec Labeling |  |  |  |  |
| Req. Effective Date: | 10 Dec 2025 |  | Originator: |  | Ethan Rao |

| Document Title | Document / Part Number | Current Revision | Description of / Reason for Change |
| --- | --- | --- | --- |
| Medical Device Reporting | QM.SLQ022 | A | Description of change: - Revised procedure incorporates the complete FDA definition of “Serious Injury” per 21 CFR 803.3, including medical or surgical intervention required to preclude permanent impairment |
|  |  |  | Reason for change: - MDR reporting nonconformity was identified as a gap in the accuracy and completeness of Silq’s internal Medical Device Reporting procedure. |

Additional Risk Assessment required?   Yes    No (Complete below section)

| Applicable Doc / Part No. | Risk Assessment activities (include doc numbers when applicable) |
| --- | --- |
| CAPA 002-2025 | A review of the Risk Management File confirmed that: No new hazards or failure modes have been introduced. No change in device risk profile has occurred. No adverse impact to product safety or performance is anticipated. |

Verification or Validation required?   Yes    No (Complete below section)

| Applicable Doc / Part No. | Corresponding V&V activities (include protocol numbers when applicable) |
| --- | --- |
| CAPA 002-2025 | Verification will consist of documented confirmation that MDRs associated with PCFs 2023001 and 2023004 were submitted to FDA within the required format and that acknowledgement receipts were received from FDA. Copies of submitted MDRs and FDA confirmation will be retained in the CAPA file. |

Training required?   Yes    No (Complete below section)

| Applicable Doc No. | Notification Type | Effectiveness Check? |
| --- | --- | --- |
| CAPA 002-2025 | Read and Acknowledge only Interactive Training | Yes No |

Material disposition required?   Yes    No (Complete below section)

| Applicable Part No. | Disposition Requirements (Scrap, Use-as-Is, Rework) |  |  |  | Additional Info |
| --- | --- | --- | --- | --- | --- |
| Material disposition not appicable | Stock Material: N/A | WIP: N/A | Finished Good: N/A | Field: N/A | N/A |

Potential Regulatory Impact?   Yes    No (Complete below section)

| Regulatory Assessment: | The corrective actions implemented under this DCO are procedural and administrative in nature and do not reflect changes to device design, manufacturing processes, specifications, materials, labeling, or performance characteristics. |
| --- | --- |

| Approvals: |  |  |
| --- | --- | --- |
| Printed Name | Department | Signature/Date |
| Ethan Rao | QA/RA/Originator | View sign-off sheet |
| Verne Sharma | Exec. Management | View sign-off sheet |
| Brian McVerry | Manufacturing/ R&D/Engineering | View sign-off sheet |