Document Control

Standard Operating Procedure

# Purpose

## The purpose of this standard operating procedure is to define the overall document control process to add, remove, change, and control access to SILQ documents.

## This formal document control system provides a process to:

### Review, approve, and re-approve documents for adequacy by the appropriate and pre-assigned approvers prior to release, including the date and signature of the individual signing shall be documented;

### Allow reviewers to assess the impact a change has on product(s) efficacy, product safety or a Quality Management System (QMS) process;

### Document the description and reason for changes to documents with each revision of the document;

### Maintain a consistent document numbering system to ensure all documents include a unique identifier;

### Maintain access to only relevant versions of documents to applicable users;

### Ensure documents introduced into the system are legible and readily identifiable;

### Establish controls for introducing and controlling external documents;

### Prevent unintended use of obsolete documents with sufficient control to identify their status;

### Retain and secure documents (including quality system records); and

### Ensure that any training required due to changes made to documents has been communicated to the training coordinator for completion in accordance with QM.SLQ003, Employee Training SOP.

# Scope

## The following procedure applies to all documents and quality system records to be controlled and maintained in the .SILQ document control system as part of the Quality Management System.  This will include, but is not limited to:  SILQ policies, standard operating procedures, work or build instructions, forms, design documentation, technical drawings, technical specifications, process specifications, material specifications, equipment calibration/maintenance records, environmental monitoring procedures, and manufacturing batch records.

## This procedure also applies to the retention and security of documents and quality system records.

## The procedure applies to any documents supplied by external sources or suppliers that are required to be imported into the SILQ document control system, such as agreements that have been physically signed, supplier communications, letters or correspondence from regulatory authorities, etc.

## Approval of company documents by suppliers, as deemed necessary, will be performed based on the defined terms as part of the specific supplier agreement with that supplier.

## For the handling and retention of documents holding confidential patient information see QM.SLQ028, Protection of Confidential Patient Information.

# Documents

## Reference Documents

### QM.SLQ003		Employee Training

### QM.SLQ014		Electronic Document System Work Instruction

### QM.SLQ026 		Part Number Assignment Work Instruction

### QM.SLQ028		Protection of Confidential Patient Information

### ISO 13485:2016	Medical Devices – Quality Management Systems – Requirements for Regulatory Purposes (Section 4.2 Documentation requirements)

### 21 CFR Part 11		Electronic Records and Electronic Signatures

### 21 CFR Part 820	Quality System Regulation (820.40 Document 					Controls)

### EU Directive for Medical Devices (MDD), 93/42/EEC (Annex II Section 3.2, Annex V, Section 3.2, Annex VI Section 3.2)

### NBOG’s Best Practice Guide 2014-3, Guidance for manufacturers and Notified Bodies on reporting of Design Changes and Changes of the Quality System

## Appendices / Associated Forms / Templates

### Appendix 1	Document Approval Index

### Appendix 2	Criteria for Substantial Changes

### Appendix 3 	Record Retention Requirements

### FM1-QM.SLQ001	Document Control Order (DCO) Form

### FM2-QM.SLQ001	Authorized Approvers Form

### TMP1-QM.SLQ001	Controlled Document Template

### FM1-QM.SLQ014	Electronic Signature Acknowledgement Form

# Definitions

## Abbreviations / Acronyms

### CFR: Code of Federal Regulations

### DCO:  Document Change Order

### ISO: International Organization for Standardization

### NBOG:  Notified Body Operations Group

### QA:  Quality Assurance

### QMS:  Quality Management System

### SOP:  Standard Operating Procedure

### V&V: Verification and Validation

### WI:  Work Instruction

## General Definitions

### Controlled Document:  A document that requires formal control of revision, content, and release to ensure that the most current, approved revision is available for user.

### Document Change Order (form):  The form used to add new or change existing documents in the document control system.  The required information is listed in the form to manage the document control process and record the required approvals.

### Document or Part Number:  A unique alphanumeric identifier that is used to track documents within the document control system.

### Effective Date: The date an approved document is effective for use at the company.

### FileHold: Software based document management system used to electronically store controlled documents.

### Revision Number:  Represents the version of the document that allows the system to differentiate between the newest and previous older documents.

### Release Date:  The date a document has been released for use at the company. Note: Release Date may not be the same as the Effective Date.

### Substantial Change: Any change to product or quality management system that could affect conformity to the essential requirements of the Medical Devices Directive 93/42/EEC.

# Responsibilities

## It is the responsibility of the designated document administrator to implement and maintain the overall document control process defined in this procedure.

## It is the responsibility of all SILQ personnel and subcontractors to follow this procedure for submitting and approving new documents and for all changes to existing documents.

## It is the responsibility of all document reviewers and approvers, especially within their expertise, to consider the effects adding or making changes to a document will have on product safety and efficacy.  In addition, consideration must be given on how additions or changes can affect function of the QMS to maintain product safety and quality.

## Regulatory management is responsible for ensuring the regulatory assessment, when applicable, for changes to documents and processes is sufficient and in compliance with applicable regulations.

## Quality Assurance management is responsible for ensuring compliance with the document control process and the retention of records by SILQ personnel and subcontractors.

# Procedure:  Document Categories

## SILQ has the following hierarchy of documents that defines their overall position in the company’s QMS:

### Policy:  These are the top level company policies that apply to the company as a whole and define the overall approach for a company in a specific subject area.  The primary example of this would be the Quality System Manual.

### Standard Operating Procedure (SOP):  These are the procedures that implement and translate the company policy.  An example would be this document control procedure.

### Design Documents:  These documents define the design of company products for realization into actual product for use in pre-clinical, clinical, or commercial use.  An example would be product requirements and risk analysis documents.

### Work Instructions (WI):  These documents represent the base quality system procedures that provide instructions at sufficient level to build, rework, inspect, clean, receive, and/or record critical activities of company product lines.  As an example, these would include product travelers and assembly instructions.  In addition, other areas outside of manufacturing require work instructions to maintain facilities and support administrative controls.

### Technical Specifications:  These documents represent the technical translation of the design documents to define a product in full detail that can be built in accordance with its WI’s.  As an example, these would include engineering drawings, BOM, and PCB layout schematics.

### Labeling:  These documents define labeling that will be placed on the product or included with the product.  An example would be Instructions for Use and primary package labels. Content distributed through website or social media platforms are included in this document category.

## Overall, policies represent the top level documents in the company, with SOPs and design documents translating policy into procedure or product design, and then WIs and technical specs providing the specific details to define, build, and maintain company products.

## Labeling represents an exclusive document category field that provides information to users and patients outside of the company.  These documents require additional scrutiny and maintenance to ensure safe use of company products.  Proper and extensive review is essential for approval of labeling documentation.

## Document categories will be used in combination with the Approvals Index to identify the required approvers for a given document.  The Document Approvals Index is included in Appendix 1 and provides a list of required approvers for each document category and type.

# Procedure:  Assigning and Maintaining Document Number(s)

## Every document will be assigned a unique number for tracking and identification.  Document numbers are to be assigned to all documents (except technical drawings and specs) using the following schema:

YY.SLQZZZ

Document Code

Sequential Number

### Document Codes identify the document type and are to be selected from the following:

| ### AD | ### Administrative |
| --- | --- |
| ### CL | ### Clinical |
| ### DC | ### Design Control |
| ### OM | ### Operations/Manufacturing |
| ### RM | ### Risk Management |
| ### SW | ### Software |
| ### TF | ### Technical File |
| ### QM | ### Quality Management System |
| ### VV | ### Verification / Validation |

### Sequential Number is the next available sequential number for the Document Code using the format SLQZZZ, where SLQ001 is the first document assigned a controlled document number.

QM.SLQ001

The following represents an example document number:

### where the document code “QM” indicates it is a QMS document, and the sequential number identifier is “SLQ001” (showing it is the first QMS document assigned a controlled document number).

### For technical drawings, manufacturing procedures, and specifications, a separate numbering system is to be applied. Please see Part Number Assignment Work Instruction, QM.SLQ026.

FM1-QM.SLQ001 or TMP1-QM.SLQ001

Forms (FM) and templates (TMP) are considered to be associated with the original document and will keep the same document number with an ‘FM’ or ‘TMP’ and a sequential number placed in front of the document number.  Forms and templates are typically approved in the same document change order as the associated original document. However, changes may occur to forms and templates without any changes to the associated original document through the same approval route. An example of the document number listed above would

FM1-QM.SLQ001 or TMP1-QM.SLQ001

be:

## Document Number Log

### Only the SILQ document administrator is allowed to assign document numbers.

### The document administrator will maintain a log of assigned document numbers.  The log must contain at minimum the following:

#### Assigned document number

#### Document title

#### Date assigned

#### Used/Not-used

### It is recommended that the log be organized by document code, with a segregated log for part numbers.

### Document numbers of obsolete documents are not to be reused, with every document number assigned unique, regardless if the document is obsoleted.

### The log can be maintained in a digital format on the secure network or cloud server.

### At the end of every month (at a minimum) the document administrator is to save a copy of the document number logs as a pdf, import into FileHold document management software, and file in “Document Number Log Archive” folder. This is stored on a secure network server.

# Procedure:  Revision Control

## The revision of a document is to be assigned by the document administrator.  Revisions will be designated by sequential letters starting with ‘A’ for all documents except for technical drawings, specs, and manufacturing documentation.

### The following letters are not to be used: I, O, Q, S, X.

## For technical and manufacturing documentation, if the document is considered pre-production the document will be assigned a sequential number starting with ‘01’.

## Once a technical or manufacturing document is being changed to a production level document the revision will be reassigned to a sequential letter starting with ‘A’.

## In certain cases where a letter or number revision has reached the end (Ex. Z or 99), an additional character can be added to the revision designation (Ex. AA or 100).

## Released documents are to have the current revision in the header.

## The revision is to be updated on each document by the document administrator upon approval of a document change order (see Section 13).

## Previous revisions of documents are to be removed from circulation by the document administrator to ensure the most up to date documents are available to users.

## All previous revisions of documents are to be archived in secure file storage.

## Previous versions of document can be requested from the document administrator, however all previous version copies must be designated with “archive”, “obsolete” or as “previous revision”.

# Procedure:  Document Approvals Index

## An index with required approvals for any given document has been included in a table within Appendix 1.  The required approvals for any given document are based on the document category (the x-axis of the table) and the document code (the y-axis of the table).  The combination will indicate a specific group set (A, B, C, etc.) that is correlated to the approvals group table in Appendix 1.  This table lists the required department approvals for any given group set listed in the approvals index.

## As an example, a quality management system SOP would require a ‘Group D’ set of approvals that indicate approvers from Executive Management, QA, RA, R&D/Engineering, and Manufacturing would need to sign off on the document in the DCO.

## If necessary, the document administrator can add additional approvers as requested by the DCO originator or the current department approvers.

## Changes to the approval index require an approved DCO for this Document Control SOP.

## Specific personnel associated with each of the department approvals are assigned by executive management with the Authorized Approvers form (FM2-QM.SLQ001) of this procedure.  Each department is assigned a primary approver and secondary approver.  The form is to be approved by Executive Management and maintained on a secure network.

# Procedure:  Creating and Releasing New Documents

## New documents (except technical drawings and specs) are to be created based on the SILQ controlled document template (TMP1-QM.SLQ001).

### For detailed steps, please see Electronic Document System Work Instruction, QM.SLQ014.

## At a minimum new documents are to include the following sections:  Purpose, Scope, Documents, Definitions, Responsibilities (if applicable), and Procedure (or the primary topic, e.g., design requirements, safety review).  Additional sections can be added as needed.

## It is recommended that documents incorporate the font type, numbering style, headers, and footers as demonstrated in the template.

## Technical specs and drawings are expected to come in many different formats however require at a minimum a title, document/part number, revision, and a release / effective date.

## Prior to submission of the document(s) the originator can request a document number(s) be pre-assigned.

## If a document number is not pre-assigned, the document administrator will assign a document number while processing the DCO.

## Upon completion of the document(s), the originator is to complete a DCO form (FM1-QM.SLQ001) and submit the form along with an electronic copy of the native document(s) to the document administrator via email.  Instructions on completing a DCO form are included in Section 14.

## Upon receipt of the form the document administrator is to assign a sequential DCO number (starting with 001) to the DCO form.

## The document administrator is to review the completed DCO form and document(s) submitted for any errors or discrepancies. The administrator can either correct the errors (if minor) or return to the originator via email for corrections.

## Once documents are set for routing to approvers, the document administrator is to determine the required approvals defined in the Approvals Index (see Appendix 1) for each document.

## The document administrator is to populate the approvals box on the DCO form with the required approvers for each document and then route the DCO and the document(s) to each of the approvers for review and their signature for approval.  Multiple approval boxes for different document sets can be made on single a DCO.

### Multiple approval boxes can be used to organize groups of documents from the DCO that have the same approval sets.

## Once ready for review, Approvers are alerted by FileHold via email and by an alert notification within FileHold.

## In addition to reviewing for appropriate and correct content, approvers are to consider the impact changes could have on product efficacy, quality, and safety.  Approvers are to consider if additional testing, training, risk mitigation, or material disposition activities are sufficiently in place.

## If an approver during his review requires changes due to discrepancies or disagreement with the documents content, the approver is to redline or note required changes that need to be addressed.  The documents need to then be returned to the document administrator via email and/or through FileHold’s review workflow as a feedback document.

## The document administrator is to return the redlined or marked up documents to the originator via email for corrections or changes.

## Once changes have been made, the document administrator is to update the DCO number with an alpha character (ex. DCO# 001A) and reroute the form with the updated documents.  If necessary the DCO form can be updated to reflect changes for the re-route.

## If no additional changes are required and the DCO is signed by all the approvers, the documents are to be released at their new revision level (starting at A or 001 for new documents) into the FileHold document control system.

## Signatures associated with scanned or faxed documents are considered acceptable.

## For the release of new documents see Section 13 of this procedure.

## If a DCO cannot be approved, the document administrator can permanently cancel the DCO.  In these cases, the document administrator checks out the DCO from FileHold, updates file name to include “cancelled” (for example, DCO004 cancelled.docx), is to cross out the approvals box and indicate ‘DCO cancelled’ (using stamp or watermark). The cancelled DCO is checked back into FileHold and the document administrator makes note of cancellation within the DCO log.

# Procedure:  Changing Existing Documents

## For changes to existing documents, the originator is to either redline or summarize the changes made to the document(s).  A copy of the electronic native file is to be obtained from within FileHold. See Electronic Document System Work Instruction, QM.SLQ014 for detailed steps.

## The originator is to complete a DCO form (FM1-QM.SLQ001) and submit the form to the document administrator via email with the redlined document(s) attached or with a description of changes in the DCO.  Instructions on completing a DCO form are included in Section 14.

### See Electronic Document System Work Instruction, QM.SLQ014 for steps.

## If verification or validation activities are required as a result of the proposed changes in the DCO, the originator is to reference or attach any required V&V documentation for review.

## Upon receipt of the form the document administrator is to assign a sequential DCO number (starting with 001) to the DCO form.

## The document administrator is to review the completed DCO form and document(s) submitted for any errors or discrepancies.  The administrator can either correct the errors (if minor) or return to the originator for corrections.

## Once documents are set for routing to approvers, the document administrator is to determine the required approvals defined in the Approvals Index (see Appendix 1) for each document.

## The document administrator is to populate the approvals box on the DCO form with the required approvers for each document and then, using FileHold, route the DCO and the document(s) to each of the approvers for review and approval.  Multiple approval boxes for different document sets can be made on single a DCO.

## In addition to reviewing for appropriate and correct content, approvers are to consider the impact changes could have on product efficacy, quality, and safety.  Approvers are to consider if additional testing, training, risk mitigation, or material disposition activities are sufficiently in place.

## If an approver during his review requires changes due to discrepancies or disagreement with the document changes, the approver is to redline or note required changes that need to be addressed.  The documents need to then be returned to the document administrator.

### Please see Electronic Document System Work Instruction, QM.SLQ014 for more detailed steps.

## The document administrator is to email the originator, alerting them of approver updates/comments to the DCO. The document administrator emails all of the updated redlines to the originator. Originator makes corrections or changes and works with the approvers to come to a consensus on the document update(s).

## Once changes have been made, the document administrator is to update the DCO number with an alpha character (ex. DCO# 001A) and reroute the form with the updated documents.

## Approvers are to also review any referenced or included V&V documentation to support a proposed change listed in DCO.  If insufficient V&V documentation is included, a request is to be made to originator for additional information before approval can occur.

## Approvers are to review the effects on material disposition the proposed changes in the DCO can have on stock materials, WIP, or with finished goods.  Approval of the DCO indicates the listed material disposition is to go in affect.

## If no additional changes are required and the DCO is signed by all the approvers, the documents are to be released at their new revision level (the next sequential number or letter, ex. A → B or 001 → 002) into the document control system utilizing FileHold.

## Signatures associated with scanned or faxed documents are considered acceptable.

## For the release of updated documents see Section 13 of this procedure.

## If a DCO cannot be approved, the document administrator can permanently cancel the DCO.  In these cases, the document administrator checks out the DCO from FileHold, updates the file name to include “cancelled”, is to cross out the approvals box and insert stamp or watermark ‘DCO cancelled’. The cancelled DCO is checked back into FileHold and the document administrator makes note of cancellation within the DCO log.

# Procedure:  Administrative Changes to Documents

## In specific cases, administrative changes can be made to a document with limited approvals required.

## These administrative changes can be corrections to spelling or typographical errors, format errors, updates incorrectly accepted from redlines, and cut/paste errors.

## Administrative changes do not include changes to the content or structure of the document.

## Administrative changes can be made with the DCO form (FM1-QM.SLQ001) by checking the Administrative box for the DCO type and attaching the redlined document(s) or a summary of changes made in the form.

## Only the originator, which can be the document administrator in this case, and QA are required to approve any administrative changes.

## For release of updated documents see Section 13 of this procedure.

# Procedure:  Releasing Documents into Electronic Document Management System

## Upon approval of a DCO, the document administrator is to indicate the effective date and provide their signature as the person responsible for release.

### Please see Electronic Document System Work Instruction, QM.SLQ014 for detailed steps.

## Although documents are considered released with the final approval signature, documents are not effective until the date entered into the DCO by the document administrator.  By default, if no date is specified, released documents are effective as of 3 days of their approval.

## The signed DCO forms are to be filed in appropriate “DCOxxx” folder (where “xxx” is the assigned DCO number). Electronic files are maintained in secure location on a network with limited access.

## The newly released document is to be filed within appropriate Document Type area within FileHold (RM, QM, etc.).  Documents that are Policy or Standard Operating Procedure document categories are to be marked “master copy” using a stamp or watermark. The superseded copy is moved to a “previous revisions” folder within FileHold.

## Current versions of documents are to be maintained in FileHold which is on a secure server. Electronic document file to include document number.

## The previous version of document, if applicable, is to be marked as “previous revision” and moved into an archive file folder within FileHold (title with the document number) in secure server in the cloud.

## In the case of pre-printed forms or logs that are in circulation, the document administrator is to review if changes have an impact on these forms/logs and replace them as necessary.

## In the case of currently issued controlled copies that have been affected by the DCO, the document administrator is to notify the party to which the controlled copies were issued to destroy the copies and replace them with updated revisions.  Email confirmation is required that the notification has been received.

## If training is required has been indicated on the DCO, the document administrator is to notify the training coordinator of the documents and type of training (including if effectiveness checks are required) in an email .  A copy of the DCO can be used in the notification if necessary.

## If effects to material disposition have been indicated, the document administrator is to contact a Manufacturing and/or QA representative with the changes to be made in an email .  A copy of the DCO can be used in the notification if necessary.

## Once DCO has been reviewed for completeness, as a final step, the document administrator files DCO within the “DCO Complete” folder in FileHold.

## Document Change Order Log

### Every approved DCO (and in certain cases cancelled DCOs) is to be recorded into a DCO log that includes, at a minimum, the following information.

#### DCO#

#### Listed Document or Part Numbers on DCO

#### Release date / Effective date

#### Revision

#### Document Originator

#### If risk assessment, validation, training, material disposition, and/or regulatory assessment was required.

### The log shall be maintained in a digital format on the secure network or cloud server.

### At the end of every month (at a minimum) the document administrator is to print a hard copy of the document change order log, initial/date, scan, and import into FileHold’s “DCO Log Archive” folder where it is kept in secure location with limited access.

# Procedure:  Completing DCO form

## The originator is to complete the DCO form for submission to the document administrator utilizing FileHold (Please see Electronic Document System Work Instruction, QM.SLQ014 for detailed steps).  The following are instructions to complete a DCO form (with certain fields to be completed by the document administrator only):

### DCO#:  (Document administrator only; assigns from DCO log.)

### Change Priority:  Indicate the priority with which the originator is targeting to get approval.  High: 1 day, Medium: 2-4 days, or Low: 5-7 days.

### DCO Type:  Indicate if this is permanent DCO (most typical), an Administrative DCO, or a temporary deviation order.

### Document category:  Indicate the document categories relevant to the documents listed in the DCO.  Definitions of these categories are listed above in Section 6.

### Required Effective Date:  Indicate the date the document(s) are required to be effective upon their release date.  Indicate the numbers of days past the release date as such ‘R + X days’.  For example R+0 would indicate the effective date is the same as the release date, whereas R+7 would indicate the effective date is 7 days after the release date.

### Document Title:  Type in the exact document title that will be in the header of the document, on the cover page, and be used to identify the document in the system.

### Document / Part Number:  Type in the assigned unique document or part number for each of the documents included in the DCO.  If a number is left blank, the document administrator will assign the next sequential available number.

### Current Revision:  Type in the current revision of the document to be changed, if a new document type in ‘New’.

### Description / Reason of Change:  Provide a summary of changes to the document and the reasons for the change.  When appropriate, redlines can be attached as the change description with ‘see Redlines’ listed in the DCO field.  A reason for change is required regardless.  Additional pages can be added to the DCO for additional space if necessary.

### Additional Risk Assessment Box:  Indicate if changes to the documents listed in the DCO require additional risk management activities.  If ‘Yes’, type in the applicable document or part number that requires additional risk management activities.  Record the corresponding activities that have been performed to address to assess or manage additional risks.  Reference any applicable documentation. If ‘No’, provide the justification why risk assessment is not required.

### Verification / Validation Required Box:  Indicate if changes to the documents listed in the DCO require V&V activities.  If ‘Yes’, type in the applicable document or part numbers that require V&V activities.  Record the corresponding activities that were performed (ex. Electrical safety testing, biocompatibility, etc.) and reference any applicable test protocol numbers. . If ‘No’, provide the justification why verification / validation is not required.

### Training Required Box:  Indicate if the addition of new documents or changes to existing documents requires training.  If ‘Yes’, type in the applicable document numbers that require training and indicate the type of training required.  In addition, indicate if an effectiveness check is required for the additional training. . If ‘No’, provide the justification why training is not required.

### Material Disposition Required Box:  Indicate if changes to the part numbers listed in the DCO have an effect on the material disposition of materials or product.  If ‘Yes’, indicate whether stock material, WIP, finished goods, or field product for a part number need to be reworked, scrapped, or used-as-is.  If rework is required, attach rework instruction and reference in the ‘Additional Info’ section. . If ‘No’, provide the justification why material disposition is not required.

### Potential Regulatory Impact Box:  Indicate if the changes will have a potential regulatory impact and require a regulatory assessment.  If ‘Yes’, discuss the regulatory impact of the proposed changes with a regulatory representative and provide a preliminary assessment. If ‘No’, provide an explanation of why the changes have no potential regulatory impact.

#### In accordance with Medical Devices Directive 93/42/EEC, SILQ’ Notified Body must be informed of any planned substantial changes to a product or the quality management system.  See Appendix 2 for criteria for substantial changes.

### Approvals Box:  The document administrator enters the names and departments of approvers.

## The DCO needs to be complete before submission to the document administrator for routing.

### Please see Electronic Document System Work Instruction, QM.SLQ014 for detailed steps.

# Procedure:  Temporary Deviation Order

## A temporary deviation order provides a process to make temporary changes or allowances from released specifications or procedures.

## To initiate a Temporary Deviation Order, the originator must complete the DCO as listed above.

## In addition, the duration or termination event for the deviation order is to be specified in the DCO.

## A complete description and reason for deviation is required and is to be included in the DCO.

## The deviation order will cancel upon reaching the termination event or specified duration.  In addition, the deviation order can be terminated by written notification or an email at any time by the originator or their direct supervisor to the document administrator.

## A deviation order is to be released by the document administrator following the same process as a normal DCO.

## Upon cancellation, the deviation order DCO will be marked “cancelled” followed by electronic signature of the document administrator and the date.

## During the manufacturing of pre-clinical product, it is allowed to make redlines to manufacturing procedures during builds without a deviation order if necessary.  All redlines are to include the initial/date of the person making the change(s), and are to be attached to the device history record.

# Procedure:  Obsoleting Documents

## A document can be obsoleted at any time, which removes the document from circulation and removes its access to users.

## To obsolete a document, the DCO form is to be completed with the applicable documents listed and ‘Obsolete Document’ in the description for change followed by the reason for the action.

## The normal DCO process is to be followed for approval and release of the DCO.

## The document is to be marked as ‘Obsolete’ and with the effective date of the action.  The obsoleted document is to be placed into an archive folder (titled Obsolete within FileHold).

# Procedure:  Document Access

## Controlled documents are accessed with FileHold document management system software.

## Access to documents is securely limited and based on the job function or role of the employee. All documents in the document management system are to be marked near document footer as either: Master Copy, Controlled Copy, or Reference Only.

## Master copies are to be maintained on a secure network and are only to be removed for audit purposes, or for review by company management for internal audit or management review purposes.  A master copy can only be changed through the DCO approval process.

## Completed lab notebooks or quality records such as completed forms or logs are also considered master document copies.

## Controlled copies of a document can be requested from the document administrator as required for designated hard copy procedure books.  These documents are to be marked ‘Controlled Copy’ and must be maintained up to date as documents are revised.

### The document administrator will maintain a log of controlled copy documents within the “Document Number Log”.  The log must contain at minimum the following:

#### Assigned document number

#### Document title

#### Revision

#### Location of hard copy

### Upon release of documents, document administrator forwards updated “Controlled Copy” versions to parties that have controlled copies of superseded versions and notifies them that the old versions are to be destroyed.

#### A confirmation email is sent to the document administrator to confirm the old versions have been destroyed and replaced with the updated version

## All local copies, both electronic and hard copies, of any document can be made available to appropriate personnel upon request to the document administrator or through pre-approved access to FileHold.  All copies are to be marked ‘Reference Copy’ or ‘For Reference only’ and are considered uncontrolled documents from that point on.

# Procedure:  External Documents

## In many cases, SILQ will use suppliers or contractors that will provide documentation that will need to be incorporated into the document control system.  These are documents that have either been physically signed or represent communication, notifications and correspondence from external parties.

### To retain an external document into the SILQ document control system, a document cover page (see document TMP1-QM.SLQ001) is to be created and submitted to the document administrator.

### The document administrator is to sign the cover page and include the date of receipt upon receiving the document.

### By default, the document is to be placed in FileHold “External Documents” file folder (titled with the document number) of the “External Documents” drawer by the document administrator. (If external document(s) provided as a hard copy, they will be scanned and imported into FileHold by the document administrator.) Access to the “External Documents” file folder is determined by the individual’s job function or role within the organization.

#### The external document may be placed in other locations in FileHold as warranted.

### In cases where internal company documents have been identified as requiring external approval by suppliers, this will be accomplished by adding the requisite supplier representative to the DCO approval form.

### The process for approving documents with external suppliers are to be defined in the specific supplier agreement between the company and the supplier.

### Document administrator to update “External Documents and Standards Log” located in FileHold to list external document.

## External Regulatory standards shall be controlled and maintained in FileHold.

### External regulatory standards purchased by SILQ are the property of Silq Technologies Corporation and may not be shared or distributed for use outside of SILQ.

### The external standard is to be placed in FileHold in the “Standards Folder” file folder (titled with the standard designation number) of the “External Documents” drawer by the document administrator.  (If external standard(s) provided as a hard copy, they will e scanned and imported into FileHold by the document administrator.) Access to the “Standards Folder” file folder is determined by the individual’s job function or role within the organization.   A DCO is not required to place a standard in FileHold.

### Document administrator to update “External Documents and Standards Log” located in FileHold to list standard.

# Procedure:  Document Storage and Retention

## All documents, including records, are to be maintained electronically within FileHold on a secure network preventing any damage or deterioration.

## Only the document administrator and a QA management representative are to have access to all FileHold’s change management functions (system administrator).

## Documents are to be kept in file folders (digital) within FileHold, with one for the current version and the second for archived copies.

## All files are to be labeled with the document number or DCO number for DCO’s.

## Completed lab notebooks or quality records such as completed forms or logs (from an approved SOP) are also considered to be quality documents and treated as such.

## All quality system documents (and records) are to be retained for a minimum of 5 years based on the maximum potential lifetime of the medical device products at SILQ as required by ISO 13485 (Section 4.2.5 Control of Records), 21 CFR Part 820 (Subpart M Records) and MDD (Annex II Section 6.1, Annex V/VI Section 5/1). See Appendix 3 for retention requirements for all quality record types.

### Obsolete documents are to be retained under the same retention requirements specified above.

## All regulatory submissions and correspondence with notified bodies are to be retained indefinitely.

## Hard copies of quality records are stored and maintained in such a way that they are secure (e.g. lock and key access) and retrievable in facilities that provide a suitable environment to minimize deterioration or damage and to prevent loss.

## All documents, including records, maintained in FileHold shall be backed up nightly by the applicable FileHold hosting service provider. Backed up data shall be available for four weeks after the backup.

Appendix 1 - Document Approvals Index

| Document Code | Document Category |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | Policy | SOP | Design Document | Work Instruction | Technical Spec | Labeling |
| ### AD - Administrative | E | E | ----- | E | ----- | ----- |
| ### CL - Clinical | D | D | C | C | ----- | D |
| ### DC - Design Control | ----- | ----- | C | C | C | D |
| ### OM – Operations/Manufacturing | ----- | E | ----- | B | ----- | E |
| ### RM - Risk Management | D | ----- | D | ----- | ----- | D |
| ### SW - Software | ----- | C | A | A | A | D |
| ### TF – Technical File | ----- | ----- | F | ----- | ----- | ----- |
| ### QM - Quality Management System | D | D | ----- | C | ----- | D |
| ### VV - Verification / Validation | ----- | ----- | C | C | ----- | ----- |
| ### Technical drawings, specifications and manufacturing documents | ----- | ----- | C | C | C | D |
| ### WC – Web Content (Website, Social Medial) | ----- | ----- | ----- | ----- | ----- | F |

| Approval Group | Department Approvals |
| --- | --- |
| A | Originator, QA, RA, R&D/Engineering |
| B | Originator, QA, RA, Manufacturing |
| C | Originator, QA, RA, R&D/Engineering, Manufacturing |
| D | Originator, Executive Mgmt., QA, RA, R&D/Engineering, Manufacturing |
| E | Originator, Executive Mgmt., Manufacturing, QA, RA |
| F | Originator, QA, RA |

Appendix 2 – Criteria for Substantial Changes

Product changes

(Relates to Medical Devices Directive 93/42/EEC Annex II section 4.4 and Annex III section 6)

Product changes should be considered substantial if the change may affect:

- The conformity with the essential requirements and/or

- The indications and/or contraindications and/or warnings determined by the manufacturer to be appropriate to ensure the clinical performance of the device.

When determining whether or not a particular product change is “substantial” following considerations should be made (the list is not exhaustive):

- Changes of the intended purpose and/or the performance specification of the device?

- Changes of the materials of the device?

- Changes of suppliers?

- Are new hazards introduced which have not previously been addressed?

- Are risks associated with existing hazards affected?

- Does the change alter the details on intended use given in the design/type approval technical file submitted to the Notified Body?

- Does the change trigger a need to alter the indications or contraindications for use or warnings necessary to ensure safety and efficacy for the intended use of the device?

- Does the change mean that the device will have different end users or be used in a different manner?

- Does the change mean that the clinical data/performance evaluation data for the original device is not sufficient to assure conformity of the changed device with the required characteristics and performance?

- Is the change a direct result of actions taken related to concerns arising from Post Market Surveillance including incidents/recalls/complaints?

- Is the change driven by the development of the state of the art?

- Does the change relating to a manufacturing process, facility or equipment impact the product´s safety or performance?

Quality System changes

(Relates to Medical Devices Directive 93/42/EEC Annex II section 3.4, Annex V section 3.4, Annex VI section 3.4)

Quality system changes should be considered substantial if (list is not exhaustive):

- The change affects compliance of the devices covered by the quality system with the essential requirements or the approved type/design.

- The change affects the compliance of the quality system with its own regulatory requirements (e.g. if the manufacturer changes archiving from paper to electronic storage).

When deciding whether or not a particular quality system change is “substantial” the following considerations should be made:

- Does the change relating to a manufacturing process, facility or equipment impact the product´s safety or performance?

- Does the change affect the manufacturing technologies (processes and equipment)?

- Does the change affect the location of the manufacturer´s activity?

- Does the change affect product conformity with the essential requirements or the approved type/design (e.g. change of the supplier)?

- Does the change affect the arrangements implemented to achieve continued compliance of the quality system with the relevant harmonized standards and/or the requirements of the directive (e.g. design verification, design validation, organizational structure, process interaction, quality control procedures, addition or deletion or move of a facility)?

- Does the change affect the organization of the manufacturer?

Please refer to NBOG’s Best Practice Guide 2014-3, “Guidance for manufacturers and Notified Bodies on reporting of Design Changes and Changes of the Quality System” for additional details and examples of substantial changes.

Appendix 3 – Record Retention Requirements

| Quality Record Type | Storage Location | Retention Period |
| --- | --- | --- |
| Management Review and Meeting Minutes | FileHold | 5 years |
| Risk Management Files | FileHold | 5 years after last product is distributed |
| Contracts | FileHold | 5 years past the end of the term of the contract |
| Patent Records | FileHold | Permanently |
| Clinical Investigation Records | FileHold | 10 years after approval of the product |
| Quality System Documents (Quality Manual, SOPs, Specifications etc.) | FileHold | 5 years after last product is distributed |
| External Documents (External/customer specification, ISO standards/guidelines, ANSI/AAMI standards, etc.) | FileHold | 5 years after last product is distributed |
| Product Literature | FileHold | 5 years |
| Purchasing Records | FileHold | 5 years |
| Supplier Records Services, Contractors, Consultants | FileHold | 5 years |
| Acceptance Activity Records - QC test data reviews - Inspection and test records - Receiving inspection - First Article Inspections | FileHold | 5 years after last product is distributed |
| Design History File Records - Design Reviews - Plans Team minutes | FileHold | 5 years after last product is distributed |
| Device History Records | FileHold | 5 years after last product is distributed |
| Laboratory Notebooks | FileHold | 10 years |
| Lot History Records / Travelers | FileHold | 5 years after last product is distributed |
| Technical Reports | FileHold | 10 years |
| Environmental monitoring records - Temperature / humidity - Microbiological, etc. | FileHold | 5 years |
| Equipment Maintenance | FileHold | 3 years |
| Equipment Calibration | FileHold | 3 years |
| Equipment Logs | FileHold | 3 years |
| Validation protocols or reports | FileHold | 5 years after last product is distributed |
| IQ, OQ, PQ | FileHold | 5 years |
| Nonconforming Material Reports | FileHold | 5 years after last product is distributed |
| Deviations | FileHold | 5 years after last product is distributed |
| Corrective and Preventive Action | FileHold | 5 years after last product is distributed |
| Complaints | FileHold | 5 years after date of report |
| MDRs (Medical Device Report) | FileHold | 10 years after date of report |
| Vigilance Reports | FileHold | 10 years after date of report |
| Adverse Incidents | FileHold | 10 years after date of report |
| Shipping Records | FileHold | 5 years |
| Internal Audit Records | FileHold | 5 years |
| QS Audit Reports Notified Body (NB) | FileHold | 10 years after last product is distributed |
| Records of QS Surveillance | FileHold | 5 years after last product is distributed |
| Records of FDA Inspections | FileHold | 10 years after last product is distributed |
| Training Records | FileHold | Employee duration + 5 years |
| Regulatory Submissions: - Submissions to FDA (510Ks, Supplements, PMAs etc.) - Technical Files/Design Dossiers | FileHold | Permanently |
| Declarations of Conformity | FileHold | 5 years after last product is distributed |
| Copies of EU / ISO Certificates | FileHold | 5 years after last product is distributed |
| Regulatory Agency Correspondence (Notified Body, FDA etc.) | FileHold | Permanently |