Design Project Planning

Standard Operating Procedure

# Purpose

## The purpose of this standard operating procedure is to define the design project planning process of the SILQ Design Control Program.  Design project planning provides the road map of required design control activities, deliverables, and resources based on the extent of the project scope.

# Scope

## The following procedure applies to all design project plans developed in support of design and development projects for SILQ products.

## In addition, all design project planning activities performed by approved subcontractors for or in support of design projects at SILQ must be in-compliance with the requirements defined in this procedure.

# Documents

## Reference Documents

### QM.SLQ001	Document Control

### QM.SLQ004	Design Control Program

## Associated Forms / Templates

### TMP1-QM.SLQ005	Design Project Plan Template

# Definitions

## Abbreviations / Acronyms

### BOM:  Bill of Materials

### DCO:  Document Change Order

### DFMEA:  Design Failure Modes and Effects Analysis

### DHF:  Design History File

### DMR:  Device Master Record

### PHA: Preliminary Hazard Analysis

### QA:  Quality Assurance

### QMS:  Quality Management System

### SOP:  Standard Operating Procedure

### V&V: Verification and Validation>

## General Definitions

### Design History File (DHF): A compilation of records that describes the design history of a finished device.

### Device Master Record (DMR):  The compilation of all the instructions, drawings and other documents that required to produce a completed product.

### Design Input: The physical and performance requirements relating to the device and its intended use. These requirements establish a basis for performing subsequent design tasks, and focus on the needs of the users and patients.

### Design Output: The results of a design effort at each design phase and at the end of the total design effort.  The finished design output is the basis for the device master record.  The total finished design output consists of the device, its packaging and labeling, and the device master record.

### Design Review:  A documented, comprehensive, systematic examination of a design to evaluate the adequacy of the design requirements, to evaluate the capability of the design to meet these requirements, and to identify problems.

### Document Change Order (form):  The form used to add new or change existing documents in the document control system.  The required information is listed in the form to manage the document control process and record the required approvals.

### FileHold: Software based document management system used to electronically store controlled documents.

### Validation:  Confirmation, by examination and provision of objective evidence, that the particular product requirements (design inputs) for a specific intended use (user needs) can be consistently fulfilled.

### Verification:  Confirmation, by examination and provision of objective evidence, that specified requirements (design inputs) have been fulfilled by the product design (design outputs).

# Responsibilities

## It is the responsibility of Quality Assurance management to ensure this procedure and its referenced procedures are followed during design and development of product line at SILQ.

## It is the responsibility of SILQ Manufacturing and R&D/Engineering management in conjunction with the assigned design team or group to ensure design project plans are completed and correct.

## It is the responsibility of project managers to communicate roles, responsibilities, and project requirements and timeline to the project team.

## It is the responsibility of Quality Assurance management to perform periodic review or audits to ensure design and development activities are in-compliance with the defined design control program.

# Procedure: Design Project Planning

## Design project planning occurs as the initial phase of an identified design control project.  The output of the project planning is a formalized Design Project Plan.

## The Design Project Plan provides a high-level outline of activities, deliverables, and resources required for the development and design of medical devices at SILQ in compliance with QMS regulations.

## Required design control activities and deliverables of a Design Project Plan are to be organized and tabulated to fall into one of the following phases, unless specified differently by the design project scope:

### Project Planning

### Design Inputs

### Design Outputs

### Design Verification

### Design Validation

### Clinical Investigation

### Design Transfer

## In addition, the project plan is to include:

### Project overview or background

### Summary of project phases (w/ key activities)

### List of project milestones

### Resources and responsibilities (including interfaces between resources) for the listed activities and deliverables

## A design project plan template has been provided as TMP1-QM.SLQ005 of this procedure.  The design project plan template is a recommended format; however, project plans are to maintain at a minimum the topics and required content described in the template.

## Prior to development of a formal project plan, SILQ management and designated project managers are to define the project scope as described in the Design Control Program SOP, QM.SLQ004.

# Procedure: Design Project Activities and Deliverables

## The required set of design control activities and deliverables for each project phase are to be based on the defined design project scope.

## SILQ management and/or the designated project manager are to use the design project scope to define the required set of activities and deliverables to accomplish the objectives of the project.

## To ensure essential design control activities and deliverables are included in the design project plan, a comprehensive list of design control activities and deliverables organized by phase has been provided in the sub-section below.

## New design projects will typically be comprised with the majority of activities and deliverables listed, whereas a project to update an existing design could only require specific set of deliverables depending on the scope of the design update.

## It is up to the discretion of SILQ management and/or the designated project manager to determine the required activities and deliverables based on the requirements of the project.  If certain key deliverables (noted with an asterisk*) are not included, a rationale needs to be given in the body of text of the design project plan.

## Project deliverables are not restricted to only regulatory requirements and may include other optional deliverables such as budgets, detailed schedules, product costing, etc.  If needed, inclusion of business related activities and milestones are appropriate in project plans.  These items may be placed within appropriate phases of the project or may have additional phases added to support specific groupings of activities.

Design Control Activities and Deliverables by Phase

## Project Planning phase:

### Design Project Plan (*)

### Risk Management Plan (*)

### Clinical Strategy, draft

### Regulatory Strategy, draft

### Development team resource allocation

### Identification of critical suppliers, contractors, and consultants

### Draft Software Development Plan, if applicable

### Software Life Cycle, if applicable

### Phase Review – Project Planning (*)

## Design Input phase:

### Definition of Customer Requirements or User Needs (*)

### Identification of regulatory constraints and standards (*)

### Product Requirements: System (*)

### Product or Functional Requirements: Sub-System

### Preliminary Hazard Analysis, initial draft (*)

### Software Requirements Specification, if applicable

### Phase Review – Design Input (*)

## Design Output phase:

### Detailed Technical Specification Documents (*)

### Technical Engineering Drawings (*)

### Source Control Specifications

### Bill-of Materials (BOM) (*)

### Assembly and Build Instructions, draft (*)

### Software Architecture, if applicable

### Software Design Documents, if applicable

### Software Construction and Coding, if applicable

### Manufacturing functional prototypes and device evaluation

### Prototype engineering testing

### DFMEA or other risk analysis, initial draft (*)

### Software Risk Analysis, if applicable

### Updates to risk management plan

### Updates to PHA

### Design Verification and Validation Test Plan (*)

### Software Test Plan, if applicable

### Identify/Qualify test vendors and service providers

### Qualification of suppliers, contractors, and consultants (*)

### Design Verification Protocols, draft (*)

### Design Validation Protocols, draft (*)

### Human Factors Analysis / Study, draft

### Software Test Protocols, draft (*)

### Traceability Matrix, draft (*)

### Technical Design Review (*)

### Phase Review – Design Output (*)

## Design Verification phase:

### Release system specifications and drawings

### Finalized Verification Protocols (*)

### Finalized Software Test Protocols - Code Review, Unit Testing if applicable

### Execution of Verification Protocols (*)

### Verification Test Summary Reports (*)

### Updates to design output deliverables as results of testing

### Updates to Traceability Matrix, Risk Analysis

### Animal Safety Test Protocol, draft

### Clinical Study Protocol and associated documents, draft if applicable

### Clinical Labeling, draft if applicable

### Phase Review – Design Verification (*)

## Design Validation phase:

### Finalized Validation Protocols (*)

### Finalized Software Test Protocols - Integration, System Testing if applicable

### Execution of Validation Protocols (*)

### Validation Test Summary Reports (*)

### Finalized Human Factors Analysis

### Execution of Human Factors Study and Summary report

### Animal Safety Study Execution and Summary report

### Design changes as results of testing

### Baseline release of software and documentation for clinical use

### Finalized Clinical Study Protocol and associated documents

### Finalized Clinical Labeling

### Preparation and Submission of Regulatory package for Clinical use

### Updates to Traceability Matrix, Risk Analysis

### Phase Review – Design Validation (*)

### Clinical Transfer checklist, if applicable (*)

## Clinical Investigation phase (if applicable):

### Approval of regulatory package for clinical use

### Clinical unit build and release

### Execution of clinical study

### Clinical study data analysis and summary report

### Updates to Traceability Matrix, Risk Analysis

### Meetings with applicable regulatory authorities

### Finalizing commercial labeling

### Preparation of Essential Requirements / Technical File, if applicable

### Clinical Evaluation Report, if applicable

### Preparation of FDA submission, if applicable

### Phase Review – Clinical Investigation

## Design Transfer

### Design V&V Summary Report

### Release baseline of production software and documentation

### Audit of Design History File (*)

### Audit of manufacturing operations

### Release of Device Master Record, including labeling and packaging (*)

### Risk Management Report (*)

### Finalize traceability matrix (*)

### Product Specific Quality Plans (*)

### Process FMEAs

### Process Validation

### Internal audit to evaluate QMS compliance

### Submission of regulatory packages

### Regulatory assessments and audits

### Product clearance by regulatory authorities

### Design transfer checklist

### Product release for distribution

# Procedure: Design Project Plan Approval and Updates

## Design Project Plans are to be approved in accordance with the Document Control process (see SOP QM.SLQ001), with additional approvers added as necessary.

## As the project is implemented, the Design Project Plan is to be reviewed at minimum with each scheduled phase review and updated to reflect completion of activities and deliverables.

## If necessary, updates to the Design Project Plan status can be made at any time.

# Procedure:  Design History File (DHF)

## Approved Design Project Plans are to be retained in the DHF for each specific product.

### Plans are to be scanned and imported into FileHold and filed in the appropriate DHF.