Design Output

Standard Operating Procedure

# Purpose

## The purpose of this standard operating procedure is to define the process for translating the design input requirements (design inputs) into design output to ensure design activities are in compliance with 21 CFR 820.30 and ISO 13485:2016.

## This procedure describes the form for design outputs that enables verification against the design and development input requirements and validation to the user needs and the intended use of the device.

# Scope

## The following procedure applies to all design and development activities for medical device product lines at SILQ.

## In addition, all design output activities performed by approved subcontractors for, or in support of, design projects at SILQ must be in compliance with the requirements defined in this procedure.

# Documents

## Reference Documents

### QM.SLQ001	Document Control

### QM.SLQ004	Design Control Program

### QM.SLQ009	Design Verification and Validation

### QM.SLQ013	Risk Analysis

## Associated Forms / Templates

### TMP1-QM.SLQ007 	Source Control Specification Template

### TMP2-QM.SLQ007	Design Input / Output Matrix Template

# Definitions

## Abbreviations / Acronyms

### BOM: Bill of Materials

### DCO: Document Change Order

### DFMEA: Design Failure Modes and Effects Analysis

### DHF: Design History File

### DMR: Device Master Record

### I/O: Input/Output

### V&V: Verification and Validation

## General Definitions

### Design History File (DHF): A compilation of records that describes the design history of a finished device.

### Device Master Record (DMR):  The compilation of all the instructions, drawings and other documents that are required to produce a completed product.

### Design Input: The physical and performance requirements relating to the device and its intended use. These requirements establish a basis for performing subsequent design tasks, and focus on the needs of the users and patients.

### Design Output: The results of a design effort at each design phase and at the end of the total design effort.  The finished design output is the basis for the device master record.  The total finished design output consists of the device, its packaging and labeling, and the device master record.

### Document Change Order (form):  The form used to add new or change existing documents in the document control system.  The required information is listed in the form to manage the document control process and record the required approvals.

### FileHold: Software based document management system used to electronically store controlled documents.

### Specification: Detailed description of a design requirement with which a product, process, service, or other activity must conform.

### Validation:  Confirmation, by examination and provision of objective evidence, that the particular product requirements (design inputs) for a specific intended use (user needs) can be consistently fulfilled.

### Verification:  Confirmation, by examination and provision of objective evidence, that specified requirements (design inputs) have been fulfilled by the product design (design outputs).

# Responsibilities

## It is the responsibility of Quality Assurance management to ensure this procedure is followed during design and development of products at SILQ.

## It is the responsibility of R&D/Engineering management in conjunction with the assigned design team or group to ensure the product design (design output) and resultant DMR deliverables are completed and correct.

## It is the responsibility of Manufacturing management to ensure product designs are suitable for manufacture.

# Procedure: Overview of Design Outputs

## The culmination of product design and development results in the transformation of product design inputs into design outputs. These design outputs define the actual device design from a combination of product specifications and descriptive materials.

## The finished design outputs are the basis for the Device Master Record (DMR).

## The design development phase is comprised of a variety of activities with a range of formality; as such, not all design output documents require formal document control. Documents or document types requiring formal control will be identified in the approved design project plan as requiring review and approval.

## The application of the activities discussed in this document is dependent upon the scope and complexity of the design project.   The approved design project plan will specify the activities and deliverables relevant to the specific design project.

## This procedure does not provide the specific presentation or format for the design output documents.  The types of design output discussed in this document are generally parts of other processes that have procedures of their own with associated forms or templates (i.e., risk analysis, design verification and validation).  Ultimately, the design output must characterize all important aspects of the design and be expressed in terms which allow adequate interpretation.

# Procedure: Design Output Activities

## As described in Section 6, as part of the product development activities, the project team transforms the design input into design output. The process of creating design output includes all the following general activities;

### Prototyping

### Bench/engineering tests

### Engineering analyses

### Risk analysis / risk management

### Specification/ drawing development

### Design verification and validation planning

### Developing production specifications

## Preliminary prototyping and bench testing is often a part of the design effort.  As long as the testing is not being relied upon to verify or validate the design, it may proceed as formally or informally as needed to meet the objectives of the design project plan.  However, prior to starting the design verification / validation activities, the parts, subassemblies, and products tested must be described in change controlled documents.  Additionally, the design verification / validation testing documentation must be formalized in accordance with the requirements described in the Design Verification and Validation SOP, QM.SLQ009.

## As the design development progresses, the design output activities become more defined, comprehensive and formalized.

## Design Output deliverables are typically comprised of, but not limited to, the following documents:

### Engineering drawings

### Assembly and/or component technical specifications

### Component / Raw material source control specs

### Bill-of-Materials

### Detailed system architecture documents

### Principle/theory of operations documents

### Assembly instructions

### Preliminary bench test reports

### Firmware or software source code

### Packaging and labeling specifications

### Safety/Technical reports or white papers

## A template for source control specifications is provided as TMP1-QM.SLQ007 of this procedure.  The source control specification provides details on components, raw materials, and supplies used in or for the manufacturing of the product.

## As part of design output phase a detailed design verification and validation plan is to be created.  The plan should establish that the required testing verifies design outputs meet design inputs.  In addition, validation activities are to be established to qualify the device design meets the user needs and intended use of the device. For more specific details see the Design Verification and Validation SOP, QM.SLQ009.

## Risk analysis activities are to be performed in tandem with product design and development to identify failure modes of the device and to ensure design, protection, and labeling mitigations are sufficiently put in place during product design.  For more specific details see the Risk Analysis SOP, QM.SLQ013.

## As part of the design output phase, a design input/output (I/O) matrix is to be developed as part of the trace matrix activities, tying design inputs, design outputs, and design verification / validation.  A template for the design I/O matrix is provided as TMP2-QM.SLQ007 of this document.  Although complete testing information will not be available at the I/O matrix inception, the matrix is to be updated with the information as it becomes available.  The design I/O matrix represents the key traceability document for design control activities.

# Procedure: Design History File (DHF)

## All output documents resulting from the activities discussed in this procedure constitute design output and should be retained in the DHF or DMR for each specific product line.

### Output documents are to be scanned and imported into FileHold and filed in appropriate DHF or DMR.