Design Transfer

Standard Operating Procedure

# Purpose

## The purpose of this standard operating procedure is to define the process for ensuring that the device design has been correctly translated into production specification and that all supporting activities have been identified, completed and documented to allow proper transfer of the design changes to production and subsequent release for human use.

## This procedure encompasses requirements for transfer and release for both clinical trial use and commercial distribution.

# Scope

## The following procedure applies to all design and development activities for medical device product lines at SILQ.

## In addition, all design and development activities performed by approved subcontractors for or in support of design control projects at SILQ must be in compliance with the requirements defined in the design control program.

## If a subcontractor does not have sufficient design controls (missing elements, deficient, etc.), the subcontractor is to follow the applicable procedures in the SILQ design control program.

# Documents

## Reference Documents

### QM.SLQ001		Document Control

### QM.SLQ009		Design Verification and Validation

### QM.SLQ012		Risk Management

### QM.SLQ020		Purchasing Controls

## Appendices / Associated Forms / Templates

### TMP1-QM.SLQ010	Clinical Transfer Checklist Template

### TMP2-QM.SLQ010	Design Transfer Checklist Template

# Definitions

## Abbreviations / Acronyms

### BOM:  Bill of Materials

### DHF: Design History File

### DMR: Device Master Record

### QMS:  Quality Management System

### SOP:  Standard Operating Procedure

### V&V:  Verification and Validation

## General Definitions

### Design History File (DHF): A compilation of records that describes the design history of a finished device.

### Device Master Record (DMR):  The compilation of all the instructions, drawings and other documents that required to produce a completed product.

### Design Output: The results of a design effort at each design phase and at the end of the total design effort.  The finished design output is the basis for the device master record.  The total finished design output consists of the device, its packaging and labeling, and the device master record.

### Design Review:  A documented, comprehensive, systematic examination of a design to evaluate the adequacy of the design requirements, to evaluate the capability of the design to meet these requirements, and to identify problems.

### FileHold: Software based document management system used to electronically store controlled documents.

### Validation:  Confirmation, by examination and provision of objective evidence, that the particular product requirements (design inputs) for a specific intended use (user needs) can be consistently fulfilled.

### Verification:  Confirmation, by examination and provision of objective evidence, that specified requirements (design inputs) have been fulfilled by the product design (design outputs).

# Responsibilities

## It is the responsibility of Quality Assurance management to ensure this procedure and its referenced procedures are followed during design transfer activities of product line at SILQ .

## It is the responsibility of R&D/Engineering management in conjunction with the assigned design team or group to ensure the DHF and DMR deliverables are completed and correct.

## It is the responsibility of Executive management to ensure that the appropriate regulatory approvals have been obtained before releasing product for use in clinical trials or for distribution in commercial markets.

## It is the responsibility of Manufacturing management to ensure product designs are suitable for manufacture and the DMR is formally released prior to production of commercial units.

## It is the responsibility of Quality Assurance management to perform periodic review or audits to ensure design transfer activities are in compliance with the this procedure.

# Procedure: Design Transfer

## Upon completion of verification and validation activities, the design is ready for the process of transfer into production units under manufacturing control for either clinical use or commercial distribution.

## It must be noted, that product can be transferred for clinical production and used in clinical investigations prior to completion of all validation activities.  Clinical investigations are considered part of the overall validation phase; however prior to use in humans sufficient verification and validation activities is still required to ensure patient/user safety.  A clinical transfer process is required to verify sufficient design control activities have been performed.

## Design transfer represents the following critical processes:

### Review of the DHF to ensure completion and design control requirements have been fulfilled, including risk management activities;

### Review / comparison of the proposed production DMR to assure it is representative of that used for design verification / validation testing;

### Approval by all necessary regulatory authorities for the distribution of commercial product;

### Review of manufacturing capabilities and readiness;

### Translation of the device design into a product ready for manufacture through a reliable and repeatable process.

## Clinical transfer represents a more limited process to ensure initial units for use in investigational studies have:

### Completed all necessary testing requirements;

### Review of clinical product DMR is representative of that used for design verification and validation testing;

### Obtained approval from regulatory authorities for investigational use;

### Clinical labeling that is complete and meets regulatory requirements;

### Sufficient device history recorded for the production of clinical units;

### Completed all required design control deliverables as defined by the project plan.

## A design transfer plan / checklist is created to identify and track the completion status of the activities required to transfer the product.  Templates for Clinical Transfer Checklist and Design Transfer Checklist have been provided in TMP1-QM.SLQ010 and TMP2-QM.SLQ010 of this document, respectively.  The deliverables associated with the design transfer process will vary based on the type and scope of the design changes as identified in the design project plan.  As such, the activities identified in the templates constitute recommended actions which should be assessed for applicability and completeness and adapted as necessary for each product that undergoes design transfer.

## The DHF review / audit conducted as part of the design transfer phase, it to ensure that all elements of the DHF identified in the design project plan have been completed, all identified action items from design reviews have been closed and that the DHF archives are in a retrievable format.  This review does not verify appropriate content of each document, which is the responsibility of the original document approvers and/or design reviews.

## Materials obtained for R&D use during the design of a product may not generally be used in the production of finished product for human use unless the requirements for transfer of R&D materials specified in the Purchasing Controls SOP, QM.SLQ020 have been met.

## Completion and sign-off of the associated Transfer Checklist will be performed in lieu of a final design review to document the closure of the design transfer phase.

# Procedure: Design History File

## The closed DHF audit and the completed / signed Design Transfer Checklist are to be retained in the DHF for each specific product line.

### These documents are to be scanned and imported into FileHold and filed within the DHF for each specific product line.