Risk Management

Standard Operating Procedure

# Purpose

## The purpose of this procedure is to define the risk management process implemented at SILQ to ensure risk management concepts and requirements as defined in ISO 14971, are integrated into design control and the Quality Management System.

## This procedure specifies the process to identify the hazards associated with SILQ medical devices, to estimate and evaluate the related risks, to reduce these risks as far as possible through risk controls, and to ensure the effectiveness of these risk controls.

# Scope

## This procedure applies to all products, ancillary components, and accessories designed and/or manufactured under the QMS at SILQ.

## Risk management is applicable to all stages of the product life cycle at SILQ.

## In addition, all risk management activities performed by approved subcontractors for, or in support of, design control projects at SILQ must be in compliance with the requirements defined in this procedure.

# Documents

## Reference Documents

### QM.SLQ001			Document Control

### QM.SLQ004			Design Control

### QM.SLQ011			Statistical Techniques

### QM.SLQ013			Risk Analysis

### QM.SLQ016 			Corrective and Preventive Action

### QM.SLQ018			Management Review

### QM.SLQ022			Medical Device Reporting

### ISO 14971:2019	Medical devices – Application of risk management to medical devices

### ISO/TR 24971:2020	Medical devices – Guidance on the application of ISO 14971

### FDA: Guidance for the Content of Premarket Submissions for Software Contained in Medical Devices (11 May 2005)

## Appendices / Associated Forms / Templates

### Appendix 1		Recommended Risk Rating Definitions

### Appendix 2 		Recommended Risk Categories

### TMP1-QM.SLQ012 	Risk Management Plan Template

# Definitions

## Abbreviations / Acronyms

### DCO:  Document Change Order

### DHF:  Design History File

### FMEA:  Failure Modes and Effects Analysis

### QA:  Quality Assurance

### QMS:  Quality Management System

### PHA:  Preliminary Hazard Analysis

### RA:  Regulatory Affairs

### SOP:  Standard Operating Procedure

### V&V: Verification and Validation

## General Definitions

### Benefit: Positive impact or desirable outcome of the use of a medical device on the health of an individual, or a positive impact on patient management or public health.

### Design History File (DHF): A compilation of records that describes the design history of a finished device.

### Document Change Order (form):  The form used to add new or change existing documents in the document control system.  The required information is listed in the form to manage the document control process and record the required approvals.

### FileHold: Software based document management system used to electronically store controlled documents.

### Harm: Injury or damage to health of people, or damage to property or the environment.

### Hazard: A potential source of harm.

### Level of Concern: The classification of a software-containing medical device based on the severity of hazard that could be directly or indirectly permitted or inflicted by the operation of the software in the absence of risk mitigations. Levels are Major, Moderate, or Minor.

### Likelihood:  A numeric ranking of the probability of the cause of the hazard occurring.

### Reasonable Foreseeable Misuse:  Use of a product or system in a way not intended by the manufacturer, but which can result from readily predictable human behavior.

### Residual Risk: The risk remaining after control measures have been implemented.

### Risk: The combined likelihood and severity of that harm.

### Risk Analysis: The systemic use of available information on a device or process to identify hazards and to estimate the risk.

### Risk Assessment: The overall process comprised of a risk analysis and risk evaluation.

### Risk Evaluation: The process of comparing estimated risk against given risk criteria to determine its acceptability of the risk.

### Risk Management: The systemic application of management policies, procedures, and practices aimed at analyzing, evaluating, controlling, and monitoring risk.

### Risk Management File: The part of the DHF that contains the set of records and other documents that are produced by risk management activities, including risk management plan, risk assessments, and records associated with the periodic review of the adequacy of the Risk Management Process.

### Risk Mitigation:  An action taken to reduce risks to, or maintained within, specified levels.  Mitigation is synonymous with risk control.

### Severity: The measure of the possible consequences of a hazard.

### State of the Art:  Developed stage of technical capability at a given time as regards products, processes, and services, based on the relevant collective findings of science, technology, and experience.

### Validation:  Confirmation, by examination and provision of objective evidence, that the particular product requirements (design inputs) for a specific intended use (user needs) can be consistently fulfilled.

### Verification:  Confirmation, by examination and provision of objective evidence, that specified requirements (design inputs) have been fulfilled by the product design (design outputs).

# Responsibilities

## It is the responsibility of the SILQ management team to periodically review the risk management process to ensure its continuing effectiveness.

## It is the responsibility of the SILQ management team to assign adequate resources and qualified personnel to execute effective risk management activities.

## It is the responsibility of the SILQ management team to determine the criteria for risk acceptability.

## It is the responsibility of a designated team of R&D/Engineering, Manufacturing, QA, RA, Clinical/Medical, and Executive Management representatives to:

### Define the scope of risk management activities for a design or manufacturing project;

### Develop a risk management plan;

### Integrate risk management activities into product design and development;

### Integrate risk management activities into production and process development and control;

### Review production and post-production information to support risk management activities.

## It is the responsibility of Quality Assurance management to ensure this procedure and its referenced procedures are followed during the design and development of products at or for SILQ.

## It is the responsibility of Quality Assurance management to ensure risk management activities are documented properly and retained in the Risk Management File.

# Policy for establishing Criteria for Risk Acceptability

## The purpose of this policy is to provide guidance for establishing the criteria for risk acceptability used in the evaluation of residual risks associated with medical devices developed and/or manufactured by SILQ.

## This policy applies to all SILQ employees, consultants, and contractors involved in establishing, reviewing, updating, and approving the criteria for risk acceptability in the risk management plans for medical devices developed and/or manufactured by SILQ for commercial distribution.

## The following factors should be considered when establishing the criteria for risk acceptability:

### Applicable local regulatory requirements where the medical device is to be marketed;

### Applicable international standards for the type of medical device;

### The generally acknowledged state of the art, which can be determined from a review of international standards, best practices in technology, results of accepted scientific research, publications from authorities, and other information for similar medical devices; and

### Validated concerns from stakeholders which include communication from users, clinicians, patients, and regulatory bodies.

# Procedure: Overall Risk Management Process

## The critical elements of the risk management process, as identified in ISO 14971, and their interrelationships are shown below:

(

1

)

Risk Management

Planning

(

2

)

Risk Analysis

(

3

)

Risk Evaluation

(

4

)

Risk Control

(

5

)

Evaluation of overall

residual risk

(

6

)

Risk Management

Report

(

7

)

Production and Post

-

Production Information

Risk Assessment

Risk Management

Process

## The risk management process produces living documents that are updated as the understanding of a product or process changes.

## Through a product’s life cycle, risk management is reconsidered when changes occur in the design, process, or application of a device.

## In addition, cumulative knowledge of the device (and similar devices) post-production is reconsidered as new information is found.

## Risk management activities integrate directly into the different phases and processes of the design control program.

# Procedure: Risk Management Planning (1)

## A formal Risk Management Plan must be developed for new products or product families.

## For changes to a product design or manufacturing process that could significantly affect safety or essential performance, two (2) paths are available to address the changes:

### The original risk management plan can be updated to reflect these changes through the DCO process.  Changes in scope and the addition on new risk management activities can be added to the original risk management plan as required.

### A new formal risk management plan can be developed instead of updating a current plan. This new plan can replace the old plan or be supplemental.

## Only a single risk management plan is required for a product or product family.

## Risk management planning is performed as part of design project planning efforts to ensure that sufficient risk management activities are identified for the project scope. A risk management plan template is provided in TMP1-QM.SLQ012.

## The Risk Management Plan establishes the risk management required to be performed for the project. Risk Management Plans must describe:

### Responsibilities for risk management activities

### Planned risk management activities

### Requirements for review of risk management activities

### Criteria for risk evaluation (with risk rating definitions and risk categories)

### Verification activities (including confidence and reliability levels for judging mitigations for each risk category)

### Criteria for risk acceptability

### Post-Production information review activities

## A new, formal Risk Management Plan is unnecessary when a design or process change occurs during the development or testing phase of product life cycle, or if there are changes to a product line that do not significantly impact safety or efficacy. In this case, the DCO process is used to identify required risk management activities. (See QM.SLQ001, Document Control)

# Procedure: Risk Analysis (2)

## Based on the Risk Management Plan (or applicable DCO), risk analysis activities are performed for the device and/or process.

## For new products, a Preliminary Hazard Analysis (PHA) shall be performed in the Design Input phase.

### PHA is a top-down approach that compiles a list of known and foreseeable hazards associated with the product based on its characteristics. Fault conditions are then identified which might cause these hazards.

### When conducted in the Design Input phase, a PHA can drive a product design to minimize its potential risks and hazards.

### Once a PHA has been completed with the development of a product under the design control program, the PHA will be considered the products Master Hazard Analysis.

### The Master Hazard Analysis represents the primary risk analysis for a product to be referenced and updated during the post-market life cycle of the product.

### Only a single PHA (Master Hazard Analysis) is required for an entire product or product family.

### Details on conducting a PHA is provided in Risk Analysis SOP, QM.SLQ013.

## Failure Modes and Effects Analysis (FMEA) is recommended to be performed on all products in the Design Output phase.

### FMEA is a bottom-up approach that identifies hazards based on fault conditions or failure modes identified for specific subunits of a product (Design FMEA) or steps of a process (Process FMEA).

### The FMEA establishes failure modes based on the design of component or process step and then works up to the potential harm to a user.

### Performing an FMEA as a design or process matures helps determine if failure-induced hazards have been sufficiently mitigated.

### To be defined in the Risk Management Plan, as many FMEAs can be developed as required based on the complexity of the design and manufacturing process

## Other risk analysis tools may also be used when judged appropriate, as will be outlined in the Risk Management Plan. Examples of other risk analysis tools include human factors analysis, health hazard evaluations, and fault tree analysis.

## Additional details about performing risk analyses are in the Risk Analysis SOP, QM.SLQ013.

# Procedure: Risk Evaluation (3)

## Each identified risk is to be evaluated by assigning a rating to their severity of potential harm and probability of occurrence (likelihood).

## A recommended rating scale for ‘Severity’ and ‘Likelihood’ is defined in Appendix 1.  These rating scales can be adjusted as more information is gained throughout the life cycle of the product if necessary.  The actual ratings scale definitions used for a product must be defined in the corresponding Risk Management Plan.

## Once ratings are assigned, each risk is classified under a risk category as either Major, Moderate, or Minor as defined in Appendix 2. The actual risk category definitions used for a product must be defined in the corresponding Risk Management Plan.

## Any risk that may result in likely death or irreversible injury must be classified as Major.

## The results of risk evaluation are recorded in the risk analysis documents.

# Procedure: Risk Control (4)

## Every effort must be made to reduce each risk as much as possible.

## Risk is reduced by selecting and implementing mitigations from the following risk control options:

### Inherent safe design and manufacture: Create an inherently safe design and construction.

### Protective measures: Add a protective measure (like a shield or alarm).

### Information for safety and training: Provide information on safety (such as labeling or training).

## For each identified risk, all available risk control options must be applied until additional controls do not improve safety.

## Although providing information on safety to users is required, it’s an unreliable enough mitigation that its impact on risk is small. It therefore must be assumed not to quantifiably reduce residual risk.

## All implemented risk mitigations are considered part of the design, so must themselves be evaluated to determine if they introduce new hazards or affect the risk ratings of previously identified ones.

## The specific mitigations for identified risks are recorded in the risk analysis documents.

## Implementation of risk mitigations shall be verified to be complete and recorded in the risk analysis documents.

## The effectiveness of implemented risk mitigations (e.g., in reducing the severity or likelihood of a hazard) must be verified.

### Verification protocol design (including sample sizes) must be statistically sound, with confidence and reliability levels selected based on the category of the related risk. (Appropriate levels are defined in each project’s Risk Management Plan; SILQ’s recommended values are in the Statistical Techniques work instruction, QM.SLQ011.)

### Records of verification activities for each mitigated risk are referenced in the corresponding risk analysis documents.

### In the rare case when testing limitations prevent a mitigation’s effectiveness from being verified directly, alternative indirect evidence is documented and referenced in the risk analysis.

# Procedure: Evaluation of Overall Residual Risk (5)

## After risk mitigations have been implemented, each identified risk is reevaluated, creating new ratings for the residual risks.

## The updated ratings are evaluated to determine updated risk categories as defined in Section 10.

## Results of the residual risk assessment are documented in the appropriate risk analysis.

## Residual risks are reviewed as a part of the review and approval of risk analysis deliverables. The residual risks of greatest concern are also discussed in design reviews to determine if additional controls need to be implemented.

## Once all residual risk has been determined, an overall benefit-risk analysis is performed considering all identified risks (regardless of risk category).

### The benefit-risk analysis uses data gathered from literature, post-market data for similar devices, and/or internal post-market surveillance.

### If any residual risk remains Moderate or Major, more targeted analysis is also performed to ensure that these risks are substantially outweighed by the benefit of the device.

### The SILQ management team reviews and approves the analysis.

## Acceptability of the residual risk is determined based on the criteria for risk acceptability as defined in the Risk Management Plan and summarized in the Risk Management Report.

# Procedure: Risk Management Report (6)

## Upon completion of the risk management process, a summary report of the risk management activities is created.

## The report contains sufficient information to ensure:

### The activities defined in the Risk Management Plan have been appropriately implemented;

### The overall residual risk is acceptable;

### Appropriate methods are in place to collect and review relevant production and post-production information (see Section 14).

## The completed Risk Management Report is retained in the Risk Management File.

# Procedure: Production and Post-Production Information (7)

## Information about the product in manufacturing or in the field that has relevance to its safety or identifies unrecognized hazards is collected from public and internal sources (e.g. product complaints or adverse event reports).

## Internal sources of information shall include, where applicable, the following:

### Corrective and Preventative Actions per QM.SLQ016

### Non-Conforming Product Reports

### Medical Device Reports (MDR) per QM.SLQ022

### information generated during production and monitoring of the production process

### information generated by the user

### information generated by those accountable for the installation, use and maintenance of the medical device

### information generated by the supply chain

### publicly available information

### information related to the generally acknowledged state of the art

## Safety information about similar devices on the market is also collected.

## Details about obtaining production and post-production information can be found in Post Market Surveillance procedure, QM.SLQ033. A post-market surveillance plan is developed for each product in which the requirements of the plan is directly proportional to the risk associated with the product based on its intended use.

## SILQ evaluates the information for impact on previous risk management activities and provides it as an input back into the risk management process, revising risk management deliverables such as the Master Hazard Analysis as needed.

## Review of post-production information for new or changes to hazards is to occur as an on-going activity as data from post-market surveillance becomes available.

## At a minimum, analysis of production and post-production information as part of the risk management process is to be performed once per year prior to management review meetings.

## A status review of the SILQ risk management process and current risk management file is to be included as part of management review meetings.  A summary of activities and status is to be included as part of the review meeting minutes.

# Procedure: Risk Management File

## The Risk Management File is comprised of all risk management deliverables for a given product or project. It is a complete record of:

### Risk management planning

### Risk analysis

### Risk evaluation

### Risk control (risk mitigation implementation and verification)

### Results of the evaluation of the residual risks (including benefit-risk analysis)

### The final Risk Management Report

## The Risk Management File is a part of the appropriate Design History File.

### Risk Management documents are to be scanned and imported into FileHold and filed into the appropriate DHF.

Appendix 1 – Recommended Risk Rating Definitions

The tables below describe recommended ratings for severity and likelihood for an identified hazard. These ratings enable categorization for the evaluation of risk acceptability.

For a process assessment, the Severity rating considers the behavior of a unit arising from a particular process failure.

| SEVERITY |  |  |
| --- | --- | --- |
| Rating | Qualitative Term | Effects |
| 1 | Negligible/Cosmetic | No or virtually no injury to patient or user; no or virtually no negative effect on the environment. No impact on product performance or user confidence in the product or company; user may or may not even notice the failure. |
| 2 | Minor | Minor injury to the patient or user; minor negative effect on the environment. Slight decline of product performance or of user confidence* in the product or company (e.g. a delay causes the customer to be slightly annoyed or inconvenienced). |
| 3 | Moderate | Moderate injury to the patient or user; moderate negative effect on the environment. Decline of product performance or of user confidence* in the product or company (e.g. customer is very annoyed or dissatisfied). |
| 4 | Critical | Serious but reversible injury to the patient or user (e.g. requires an invasive procedure to treat or increases case complication); severe negative effect on the environment. |
| 5 | Catastrophic | Serious irreversible injury to or death of the patient or user; very severe negative effect on the environment. |

* Loss of user confidence in the product or company is considered a hazard because it may lead to the user ignoring instructions or alarms, or creating workarounds or other uncontrolled procedures.

| LIKELIHOOD |  |  |  |
| --- | --- | --- | --- |
| Rating | Qualitative Term | Design Failure Frequency | Process Failure Frequency |
| 1 | Remote | Probability of failure or defect occurrence is so remote that event is not expected to occur during the lifetime of the medical device. | Likely to occur once during the life of the process. |
| 2 | Low | Unlikely probability of failure or defect occurrence during a procedure or use of the medical device. | Likely to occur less than once per year at anticipated volumes. |
| 3 | Moderate | Occasional probability of failure or defect occurrence during a procedure or use of the medical device. | Likely to occur a few times per year at anticipated volumes. |
| 4 | High | Moderate probability of failure or defect occurrence during a procedure or use of the medical device. | Likely to occur many times per year at anticipated volumes. |
| 5 | Very High | High probability of failure or defect occurrence during a procedure or use of the medical device. | Likely to occur many times per process. |

Risk assessments of product designs should use the Design Failure Frequency ratings.

Because process failures lead to the creation of defective product that will itself have its own likelihood of failure, assessments of processes risks might reasonably use a composite of Design Failure Frequency and Process Failure Frequency.  However, for the sake of simplicity, a process risk assessment may take the more conservative approach of considering only the Process Failure Frequency.  This equates to an assumption that all types of defective product will fail at equivalent rates.

Appendix 2 – Recommended Risk Categories

| RISK CATEGORIES BASED ON SEVERITY AND LIKELIHOOD |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Likelihood | Severity |  |  |  |  |
|  | Catastrophic (5) | Critical (4) | Moderate (3) | Minor (2) | Negligible (1) |
| Very High (5) | Major | Major | Major | Major | Moderate |
| High (4) | Major | Major | Major | Moderate | Moderate |
| Moderate (3) | Major | Major | Moderate | Moderate | Minor |
| Low (2) | Major | Moderate | Moderate | Minor | Minor |
| Remote (1) | Moderate | Moderate | Minor | Minor | Minor |

The categories, as defined here, cannot be assigned based on the product of the Severity and Likelihood ratings, but they can be defined by their sum:

| Severity + Likelihood | Category |
| --- | --- |
| 7 or higher | Major |
| 5 or 6 | Moderate |
| 4 or lower | Minor |

This may be useful for calculating category assignments in a spreadsheet.