Risk Analysis

Standard Operating Procedure

# Purpose

## This procedure defines the following risk analysis methods that may be used to evaluate products and processes developed at or for SILQ:

### Preliminary Hazard Analysis (PHA), a top-down approach to identify and evaluate system-level hazards in the early stages of product design. The PHA represents the foundation of a products Master Hazard Analysis.

### Failure Modes and Effects Analysis (FMEA), a bottom-up approach to identify failure modes of specific elements of a product design or build process and quantify the resulting risks.

# Scope

## This procedure applies to all risk analysis activities carried out for products or processes developed at or for SILQ.

## The scope of risk analysis activities for any given product or project is defined in the Risk Management Plan or DCO, as directed by Risk Management process (QM.SLQ012).

### All of the methods described in this procedure might be used, or some subset, or none of them.

### Methods beyond those described in this procedure may also be used, as warranted by the project or as defined in planning activities.

# Documents

## Reference Documents

### QM.SLQ001			Document Control

### QM.SLQ004			Design Control Program

### QM.SLQ012			Risk Management

### ISO 14971:2019	Medical devices – Application of risk management to medical devices

### ISO/TR 24971:2020	Medical devices – Guidance on the application of ISO 14971

## Appendices / Associated Forms / Templates

### Appendix 1		Questions for PHA Device Characterization

### Appendix 2		Hazard Categories and Examples

### TMP1-QM.SLQ013	Preliminary Hazard Analysis Template

### TMP2-QM.SLQ013 	Hazards Analysis Table Template

### TMP3-QM.SLQ013 	FMEA Worksheet Template

# Definitions

## Abbreviations / Acronyms

### DCO:  Document Change Order

### DHF:  Design History File

### FMEA:  Failure Modes and Effects Analysis

### PHA:  Preliminary Hazard Analysis

## General Definitions

### Benefit: Positive impact or desirable outcome of the use of a medical device on the health of an individual, or a positive impact on patient management or public health.

### Design History File (DHF): A compilation of records that describes the design history of a finished device.

### Document Change Order (form):  The form used to add new or change existing documents in the document control system. It lists the information required to manage the document control process and record the required approvals.

### Effect (in the context of an FMEA):  The harm that may result from an identified Failure.

### Failure:  Termination of the ability of an item to adequately perform a required function.

### Failure Cause:  The underlying cause of a failure, or cause that initiates a process that leads to the failure.  A single cause may result in multiple effects; and several causes my all lead to the same effect.

### Failure Mode:  Manner in which the item under analysis could fail.

### FileHold: Software based document management system used to electronically store controlled documents.

### Harm:  Injury or damage to health of people, or damage to property or the environment.

### Hazard:  A potential source of harm .

### Item:  Any part, component, sub-system, or system that can be individually considered, as well as any process or sub-process.

### Likelihood:  A numeric ranking of the probability of the cause of the hazard occurring.

### Reasonable Foreseeable Misuse:  Use of a product or system in a way not intended by the manufacturer, but which can result from readily predictable human behavior.

### Residual Risk:  The risk remaining after control measures have been implemented.

### Risk:  The combination of the likelihood and severity of that harm.

### Risk Analysis:  The systemic use of available information on a device or process to identify hazards and to estimate the risk.

### Risk Assessment:  The overall process comprised of a risk analysis and risk evaluation.

### Risk Control: The process of making decisions and implementing risk mitigations such that risks are reduced to or maintained within specified levels.

### Risk Evaluation:  The process of comparing estimated risk against given risk criteria to determine its acceptability of the risk.

### Risk Management:  The systematic application of management policies, procedures, and practices aimed at analyzing, evaluating, controlling and monitoring risk.

### Risk Management File:  The part of the DHF that contains the set of records and other documents that are produced by risk management activities, including risk management plan, risk assessments, and records associated with the periodic review of the adequacy of the Risk Management Process.

### Risk Mitigation:  An action taken to reduce risks to, or maintained within, specified levels.  Mitigation is synonymous with risk control.

### Severity: The measure of the possible consequences of a hazard.

### State of the Art:  Developed stage of technical capability at a given time as regards products, processes, and services, based on the relevant collective findings of science, technology, and experience.

### Validation:  Confirmation, by examination and provision of objective evidence, that the particular product requirements (design inputs) for a specific intended use (user needs) can be consistently fulfilled.

### Verification:  Confirmation, by examination and provision of objective evidence, that specified requirements (design inputs) have been fulfilled by the product design (design outputs).

# Responsibilities

## It is the responsibility of the SILQ management team to assign adequate resources and qualified personnel to execute effective risk analysis activities.

## It is the responsibility of R&D/Engineering, Manufacturing, Clinical/Medical, Quality Assurance, and Regulatory Affairs personnel to perform risk analysis and evaluation activities in compliance with this procedure.

## It is the responsibility of Quality Assurance to ensure this procedure and its referenced procedures are followed during risk analysis activities of products at SILQ.

## It is the responsibility of Quality Assurance to ensure risk analysis activities are documented properly and retained in the Risk Management File.

# Procedure: Overall Risk Assessment Process

## The overall risk management is to be conducted as described in SOP QM.SLQ012.

## Various methods may be chosen for risk analysis. The methods described below are the primary methods recommended for SILQ products, however do not represent the only options allowed.  Other risk analysis methods are allowed to be used as necessary and when defined in the applicable risk management plan.

# Procedure: Preliminary Hazard Analysis

## PHA is a powerful design input tool that can drive the minimization of potential risks to patients, operators, or the environment in a product’s design. The PHA is the foundation for a master hazard analysis, which documents the results of the risk assessment and risk evaluation from pre-mitigation to post-mitigation for identified risks.

## PHA is a top-down analysis that compiles a list of known and foreseeable hazards associated with the product based on its characteristics. Fault conditions that may cause these hazards are then identified.

## A template for the PHA is provided in TMP1-QM.SLQ013.

## To initiate the PHA:

### Define the scope of the hazard analysis. This should reflect the scope of the project define in the Risk Management Plan (if applicable).

### Write an overview of the product. It should include the intended use, general function, principle of operation, a brief summary of potential users, and any reasonably foreseeable misuse.

### Describe the set of device characteristics that could be a source of risk, as guided by the questions in Annex A of ISO/TR 24971:2020. (These questions and the factors to consider when answering them are provided in Appendix 1, directly drawn from ISO/TR 24971; the questions are re-stated in Table 1 of the PHA Template, TMP1-QM.SLQ013.)

## Upon completion of the device characterization, identify applicable hazards for the device based on the categories listed in Table C.1 of ISO 14971 (see Appendix 2) and recorded in the hazard identifications table (see Section 8 of PHA Template, TMP1-QM.SLQ013).

## Document the specific hazards that arise from the identified hazardous situations in a Hazard Analysis Table, organized by hazard category. (See TMP2-QM.SLQ013 for a template of the Hazard Analysis Table.)

## For each hazard identified, record the following in the Hazard Analysis Table. (Severity and likelihood ratings will be based on the definitions in the Risk Management Plan, if applicable, or on ones defined in the body of the PHA document.)

### Determine the resulting hazardous situation (local effect) and potential harm (clinical effect), in consultation with clinical and/or medical officers of the company or other appropriate experts.

#### A local effect is the impact of the failure on immediate product function.

#### A clinical effect is the impact of the local effect on the patient, user, or environment.

### Assign a severity rating to the potential harm.

### Determine the potential cause(s) of each hazard by identifying the potential sequences of events that could lead to the hazard. These may be failures in design, failures in software, environmental effects, user error, etc.

### Once potential causes have been identified, rate the likelihood of the hazard absent any risk mitigations.

### Assign a risk category to the hazard based on its ratings.

### Identify specific risk mitigations to be implemented, and list the type of risk control measure each is (Inherent safety by design and manufacture, Protective measures, or Information for safety and training, as defined in QM.SLQ012 Risk Management).

## If a risk mitigation introduces a new hazard or affects an existing hazard, make the relevant changes to the Hazard Analysis Table.

## Throughout the completion of a project, update the Hazard Analysis Table to include any new risk mitigations, as well as references to verification activities that prove risk mitigations have been put in place and are effective.

## Upon completion of all risk control and verification activities, assign post-mitigation risk ratings for the Likelihood of each hazard (based on the implementation of such mitigations), update the Risk Categories, and record these in the Hazard Analysis Table.

## Lastly, summarize the hazard analysis results. Discuss the impact of the mitigations on the entirety of identified risk, determining if all feasible control methods were exhausted and were sufficient to make the residual risks meet the criteria for risk acceptability. If there are any high residual risks, provide additional discussion on the cause for these residual risks.

# Procedure: Failure Modes and Effects Analysis

## FMEA is a bottom-up approach that identifies hazards of systems or processes based on identified fault conditions or failure modes. The FMEA represents a detailed technical approach to examine the potential risks that could arise from each individual component of a product or from each step of a manufacturing process.

## The goal is to identify failures modes from the bottom up that can cause the occurrence of clinical harms already identified in the PHA (or Master Hazard Analysis).

## The following are two (2) primary types of FMEA:

### A Design FMEA (DFMEA) examines the specific assemblies, sub-assemblies, and components of a particular device design. It assumes a design is created exactly as envisioned, and then explores the resulting failures. It may be useful during the Design Output phase to evaluate a maturing device design.

### A Process FMEA (PFMEA) examines the specific steps of a manufacturing (or other) process. It initially assumes a process is executed without any in-process monitoring or quality control tests, and then uses the results of the analysis to determine where the addition of such process failure mitigations is necessary to create the desired result.

### FMEA techniques can also be helpful when interpreting field data after a product is in distribution. Detailed analysis of each cause of an observed failure mode, using information from the field on likelihoods of occurrence, can provide useful quantitative data.

## Based on the risk management plan, perform the requisite FMEAs using the FMEA worksheet template TMP3-QM.SLQ013, as follows:

### First, identify the product (assembly) or scenario being analyzed, or reference such a detailed description. This may include bill of materials, system architecture diagrams, engineering drawings, use case diagrams, flow charts of manufacturing processes, etc. Document the scope on the FMEA worksheet cover page.

#### If the purpose of the FMEA is to determine what additional mitigations may be needed beyond those already implemented, then the current mitigations should be included in the description. (this is typical when analyzing device designs, for which mitigations may be intimately integrated into the design itself.)

#### If the purpose of an FMEA is to determine what mitigations are needed regardless of what may already be implemented, then current mitigations should be excluded from the description. (This is typical when analyzing pilot manufacturing processes to determine what in-process monitoring or quality control testing may be necessary. Such failure detection methods are generally easily separable from the process itself, so what may have been implemented in the existing pilot process is irrelevant.)

### Use the detailed description to break the product assembly or process into one, two, or three hierarchical levels of component or process step, as appropriate.

### Describe the intended function or purpose of each component or step.

### Identify failure modes for each component or step. Consider only single point failures, assuming all other components or steps are functioning as expected. (Other risk analysis methods may be better suited for exploring multi-point failures if necessary.)

### Describe the local and clinical effects of each failure, as described for the PHA in Section 7.7.1.

### Assign a Severity rating based on the scheme in the Risk Management Plan.

### Identify the likely cause or causes of each failure.

### Describe existing mitigations, if any.

### Assign a Likelihood rating based on the scheme in the Risk Management Plan, assuming the presence of only the existing mitigations.

### Determine the Risk Category for each failure mode.

### Identify mitigations that may be implemented to address the failures. If existing mitigations were previously identified, then identify new mitigations, if any, that may be required beyond those already implemented.

### Once verification testing has been completed, record references to these activities in the FMEA worksheet to ensure the mitigations have been implemented.

### Assign post-mitigation ratings for each failure mode based on the verification testing and update the Risk Category of each hazard.

### If an implemented mitigation introduces a new failure mode or affects an existing risk, update the FMEA worksheet to consider the new failure mode or the affected existing risk.

### Upon completion of the FMEA worksheet, complete a summary of FMEA results on the FMEA summary page.

### Discuss the impact of the mitigations on the entirety of identified risk, determining if all feasible control methods were exhausted and were sufficient to make the residual risks meet the criteria for risk acceptability.

### Specifically identify risks which were not subjected to risk controls or did not have mitigations available for implementation.  List any high residual risks that remained in the Major risk category regardless of implemented mitigations and provide additional discussion on the cause for these residual risks.

### The FMEA results summary will be referenced in the risk management report and used as input in the product benefit-risk analysis.

## Additional considerations:

### A process failure is irrelevant until the faulty output is delivered to the final recipient. Because of this, some prefer to split Likelihood into Likelihood and Detectability in process FMEAs.

#### Option 1: Likelihood. In this case, Likelihood is the probability that the process output delivered to the final recipient is faulty. This option may be preferred because of its simplicity and its consistency with other types of risk analysis.

#### Option 2: Likelihood and Detectability. In this case, Likelihood is the probability that the failure occurs during the process, and Detectability is the probability that the failure isn’t detected before the faulty output is delivered to the final recipient. This option may be preferred because of its finer level of detail.

#### The choice of paradigm is left to the discretion of the project manager responsible for executing the process FMEA.

#### In the case Detection is chosen for use, the adjusted rating system in the template is to be selected and documented in the FMEA worksheet.

#### Using detection in design FMEAs is not recommended as it injects process considerations into the design analysis

## Other types of FMEAs for software, application, and maintenance are also available, using the same basic principles described above.  These FMEA approaches can be used as needed based on the risk management plan.

## When necessary, experimentation purposefully injecting faults into a system can assist in understanding the effects of failure modes.

## An examination of a relevant Verification and Validation test plan may bring to light failure modes that were overlooked in an initial risk analyses.

# Procedure: Risk Management File

## Risk assessment deliverables are retained as a part of the Risk Management File for a given product line or project.

## The Risk Management File is included as part of the appropriate Design History File.

### Risk Management documents are to be scanned and imported into FileHold and filed into the appropriate DHF.

# Appendix 1 – Questions for PHA Device Characterization

Factors to be considered in the determination of qualitative and quantitative device characteristics, as provided in (and referenced from) Annex A of ISO/TR 24971:2020, are included below:

- What is the intended use and how is the medical device to be used?

{Factors that should be considered include:

- what is the medical device’s role relative to:

  - diagnosis, prevention, monitoring, treatment, or alleviation of disease,

  - diagnosis, monitoring, treatment or alleviation of or compensation for an injury

  - investigation, replacement, modification or support of anatomy, or a physiological process, or

  - control of conception?

- what are the indications for use (e.g. patient population, user profile, use environment)?

- what are the contra-indications?

- does the medical device sustain or support life?

- can the performance of the medical device be impacted in the event of a security breach (performance degradation or loss of availability)?

- can unauthorized access, unauthorized activities, or loss of data affect the medical device safety?}

- Is the medical device intended to be implanted?

{Factors that should be considered include the location of implantation, the characteristics of the patient population, age, weight, physical activity, the effect of aging on implant performance, the expected lifetime of the implant, the reversibility of the implantation, whether the implant can be modified or configured while implanted and the access connection to perform this modification or configuration (e.g. physical access point or wireless connection to the implanted medical device.}

- Is the medical device intended to be in contact with the patient or other persons?

{Factors that should be considered include the nature of the intended contact, i.e. surface contact, invasive contact, or implantation and, for each, the period and frequency of contact.}

- What materials or components are utilized in the medical device or are used with, or are in contact with, the medical device?

{Factors that should be considered include:

- compatibility with relevant substances;

- compatibility with tissues or body fluids;

- whether characteristics relevant to safety are known;

- is the medical device manufactured utilizing materials of animal origin?}

- Is energy delivered to or extracted from the patient?

{Factors that should be considered include:

- the type of energy transferred;

- its control, quality, quantity, intensity, and duration; and

- whether energy levels are higher than those currently used for similar devices.}

- Are substances delivered to or extracted from the patient?

{Factors that should be considered include:

- whether the substance is delivered or extracted;

- whether it is a single substance or range of substances;

- the maximum and minimum transfer rates and control thereof.}

- Are biological materials processed by the medical device for subsequent re-use, transfusion, or transplantation?

{Factors that should be considered include the type of process and substance(s) processed (e.g. auto-transfusion, dialysis, blood component or cell therapy processing).}

- Is the medical device supplied sterile or intended to be sterilized by the user, or are other microbiological controls applicable?

{Factors that should be considered include:

- whether the medical device is intended for single use or re-use packaging;

- shelf-life issues;

- limitation on the number of re-use cycles;

- method of product sterilization;

- the impact of other sterilization methods not intended by the manufacturer.}

- Is the medical device intended to be routinely cleaned and disinfected by the user?

{Factors that should be considered include the types of cleaning or disinfecting agents to be used and any limitations on the number of cleaning cycles. The design of the medical device can influence the effectiveness of routine cleaning and disinfection. In addition, consideration should be given to the effect of cleaning and disinfecting agents on the safety or performance of the medical device.}

- Does the medical device modify the patient environment?

{Factors that should be considered include:

- temperature;

- humidity;

- atmospheric gas composition;

- pressure;

- light.}

- Are measurements taken?

{Factors that should be considered include the variables measured and the accuracy and the precision of the measurement results, as well as whether the measurement apparatus or data can be compromised. In addition, the need for calibration and maintenance should be considered.}

- Is the medical device interpretative?

{Factors that should be considered include whether conclusions are presented by the medical device from input or acquired data, the algorithms used and confidence limits. Special attention should be given to unintended applications of the data or algorithm, as well as unauthorized manipulation or changes to algorithms and data.}

- Is the medical device intended for use in conjunction with other medical devices, medicines, or other medical technologies?

{Factors that should be considered include:

- identifying any other medical devices, medicines, or other medical technologies that can be involved;

- the potential problems associated with interaction s(such as the medical device impacting the performance of other medical devices); and

- whether the patient follows the instructions for the therapy.}

- Are there unwanted outputs of energy or substances?

{Energy-related factors that should be considered include noise and vibration, heat, radiation (including ionizing, non-ionizing, and ultraviolet/visible/infrared radiation), contact temperatures, leakage currents, and electric or magnetic fields. Substance-related factors that should be considered include substances used in manufacturing, cleaning or testing having unwanted physiological effects if they remain in the product. Other substance-related factors that should be considered include discharge of chemicals, waste products, and body fluids.}

- Is the medical device susceptible to environmental influences?

{Factors that should be considered include the operational, transport and storage environments.  These include light, temperature, humidity, vibrations, spillage, susceptibility to variations in power and cooling supplies, and electromagnetic interference.}

- Does the medical device influence the environment?

{Factors that should be considered include:

- the effects on power and cooling supplies;

- emission of toxic materials;

- the generation of electromagnetic disturbance.}

- Does the medical device require consumables or accessories?

{Factors that should be considered include specifications for such consumables or accessories and any restrictions placed upon users in their selection of these.}

- Is maintenance or calibration necessary?

{Factors that should be considered include:

- whether maintenance or calibration are to be carried out by the user or by a specialist;

- whether special substances or equipment are needed for proper maintenance or calibration;

- traceability of the calibrator values to a higher order reference;

- how to determine when maintenance or recalibration is needed;

- how to verify that calibration is (still) acceptable.}

- Does the medical device contain software?

{Factors that should be considered include whether software is intended to be installed, verified, modified, or exchanged by the user or by a specialist, and the authenticity of a software update.}

- Does the medical device allow access to information?

{Factors that should be considered include accessible Ethernet ports, USB ports, serial ports, and removable hard drives.}

- Does the medical device store data critical to patient care?

{Factors that should be considered include the possibility of the data being modified or corrupted, unauthorized access to the data, and the consequences for the patients.}

- Does the medical device have a restricted shelf-life?

{Factors that should be considered include whether the medical device can deteriorate over time, the impact of storage conditions and primary packaging, the communication of the expiry date (by labeling or an indicator), possibility of use after the expiry date, and the disposal of expired medical devices.}

- Are there any delayed or long-term use effects?

{Factors that should be considered include ergonomic and cumulative effects. Examples could include pumps for saline that corrode over time, mechanical fatigue, loosening of straps and attachments, vibration effects, labels that wear or fall off, long term material degradation.}

- To what mechanical forces will the medical device be subjected?

{Factors that should be considered include whether the forces to which the medical device will be subjected are under the control of the user or controlled by interaction with other persons.}

- What determines the lifetime of the medical device?

{Factors that should be considered include battery depletion, deterioration of materials and failure of components due to aging, wear, fatigue or repeated use. The availability of spare parts should be considered as well.}

- Is the medical device intended for single use?

{Factors that should be considered include:

- whether the medical device self-destructs after use;

- whether it is obvious to the user that the medical device has been used.}

- Is safe decommissioning or disposal of the medical device necessary?

{Factors that should be considered include the waste products that are generated during the disposal of the medical device itself, and the proper sanitization (removal) of all sensitive data on the medical device.  For example, does it contain hazardous material (e.g. toxic chemical or biological agent), or is the material recyclable? If the medical device stores data, proper handling and security of the stored data should be considered, including data removal and retention.}

- Does installation or use of the medical device require special training or special skills?

{Factors that should be considered include the complexity and novelty of the medical device and the knowledge, skills and ability of persons installing, maintaining or using the medical device. This can include training, education, competence assessment, certification or qualification.}

- How will information for safety be provided?

{Factors that should be considered include:

- whether information will be provided directly to the end user by the manufacturer or will it involve the participation of third parties such as installers, care providers, health care professionals, laboratory directors or pharmacists, and whether this will have implications for training;

- commissioning and transferring to the end user and whether it is likely/possible that installation can be carried out by people without the necessary skills;

- based on the expected life of the medical device, whether re-training or re-certification of users or service personnel would be indicated.}

- Are new manufacturing processes established or introduced?

{Factors that should be considered include the application of new or innovative technology and changes in the scale of production. This can also involve changes in contract manufacturing, suppliers and vendors.}

- Can the user interface design features contribute to user error?

{Factors that should be considered include: control and indicators, symbols used, ergonomic features, physical design and layout, hierarchy of operation, menus for software-driven medical devices, visibility of warnings, audibility of alarms, standardization of color coding. See IEC 62366-1 for additional information on usability and IEC 60601-1-8 for alarms.}

- Is the medical device used in an environment where distractions can cause user error?

{Factors that should be considered include:

- the consequence of use error;

- whether the distractions are commonplace;

- whether the user can be disturbed by an infrequent distraction;

- whether repetitive stress can reduce the user’s awareness or attention.}

- Does the medical device have connecting parts or accessories?

{Factors that should be considered include the possibility of wrong connections, similarity to other products’ connections, connection force, feedback on connection integrity, and over- and under-tightening.}

- Does the medical device have a control interface?

{Factors that should be considered include spacing, coding, grouping, mapping, modes of feedback, blunders, slips, control differentiation, visibility, direction of activation or change, whether the controls are continuous or discrete, and the reversibility of settings or actions.}

- Does the medical device display information?

{Factors that should be considered include visibility in various environments, orientation, the visual capabilities of the user, populations and perspectives, clarity of the presented information, units, color coding, and the accessibility of critical information.}

- Is the medical device controlled by a menu?

{Factors that should be considered include complexity and number of layers, awareness of state, location of settings, navigation method, number of steps per action, sequence clarity and memorization problems, and importance of control function relative to its accessibility and the impact of deviating from specified operating procedures.}

- Is the successful use of the medical device dependent on a user’s knowledge, skills and abilities?

{Factors that should be considered include:

- the (intended) users, their mental and physical abilities, skill, and training;

- the use environment, ergonomic aspects, installation requirements;

- the capability of intended users to control or influence the use of the medical device; and,

- the personal characteristics of intended users that can affect their ability to successfully interact with the medical device. See IEC TR 62366-2.}

- Will the medical device be used by persons with special needs?

{Factors that should be considered include:

- users with special characteristics, such as disabled persons, the elderly, and children who might need assistance by another person to enable the use of a medical device;

- users having wide-ranging skill levels and differing cultural backgrounds and expectations that could lead to differences in what is considered appropriate application of the medical device.}

- Can the user interface be used to initiate unauthorized actions?

Factors that should be considered include whether the user interface allows the user to enter an operation mode with restricted access (e.g. for maintenance or special use), which increases the possibility of use error and thereby the associated risks and whether the user becomes aware of having entered such operation mode.)

- Does the medical device use an alarm system?

{Factors that should be considered are the risk of false alarms, missing alarms, disconnected alarm systems, unreliable remote alarm systems, and the user’s ability of understanding how the alarm system works. Guidance for alarm systems is given in IEC 60601-1-8.}

- In what way(s) might the medical device be misused (deliberately or not)?

{Factors that should be considered are incorrect use of connectors, disabling safety features or alarms, neglect of manufacturer’s recommended maintenance, unauthorized access to the medical device or to medical device functions.}

- Is the medical device intended to be mobile or portable?

{Factors that should be considered are the need for grips, handles, wheels or brakes, and the need for mechanical stability and durability.}

- Does the use of the medical device depend on essential performance?

- {Factors that should be considered are, for example, the characteristics of the output of life-supporting devices or the operation of an alarm. See IEC 60601-1 for a discussion of essential performance of medical electrical equipment and medical electrical systems.}

- Does the medical device have a degree of autonomy?

{Factors that should be considered include:

- awareness of the user when the medical device with a degree of autonomy generates an error, alarm or failure;

- awareness of the user when intervention in an autonomously performed action is required;

- the ability of the user to intervene in or to abort an action that is performed autonomously; and;

- the ability of the user to select and perform proper corrective actions.

See IEC TR 60601-4-1 for further guidance on medical devices with a degree of autonomy.}

- Does the medical device produce an output that is used as an input in determining clinical action?

{Factors that should be considered include whether incorrect or delayed outputs can result in direct or indirect risks to patients, e.g. an incorrect diagnosis resulting in delayed or omitted therapy for a patient. See Annex H for guidance on in vitro diagnostic medical devices.}

# Appendix 2 – Examples of Hazards

The list in Table C.1 of ISO 14971:2019 can be used to assist in the identification of hazards associated with a particular medical device, which could ultimately result in harm.