Identification and Traceability

Standard Operating Procedure

# Purpose

## The purpose of this standard operating procedure is to define the requirements and process for the identification and traceability of manufacturing processes, manufacturing materials, raw materials, components, assemblies and final product.

# Scope

## This procedure applies to all purchased product and product related items for production use and to all final product manufactured by SILQ.

## This procedure defines the minimum requirements for identification and traceability to be carried out by SILQ, Contract Manufacturing Organizations (CMOs), or service providers in regards to product and product related items for SILQ in accordance with applicable regulatory requirements.

## This procedure does not apply to general supplies such as office supplies or items.

# Documents

## Reference Documents

### QM.SLQ001		Document Control SOP

### QM.SLQ026		Part Number Assignment WI

### QM.SLQ043		Work Order SOP

### QM.SLQ045		Receiving SOP

### ISO 13485:2016 	International Standards Organization – Medical Devices – Quality Management Systems (section 7.5.3 – Identification and Traceability)

### 21 CFR 820 		FDA Quality System Regulation (820.60 – Identification; 820.65 – Traceability)

### 21 CFR 812.5		FDA Quality System Regulation – Labeling of Investigational Devices

### 21 CFR 830		FDA Quality System Regulation – Unique Device Identification

# Definitions

## Abbreviations / Acronyms

### CMO: Contract Manufacturing Organization

### DHR:  Device History Record

### DMR:  Device Master Record

### P.O.:  Purchase Order

### SOP:  Standard Operating Procedure

### UDI: Unique Device Identifier

### WI: Work Instruction

## General Definitions

### Component: any material, substance, piece, part, software, firmware, labeling, or assembly, which is intended to be included in the finished, packaged, and labeled device.

### Lot or Batch:  One or more manufacturing materials, raw materials, components, assemblies or final product that consist of a single type, model, class, size, composition, or software version that are manufactured under essentially the same conditions and that are intended to have uniform characteristics and quality within specified limits.

### Lot or Batch Number:  A unique and distinctive combination of letters or numbers, or both assigned to a production lot or batch of components, materials, or final product for traceability purposes.

### Manufacturing Material:  Any material or substance used in or to facilitate in the manufacturing process: a concomitant constituent, or byproduct constituent produced during the manufacturing process; which is present in or on the final product as a residue or impurity not by design or intent of the manufacturer.

### Non-conforming Product:  Any component, manufacturing material, in–process product or device, final product, or returned product that does not fulfill specified requirements.

### Product:  components, manufacturing materials, in-process devices, finished devices, and returned finished devices.

### Purchase Order:  a document issued by SILQ to a seller (supplier), indicating types, quantities, and agreed prices for products or services the seller will provide to SILQ. It is used to control the purchasing of products and services from external suppliers.   Issuing a purchase order to a supplier constitutes a legal offer to buy products or services. Acceptance of a purchase order by a seller usually forms a contract between SILQ and seller, so no contract exists until the purchase order is accepted.

### Raw Material:  Any material or substance used in the final product or to facilitate the manufacture of the final product.

### Specification: any requirement with which a product, process, service or other activity must conform.

### Supplier:  A contract manufacturing organization (CMO), custom fabricator, or distributor that provides purchased product.

### Unique Device Identifier: An identifier that adequately identifies a device through its distribution and use. The first part of the UDI is unique and linked to all devices of that specific type; the second part of the UDI is linked to homogenous production of devices (e.g. batch, serial number, etc.).

# Responsibilities

## It is the responsibility of Purchasing to ensure P.O. is generated in accordance with applicable sections of this procedure.

## It is the responsibility of SILQ operations personnel or contract manufacturing organization on behalf of the company to maintain logs of lot numbers in accordance with this procedure.

## It is the responsibility of all SILQ operations personnel and applicable contract manufacturing organizations to perform duties in accordance with this procedure.

## It is the responsibility of Quality Assurance to ensure this procedure and its reference procedures and activities are followed and periodic audits of compliance are performed.

# Procedure: General Requirements

## All processes performed for product identification including part number assignments and revision level changes are to be controlled in accordance with QM.SLQ001 Document Control SOP and QM.SLQ026 Part Number Assignment WI.

## Changes to product including but not limited to raw materials, components, assemblies used in the manufacture of the final product or changes to the final product itself are, at a minimum, to be documented and controlled by revision level in order to provide traceability of the changes. Note: In general, changes affecting form, fit or function should constitute a change to the part number.

## Purchase Orders are to have SILQ part numbers and revision and, where applicable, manufacturer’s part numbers (if applicable) referenced.

## Raw material, component, assembly part numbers, lot numbers and/or serial numbers (as applicable) are to be referenced on all documentation used in the manufacture of the final product.

## Lot numbers are to provide traceability to the following:

### Part number

### Revision level (if applicable)

### Raw material lot (if applicable)

### Component log (if applicable)

### Assembly lot (if applicable)

### Date of manufacture

### Quantity

## Product is to be clearly identified and controlled during all stages of handling to prevent the loss of traceability, including any containers, bins, carts, pallets, trays, plastic bags, or fixtures used to protect, store and/or move product.  These items must clearly display the associated material / component / assembly number and status.

## Any identifying information for non-conforming product such as part number or lot number is to be entered into the device history records for product status and traceability.  Product status is to be labeled clearly as non-conforming with the proper disposition on in-process product labels.

# Procedure: Purchased Product

## All purchased products including raw materials, components, fabricated parts, finished goods, manufacturing materials, and supplies are to have unique part numbers assigned and a corresponding specification developed and released.

## Part number, revision level, and lot number are to be recorded in the appropriate Receiving Inspection forms (as applicable).

## Receiving inspection documentation should, at a minimum, include traceability to measurement equipment (if applicable), measurement results (if applicable) and personnel performing the inspection.

## Product and materials are to be identified by the part number, received date, lot number, and acceptance status in one or more of the following methods:

### Stamped

### Etched

### Labeled – product or container

### Attached tag or traveler – product or container

## Part number, revision level, and lot number are to be recorded in the DHR of the final product.

## Batch or lot numbers that have been assigned by the manufacturer (supplier) are acceptable to maintain and reference in the DHR, without having to assign a new identifying number.

# Procedure: Manufactured Product

## Product (including components and assemblies) manufactured for SILQ are to have a unique part number assigned and a corresponding specification developed and released.

## All final product manufactured for SILQ must have a unique lot number assigned by SILQ or the applicable CMO.

## If it is determined that the finished device must be uniquely identified per regulatory requirements or SILQ preference, a Unique Device Identifier (UDI) will be assigned to each individual final device in accordance with 21 CFR Part 830. The UDI provides identification and traceability of each device distributed for use.

## Product is to be identified by the part number, revision level, lot number, and acceptance status.

## Part number and lot number of raw materials, components, and assemblies are to be recorded in the DHR.

## The DHR of product manufactured for SILQ should at a minimum include traceability to key process steps / parameters, equipment and personnel.

# Procedure: Assigning Lot Numbers

## For incoming raw materials, components, and assemblies the manufacturer’s assigned lot number shall be used, if present.  If a lot or batch number has not been assigned by the manufacturer, a lot number shall be assigned by SILQ or the applicable CMO using the applicable purchase order number.

## For components and assemblies manufactured by SILQ, a unique lot number shall be assigned to the lot.

## A log of lot numbers shall be maintained by SILQ operations personnel.  The log must shall (at minimum) the following:

### Lot number

### Part number of the Component / Assembly / Material or other commodity that the lot number is assigned to

### Part name or description

### Part revision

### Assigned by

### Date assigned

## The following schema shall be used to assign lot numbers internally YYMM-XX-Z, where:

- YYMM-XX = Work order number assigned in accordance with QM.SLQ043 Work Order SOP

- Z = next consecutive production run in the work order starting with 1.  For each work order the sequence will begin at 1.

    - Example 1: Three production runs are planned for the third work order of December 2022. The lot numbers associated with the work order are: 2212-03-1, 2212-03-2, and 2212-03-3.

    - Example 2: Two production runs are planned for the first work order of January 2023. The lot numbers associated with the work order are: 2301-01-1 and 2301-01-2.

## Lot numbers are to be assigned by the applicable CMO in accordance with their own internal procedures provided the requirements defined in this procedure are met.

# Procedure: Finished Device Traceability

## A consignee list is to be maintained by the applicable CMO responsible for storage and distribution of finished devices on behalf of SILQ.

## At a minimum, consignee list and distribution records for the finished devices shall include, at a minimum, the following:

### The name and address of the initial consignee;

### The identification and quantity of devices shipped;

### The date shipped; and

### Lot number, and/or unique device identifier (UDI), as applicable.