Product Complaint System

Standard Operating Procedure

# Purpose

## This procedure describes the processing of product complaints at SILQ.

# Scope

## A “complaint” is any indication of the failure of a device to meet customer or user expectations for quality or to meet performance specifications reported by customer. Thus, any written, oral, or returned goods expression of dissatisfaction relative to the identity, quality, durability, reliability, safety, effectiveness, or performance of any device manufactured by this manufacturer would be considered a complaint.

## SILQ does not directly handle product or returned products.  Returned products will be handled and evaluated by the appropriate contract organization or authorized representative.

# Documents

## Reference Documents

### QM.SLQ016		Corrective and Preventive Action SOP

### QM.SLQ022		Medical Device Reporting SOP

### QM.SLQ028		Protection of Confidential Patient Information 					SOP

### QM.SLQ030		Advisory Notices and Recalls SOP

### QM.SLQ033		Post-Market Surveillance SOP

### 21 CFR 820.198	Complaint Files

### 21 CFR Part 803	Medical Device Reporting

### ISO 13485:2016	International Standards Organization – Medical Devices – Quality management systems – Requirements for regulatory purposes (section 8.2.2 Complaint Handling)

## Appendices / Associated Forms / Templates

### FM1-QM.SLQ021	Product Complaint File (PCF)

### TMP1-QM.SLQ021	Template, Complaint Response Letter

# Definitions

## Abbreviations / Acronyms

### PCF:  Product Complaint File

### HIPPA:  Health Insurance Portability and Accountability Act

## General Definitions

### Complaint:  A customer’s written, electronic, or oral expression of dissatisfaction relative to the identity, quality, durability, safety, effectiveness, labeling, packaging or performance of any product after it is released for distribution.

### Complainant:  The person who made the initial report of the incident to a representative of the manufacturer or the importer. The complainant may be the patient, the user of the device or other person.

### Event reporting:  Generic term for any of the processes established by the various regulatory bodies (e.g., Medical Device Report for the U.S.) leading to their notification in the case of an event representing significant interest with respect to safety.

### FileHold: Software based document management system used to electronically store controlled documents.

# Responsibilities

## QA is the company’s formally designated unit responsible for the complaint system.

## Designated company representatives shall receive initial complaint reports and manage interactions with customers.

## R&D/Engineering and Manufacturing shall perform complaint evaluations and when applicable will work in conjunction with contract organizations that directly handle returned product.

## Regulatory Affairs shall determine the complaint reportability to applicable regulatory bodies and respond back to the complainant as necessary.

# Procedure

## Complaints are received from many sources including customers by letter or phone, field personnel, sales representatives, and distributors.

## Specific patient information that may be collected through the complaint process, such as images, physician descriptions, pictures or any other related materials shall be anonymized for purpose of the complaint file. This information shall be treated confidentially per QM.SLQ028 Protection of Confidential Patient Information SOP.

## Employees or agents of the company becoming aware of a complaint shall refer the complainant to QA for complaint initiation. QA shall collect the information used to complete FM1-QM.SLQ021 Product Complaint File (PCF). See section 7 for guidance on completing form.

## Complaints reported to employees or agents shall be reported to QA within two business days.

## Complaints involving possible serious injury shall be reported immediately.

## A unique complaint number (PCF number) shall be assigned by QA to each complaint received. The complaint number is a seven-digit number of the form YYYYXXX where YYYY is the year the complaint was received and XXX is a sequential number that starts at 001 each year. Example: 2019010 is the tenth complaint received in 2019.

## Returned Products

### Returning devices should be sent to the contract organization or authorized representative.

### Even though the device may not be returned, an investigation report shall be completed.

### Returned product shall be treated as a biohazard and handled by the contract organization or authorized representative in accordance with their internal procedures for the handling and decontamination of returned product.

## Product Decontamination

### Product returned from the field will be decontaminated by the contract organization or authorized representative in accordance with their internal procedures.

## Complaint Investigation

### Complaints shall be investigated in accordance with Section 7.4.

### The complaint investigation shall be completed within 30 calendar days from the date the complaint was reported.

#### If the complaint investigation requires longer than 30 days, the justification for extending the timeframe shall be documented and a new timeframe proposed.

## PCF Form

### FM1-QM.SLQ021, PCF Form may be completed electronically, and expanded as necessary (i.e. number of rows or space within a section, etc.).

### If the form is completed electronically, no form content may be deleted.  All signatures must be original.  Once the form contains original signatures, changes/additions/deletions made to that portion of the form shall be initialed and dated by the individual making the changes/additions/deletions.

### Attachments in support of the PCF report must be associated using the PCF number and cross referenced within the form.

## Complaint Trending

### Complaint trends shall be generated on a quarterly basis and in conjunction with the Quality System Management Review meeting. At a minimum, trend reports shall include the following:

#### Quantity of complaints received per quarter

#### Complaint frequencies per quarter

#### Top complaint issues

#### Complaint System Aging

# Procedure: Completing the PCF Form

## Section I: Initial Product Complaint Report

### PCF No.: Product Complaint File number. Unique number assigned by QA and entered into header of form.

### Initiated By, Date: name/date of person completing Product Complaint File report form.

### Product Name: commercial name of the product.

### Model no.: Model number of the product.

### Lot no.: Lot number of the product.

### Manufacturing date: date of manufacture of the product.

### Expiration date: date the product expires.

### Number of units: number of units that are associated with this complaint.

### Accessories and associated devices: describe any accessories to the device or associated devices that are reported as part of the complaint.

### Complainant (Name & Title): individual making initial report to SILQ/agent and their title or role.

### Hospital and/or Distributor Name: contact information such as name, address, and phone number of facility experiencing problem.

### Physician Name: name of physician, if applicable.

### Date Reported: date the complaint was first received by SILQ/agent.

### Date of Event: date the event occurred.

### Will product be returned: will product be returned for investigation; RMA no. if applicable.

### Problem Statement:  enter a concise description of issue to be addressed.

### Event Description: complaint details as reported in the words of the complainant including actions taken.

### Was a patient injured: if there was patient injury and if there was an adverse event.

### Description of patient injury: if applicable, describe the injury.

## Section II: Complaint Follow Up Report

### QA to complete this section once more information has been gathered, filling in incident information and patient information.

## Section III Complaint Reportability Determination – U.S.

### Where the information collected reasonably suggests the device has or may have caused or contributed to a death or serious injury, RA shall determine reportability by completing FM1-QM.SLQ021 Section 3  – Complaint Reportability Determination – U.S. per QM.SLQ022 Medical Device Reporting. This also applies when information collected identifies a device malfunction that would be likely to cause or contribute to a death or serious injury if the malfunction were to recur.

## Section IV: Product Complaint Evaluation Report - Complaint Investigation

### QA shall engage R&D/Engineering to evaluate the device. The investigation shall include:

#### Manufacturing Record Review

##### When the lot number is known, the device history record will be reviewed to determine whether any similar nonconformance or anomalies exist within the manufacturing records.

#### An attempt shall be made to re-create the issue reported.

#### The final test corresponding to the product may be performed to determine if the device meets release requirements.

#### Where appropriate, devices from the same lot (returned or from inventory) may be evaluated to assist in understanding the extent or root cause of a problem.

#### The results of the evaluation shall be recorded. A comparison to the specification shall be made whenever possible.

#### The investigator shall attempt to make a root cause determination. If the root cause cannot be determined, the complaint file shall note that the root cause could not be determined.

#### A recommendation for further action may be made.

#### QA working with R&D/Engineering shall determine if corrective action is required and document the decision on FM1-QM.SLQ021, Section IV – Product Complaint Evaluation Report, Product Complaint File. If corrective action is required, a CAPA shall be initiated in accordance with the CAPA procedure (QM.SLQ016) and the CAPA # recorded. If correction action is not required, the complaint file shall note the reason.

#### QA shall perform a risk evaluation and determine if the risk management file needs to be updated. If risk evaluation is not required, the complaint file shall note the reason.

#### The results of the investigation are documented using FM1-QM.SLQ021, Section IV – Product Complaint Evaluation Report, Product Complaint File. The complaint file shall note:

##### Whether the complaint was confirmed

##### Whether the device met specification

##### If the device was being used for treatment or diagnosis

##### Whether the complaint was the result of user error

##### If issue been addressed

##### If there is a relationship of the product to the incident

##### If a risk evaluation was performed

##### If a CAPA was assigned

## Section V: Response to Complainant - Customer Response

### Where it is determined that a formal complaint reply is warranted, RA or designee shall prepare the reply. The response to the customer may be documented on TMP1-QM.SLQ021. A copy of any correspondence shall be included in the complaint file.

## Section VI: Complaint Closure

### QA is responsible for complaint closure.

### QA reviews complaint file for completeness and completes Documentation Review section.

### When the returned product is not received after thirty days and all other aspects of the investigation have been completed, the complaint file may be closed. If product is received after this closure, the complaint file shall be reopened to perform the device investigation.

# Procedure: Complaint Log

## QA is to create and maintain a Complaints and Reportable Events Log to facilitate complaint tracking and provide a complete list of all complaints and their status.

## The Complaints and Reportable Events Log should contain the following minimum information: PCF #, customer/company site name, date complaint initiated, Date of event, date reported, Initiated By, Product/Device name, model number, serial number, batch and/or lot number, Date Closed and Closed By.

## The Complaint Log is archived on a secure network server.

# PCF Quality Records Retention

## The complaint file shall include all records associated with the opening, reportability, investigation, and correspondence for an individual complaint.

## Records of complaints shall be maintained for a period of five years or a period of time equivalent to the expected life of the device, whichever is greater.

### Records are to be scanned and imported into FileHold by QA and filed within appropriate complaint folder.