Medical Device Reporting

Standard Operating Procedure

# Purpose

## This document defines the requirements and procedure for the determination and implementation of electronic medical device event reporting in the U.S. in accordance with 21 CFR 803.

# Scope

## Activities related to Medical Device Reporting shall comply with the requirements and time table of 21 CFR 803.

## This procedure applies to all marketed medical devices that have been manufactured by, or for, SILQ.

# Documents

## Reference Documents

### QM.SLQ021	Product Complaint System

### QM.SLQ023	eMDR Submission Work Instruction

### QM.SLQ028	Protection of Confidential Patient Information

### US FDA Medical Device Reporting (MDR) Regulations, 21 CFR Parts 803 and 806

### CDRH Guidance Document, Medical Device Reporting for Manufacturers, November 2016.

### FDA MedWatch Form 3500A

### Instructions for Completing FDA Form 3500A

### FDA Modernization Act of 1997, Section 213(a)(2)

## Appendices / Associated Forms / Templates

### FM1-QM.SLQ021	Product Complaint File (PCF)

# Definitions

## Abbreviations / Acronyms

### eMDR:  Electronic Medical Device Report

### EOL:  End of Life

### ESG: Electronics Submissions Gateway

### FDA: Food and Drug Administration

### HIPPA:  Health Insurance Portability and Accountability Act

### MDR:  Medical Device Report

### QA: Quality Assurance

### RA: Regulatory Affairs

## General Definitions

### 5-day Report:  A medical device report that must be submitted by a manufacturer to FDA under 21 CFR Part 803.53, on FDA FORM 3500A or an electronic equivalent approved under 803.14, within 5 work days after the day you become aware of an MDR reportable event.

### Become Aware:  Manufacturers are considered to have become aware of an event when:

#### Any employee becomes aware of information that reasonably suggests that an event is required to be reported within 30 calendar days or that is required to be reported within 5 working days under a written request from FDA in accordance with 21 CFR 803.53(b).

#### Any employee, who is a person with management or supervisory responsibilities over persons with regulatory, scientific, or technical responsibilities (including consultants or contractors), or a person whose duties relate to the collection and reporting of adverse events, becomes aware, from any information, including any trend analysis, that a reportable MDR event or events necessitate remedial action to prevent an unreasonable risk of substantial harm to the public health. (In this case, you must submit a report no later than 5 working days after the day that you become aware.)

### Caused or Contributed:  A death or serious injury was or may have been attributed to a medical device, or that a medical device was or may have been a factor in a death or serious injury, including events occurring as a result of: failure, malfunction, improper or inadequate design, manufacture, labeling, or user error.

### Complaint:  Any written, electronic, or oral communication that alleges deficiencies related to the identity, quality, durability, reliability, safety, effectiveness, or performance of a device after it is released for distribution.

### Electronic Submissions Gateway:  FDA’s internet portal for accepting electronic regulatory submissions (e.g., eMDRs).

### Expected Life of a Device:  The time that a device is expected to remain functional after it is placed into use. Certain implanted devices have specified “end of life” (EOL) dates. Other devices are not labeled as to their respective EOL, but are expected to remain operational through activities such as maintenance, repairs, or upgrades, for an estimated period of time.

### Information that Reasonably Suggests:  Any information, including professional, scientific, or medical facts, observations, or opinions that would cause you to come to a reasonable conclusion that a device has caused or may have caused or contributed to an MDR reportable event.

### Intended Purpose:  The use for which the device is intended according to the data supplied by the manufacturer on the labeling, in the instructions and/or in promotional materials.

### Malfunction:  The failure of a device to meet its performance specifications or otherwise perform as intended. Performance specifications include all claims made in the labeling of the devices. The intended performance of a device refers to the intended use for which the device is labeled or marketed.

### Manufacturer Report Number:  The number that uniquely identifies each individual MDR report submitted by a manufacturer. This number consists of the following three parts:

### The FDA registration number for the manufacturing site of the reported device;

### The four-digit calendar year in which the report is submitted; and

### The five-digit sequence number of the reports submitted during the year, starting with 00001. (For example, the complete number will appear as follows: 3000000000-2018-00001.)

### MDR Reportable Event (or Reportable Event):  An event about which the manufacturers have received or become aware of information that reasonably suggests that one of their marketed devices:

### May have caused or contributed to a death or serious injury; or

### Has malfunctioned and that the device or similar device marketed by the manufacturer would be likely to cause a death or serious injury if the malfunction were to recur.

### Necessitates medical or surgical intervention to preclude permanent impairment of a body function or permanent damage to a body structure. Permanent means irreversible impairment or damage to a body structure or function, excluding trivial impairment or damage

### Permanent:  Irreversible impairment or damage to a body structure or function, excluding trivial impairment or damage.

### Persons Qualified to Make a Medical Judgement:  Persons qualified to make a medical judgment include physicians, nurses, risk managers, and biomedical engineers.

### Remedial Action:  Any action other than routine maintenance or servicing of a device where such action is necessary to prevent recurrence of a reportable event.

### Reportable Malfunction:  Those malfunctions which would likely cause or contribute to a death or serious injury if they were to recur. A malfunction falls within this definition if:

### The chance of death or serious injury occurring as a result of a recurrence of the malfunction is not remote; or

### The consequences of the malfunction affect the device in a catastrophic manner that may lead to a death or serious injury; or

### The malfunction results in the failure of the device to perform its essential function and compromises the device’s therapeutic, monitoring or diagnostic effectiveness, which could cause or contribute to a death or serious injury, or other significant adverse device experiences; or

### It involves an implant malfunction that would be likely to cause or contribute to death or serious injury, regardless of how the device is used; or

### The manufacturer takes or would be required to take a remedial action as a result of the malfunction of the device or to other similar devices.

### The malfunction necessitates medical or surgical intervention to preclude permanent impairment of a body function or permanent damage to a body structure. Permanent means irreversible impairment or damage to a body structure or function, excluding trivial impairment or damage

### Serious Injury:  An injury or illness that:

### Is life-threatening, even if temporary in nature; or

### Results in permanent impairment of a body function or permanent damage to a body structure; or

### Necessitates medical or surgical intervention to preclude permanent impairment of a body function or permanent damage to a body structure.

### Similar Device:  FDA considers a device to be “similar” to another device if the devices have the same:

### Basic design and performance characteristics related to device safety and effectiveness;

### Intended use and function; and

### Device classification and product code.

### Supplemental MDR Report:  Contains any additional information required in the 30-day or 5-day report that was not provided because it was not known or available when the initial report was submitted. Supplemental information is submitted to the FDA via MedWatch form 3500A within 30 calendar days following receipt of such information.

# Responsibilities

## QA/RA is responsible for compliance with complaint handling and MDR reporting requirements.

## QA/RA is responsible for assuring that timely MDR determinations are made.

## QA/RA is responsible for ensuring that MDR reports are filed electronically with FDA within the required time frame.

## QA/RA is responsible for maintenance of the Complaints and Reportable Events Log and management of the MDR file.

## QA/RA is responsible for trend analysis and for presenting trend data to Management Review.

## QA/RA is responsible for alerting senior management when new trends are identified and when there may be information regarding a serious threat to public health.

## Regulatory Affairs personnel are responsible for preparation and submission of electronic MDRs to FDA.

## Qualified persons or Medical Safety Consultant (or designee) is responsible for providing clinical rationales when no MDR report is required.

# Procedure

## General

### All company employees and representatives are potential recipients of information reasonably suggesting that a medical device marketed by SILQ may have been involved in a reportable event.

### Any SILQ employee who learns of a product issue that may have been involved in a death or injury shall report the information promptly (within one business day) to QA/RA.

### Any SILQ employee who learns of a product issue that may potentially lead to a patient death or injury shall report the information promptly to QA/RA.

## Determination of Reportability

### QA/RA shall assess all complaints for reportability in accordance with QM.SLQ021 Product Complaint System.

### The MDR reportability assessment shall be performed and documented using form FM1-QM.SLQ021 Product Complaint File. A copy of the form shall be maintained in the complaint file and, if applicable, the MDR file.

### QA/RA shall document whether or not the complaint event constitutes a serious public health threat or warrants remedial action, in which case QA/RA shall immediately contact senior management for notification to regulatory agencies.

### If there is a serious public health threat or situation that requires submission of a 5-Day Report, the assessment shall be completed within 48 hours of becoming aware. All other complaint events shall be assessed within 20 days of awareness.

### When performing the reportability assessment, all available complaint information shall be taken into consideration.

### When no MDR report is required, the clinical rationale for the reportability assessment shall be made by a person qualified to make a medical judgment.

#### The person qualified to make a medical judgment may be an employee or consultant of SILQ.

### If necessary, QA/RA (or designee) shall contact the treating physician or health care professional reporting the complaint to discuss information related to the circumstances of the event. All communications shall be maintained in the complaint file. Specific patient information that may be collected through the complaint process, such as images, physician descriptions, pictures or any other related materials shall be anonymized for purpose of the complaint file. This information shall be treated confidentially per QM.SLQ028 Protection of Confidential Patient Information.

### An MDR report is required if information reasonably suggests that a medical device marketed by SILQ may have caused or contributed to a:

### Death; or

### Serious injury; or

### Reportable malfunction; or

### Remedial action is necessary to prevent significant risk of substantial harm to the public health; or

### Necessitates medical or surgical intervention to preclude permanent impairment of a body function or permanent damage to a body structure. Permanent means irreversible impairment or damage to a body structure or function, excluding trivial impairment or damage

### FDA has made a written request for the submission of a 5-day report involving a particular type of medical device or type of event.

### A Medical Device Report is not required if it is determined that the information received is erroneous or the device involved was not manufactured or imported by SILQ.

### An MDR is not required if it is determined that an event is solely the result of user error with no other performance issue, and there has been no device-related death or serious injury. Supporting information shall be retained in the complaint file.

### An MDR does not need to be submitted when there is information that would cause a person who is qualified to make a medical judgment to reach a reasonable conclusion that the device did not contribute to the death or serious injury, or that a malfunction would not be likely to cause or contribute to a death or serious injury if it were to recur. This information must be recorded in the complaint file.

### Refer to the FDA MDR Guidance Document for additional considerations when making reportability decisions.

### If the complaint is questionable regarding the need to report to the FDA the assessment outcome should err on the side of caution and report the event.

### When additional information is added to the complaint file after the reportability assessment has been made, this new information shall be reassessed for reportability. In addition, a reportability reassessment shall be conducted once the device investigation is finalized.

#### If the conclusion of the reassessment remains unchanged, the rationale shall be documented on the initial assessment form.

#### If the conclusion of the reassessment changes from reportable to non-reportable (or vice versa), a new form FM1-QM.SLQ021 Product Complaint File, Section 3, shall be completed and the rationale should explain what information led to the change in the reportability outcome.

#### All initial assessments and reassessments shall be maintained in the complaint file and, if applicable, the MDR file.

### Once it has been determined that an MDR event is required, QA/RA shall enter the complaint in the Complaints and Reportable Events Log and the FDA due date shall be recorded.

### Based on the available information, if it is determined that an event is:

### Not reportable, no further action is required.

### Reportable, refer to Section 6.4 for the reporting process.

## MDR Investigations

### QA/RA shall promptly investigate MDR reportable complaints in order to adhere to required reporting timeframes.

### Investigations shall be conducted in accordance with the requirements of QM.SLQ021 Product Complaint System.

### SILQ will make a good faith effort to obtain all information required for the report. At least one written (including email) request for information should be made. If unable to obtain all required information, the MDR file will contain an explanation of why the information could not be obtained, and document SILQ’s efforts to obtain the information.

### MDR investigations (at a minimum) shall include a determination of:

### Whether the device failed to meet specifications; and

### Whether the device was being used for treatment or diagnosis; and

### The relationship, if any, of the device to the reported incident or adverse event.

## Reporting MDRs to FDA

### MDR reportable events must be reported to FDA electronically in accordance with 79 FR 8832 – Federal Register notice dated 14-FEB-2014, Medical Device Reporting: Electronic Submission Requirements.

### Regulatory Affairs will submit the initial MDR report electronically to FDA via FDA’s ESG.

### The eMDR shall be prepared in accordance with the work instructions provided in QM.SLQ023 eMDR Submission Work Instruction.

### Prior to electronically submitting the eMDR, Regulatory Affairs will forward a final draft to Executive Management to review for accuracy.

### Regulatory Affairs shall monitor the electronic communications with FDA and verify that the eMDR is successfully transmitted. If needed, Regulatory Affairs shall contact the appropriate FDA eMDR help desk to resolve any transmission errors.

### If new information is received after the initial eMDR report has been filed with FDA:

#### Supplemental eMDRs shall be prepared and submitted in accordance with the work instructions provided in QM.SLQ023 eMDR Submission Work Instruction within 30 calendar days of receiving the new information.

#### If the initial MDR report contained all of the required information and any new information received does not change the facts and/or conclusions reported in the initial report, a supplemental report is not required.

#### If an initial MDR report for an event has been filed and it is later decided that remedial action needs to be taken to address the reported device problem, a supplemental report that identifies the remedial action taken in response to the related problem is required.

### QA/RA shall update the Complaints and Reportable Events Log, recording the date(s) of submission and Core ID# (Core ID number can be found within FDA ESG Gateway acknowledgements).

## MDR Reporting Timeframes

### The timeframes for reporting MDRs to FDA are listed in Table 1.

## Table 1: MDR Reporting Timeframes

| ## 5-Day Report | ## Due within 5 work days after the day of becoming aware that a reportable event necessitates remedial action to prevent an unreasonable risk of substantial harm to the public health. ## OR ## Due within 5 work days after the day of becoming aware of an MDR reportable event for which the FDA has made a written request for the submission of a 5-day report involving a particular type of medical device or type of event. |
| --- | --- |
| ## 30-Day Report | ## Due within 30 calendar days of becoming aware of information that reasonably suggests that the device may have caused or contributed to a death, serious injury, or reportable malfunction. |
| ## Supplemental Report | ## A supplemental MDR report will be submitted electronically to the FDA to provide required or additional information that was not known or available at the time of the initial report. ## Submission of additional information must be within 30 calendar days of receipt of the new information. |

## Report Submission

### eMDRs are submitted using FDA’s Electronic Submissions Gateway (ESG) Web Trader Hosted Solution (WTHS) in accordance with the work instructions provided in QM.SLQ023 eMDR Submission Work Instruction.

## Quality Records

### Complaints with MDR reports will be readily identifiable from other complaint files by the addition of a dash followed by “MDR” (e.g. 2018001-MDR).

### The MDR report file will be numbered the same as the corresponding complaint number.

### MDR report file shall contain (as appropriate):

#### FM1-QM.SLQ021 Product Complaint File (form).

#### Core ID reference # for the eMDR.

#### Copies of any test reports, laboratory reports, service records and reports, records of investigation, etc.

#### Copies of all correspondence with the initial reporter/customer/medical professional.

#### All correspondence with FDA.

#### Investigation findings and supporting documentation.

#### Reference to any applicable corrective action.

#### Copies of all eMDR reports submitted to FDA.

#### Copies of any FDA MedWatch (3500A) forms received from FDA, User Facilities, or Distributors.

#### Documents verifying that the event has been evaluated in accordance with the applicable requirements of QM.SLQ021 Product Complaint System.

#### References to any other relevant documents or information used during assessment.

## QA/RA shall generate MDR trends (or summary of MDRs submitted) periodically and present at Management Review.

## Record keeping systems must allow access to information that facilitates timely follow-up and inspection by FDA.

## Manufacturers shall permit any authorized FDA employee during all reasonable times to access, to copy, and to verify the records required by this procedure.

## Records of complaints shall be maintained for a period of five years or a period of time equivalent to the expected life of the device, whichever is greater.

## Records of erroneous complaints along with SILQ’s conclusion that an event is not an MDR reportable event shall be maintained in the MDR files for two years from the date of the event or a period equivalent to the expected life of the device, whichever is greater.