Quality Manual

Table of Contents

1	Purpose	4

2	Scope and Introduction	4

2.1	Scope	4

2.2	Non-applicability	5

3	Documents	5

3.1	Reference Documents	5

3.2	Appendices / Associated Forms / Templates	5

4	Definitions	5

4.1	Abbreviations / Acronyms	5

4.2	General Definitions	5

5	Quality Management System	7

5.1	General	7

5.2	Applicable QMS Procedures	7

5.3	Processes	7

6	Documentation Requirements and Data Control Process	9

6.1	Control of Documents	9

6.2	Control of Quality Records	10

7	Management Responsibility	11

7.1	Management Commitment	11

7.2	Customer Focus	11

7.3	Quality Policy	11

7.4	Quality Planning Process	12

7.5	Responsibility, Authority and Communication	13

7.6	Management Review Process	14

8	Resource Management	16

8.1	Provision of Resources	16

8.2	Human Resources	16

8.3	Training Process	16

8.4	Infrastructure	17

8.5	Work Environment and Contamination Control	17

9	Product Realization	17

9.1	Planning Process	17

9.2	Customer Related Processes	18

10	Design and Development Process	19

11	Purchasing	22

11.1	Purchasing Process	22

11.2	Purchasing Information	22

11.3	Verification of Purchased Product	22

12	Production Process	23

12.1	General Requirements	23

12.2	Production Control	23

12.3	Production Process Validation	24

12.4	Identification and Traceability	25

12.5	Control of Customer-Supplied Product / Customer Property	25

12.6	Handling, Storage, Packaging, Preservation and Delivery	26

12.7	Control of Monitoring, Inspection, Measurement and Test Equipment	26

13	Measurement and Analysis Process	27

13.1	General Requirements	27

13.2	Feedback	27

13.3	Internal Audits	27

13.4	Monitoring and Measurement of Processes	28

13.5	Monitoring and Measurement of Product	28

13.6	Control of Nonconforming Product	28

13.7	Analysis of Data	29

14	Improvement Process	30

14.1	Corrective Action	30

14.2	Preventive Action	31

15	Appendix 1: ISO 13485 Section Correlations with QMS Documentation	31

# Purpose

The purpose of this Quality Manual is to provide a descriptive overview of SILQ’s commitment to quality, and to describe the Quality Management System (QMS) policies, responsibilities and requirements that have been established in compliance with applicable internationally recognized standards and regulations. This document defines the Quality Management System used for governing the design, development, and manufacture of SILQ products.  This manual specifies the policies and references the system procedures that comprise the Quality Management System.

# Scope and Introduction

## Scope

The Quality Manual applies to all personnel employed by SILQ including subcontractors, consultants, Clinical Research Organization (CRO) for a clinical trial, Contract Manufacturing Organization (CMO) for the assembly of a final product, and other service providers that could impact conformance to applicable regulatory requirements and product requirements.

The Quality Manual specifies the roles and responsibilities undertaken by the SILQ management under applicable regulatory requirements to manage, perform, and assess work affecting quality by providing the independence and authority necessary to perform the required tasks. The role of regulatory correspondence is undertaken by SILQ personnel.

SILQ is the manufacturer and holds the final responsibility to ensure the quality of all product. The manufacturing process, however, is outsourced to qualified CMO’s that manufacture product under the oversight of SILQ. The role of distributor is also outsourced.

The Quality Manual specifies the policies, processes and the interaction between processes, and documentation structure utilized by the organization to provide a high degree of quality, from initial design and development through manufacturing and distribution in compliance with:

- FDA Medical Device Quality System Regulation, Code of Federal Regulations Title 21, Part 820

- FDA Medical Device Reporting, Code of Federal Regulations Title 21, Part 803

- FDA Medical Devices Reports of Corrections and Removals, Code of Federal Regulations Title 21, Part 806

- FDA Medical Device Recall Authority, Code of Federal Regulations Title 21, Part 810

- ISO 13485:2016 International Standards Organization – Medical Devices – Quality Management Systems – Requirements for Regulatory Purposes

## Non-applicability

### Installation activities (reference 7.5.3 of ISO 13485:2016) do not apply to SILQ. No product(s) or component(s) are distributed where installation is a specified requirement. Product(s) will be shipped ready to use.

### Servicing activities (reference 7.5.4 of ISO 13485:2016) do not apply to SILQ. No product(s) or components(s) require servicing.

### Particular requirements for implantable medical devices (reference 7.5.9.2 of ISO 13485:2016) does not apply to SILQ. No product(s) or component(s) of SILQ are regulated as an implantable medical device.

### Customer Property (reference 7.5.10 of ISO 13485:2016) does not apply to SILQ. No physical customer property will be used or incorporated into the product. Patient data may be obtained for product complaint and failure investigation purposes and is safeguarded in accordance with QM.SLQ028 Protection of Confidential Patient Information.

# Documents

## Reference Documents

### 21 CFR Part 820, FDA Quality System Regulation

### ISO 13485:20016 – Medical Devices – Quality Management Systems – Requirements for Regulatory Purposes

## Appendices / Associated Forms / Templates

### Appendix 1 	ISO 13485 Clause Correlations with QMS Documentation

# Definitions

## Abbreviations / Acronyms

### CMO: Contract Manufacturing Organization

### CRO: Clinical Research Organization

## General Definitions

### Customer Complaint:  Written, electronic or oral communication that alleges deficiencies related to the identity, quality, durability, reliability, safety, effectiveness, or performance of a device after it is released for distribution.

### Design History File (DHF):  A compilation of records which describes the design history of a finished device.

### Design Validation:  Establishing by objective evidence that device specifications conform to user needs and intended use(s).

### Device History Record (DHR):  A compilation of records containing the production history of a finished device.

### Device Master Record (DMR):  A compilation of records containing the procedures and specifications for a finished device. Also referred to as the Medical Device File.

### FileHold: Software based document management system used to electronically store controlled documents.

### Labeling:  Any written, printed or graphic matter affixed to a medical device or any of its containers or wrappers, or accompanying a medical device, related to identification, technical description, intended purpose and proper use of the medical device, but excluding shipping documents.

### Management with Executive Responsibility:  Senior management employees (VP or above) of SILQ who have the authority to establish or make changes to the quality policy and quality system.

### Management Representative:  A member of management, appointed by top management, who, irrespective of other responsibilities, has authority that includes ensuring that processes needed for the quality system are established and maintained, reporting to top management on the effectiveness, suitability and adequacy of the quality system and any need for improvement, and ensuring the promotion of awareness of applicable regulatory and customer requirements throughout the organization.

### Process Validation:  Establishing by objective evidence that a process consistently produces a result or product meeting its predetermined specifications.

### Quality Management System (QMS):  The organizational structure, responsibilities, procedures, processes, and resources for implementing quality management.

### Quality Objective: A measurable quality result that the organization intends to achieve. Quality objectives are based on or derived from the organization’s quality policy and must be consistent with it.

### Quality Policy: The overall intentions and direction of the organization with respect to quality, as established by management with executive responsibility.

### Validation:  Confirmation by examination and provision of objective evidence that the particular requirements for a specific intended use can be consistently fulfilled.

### Verification:  Confirmation by examination and provision of objective evidence that specified requirements have been fulfilled.

# Quality Management System

## General

SILQ has established, documented, implemented, and maintained a Quality Management System in accordance with applicable internationally recognized standards and regulatory requirements as the means by which specified quality policies and objectives are achieved.

## Applicable QMS Procedures

A cross reference table between the ISO 13485 and the SILQ quality system documentation is provided in Appendix 1. The processes described herein are implemented through the procedures listed in the appendix.

## Processes

The processes needed for the Quality Management System and their application throughout the organization have been identified. The appropriate criteria, methods, resources, organizational structure, and quality procedures have been established proportionate to the risk involved to support the planning, operation, control, measurement, monitoring, analysis, and improvement of these processes, and to ensure their effectiveness. SILQ takes a risk-based approach in determining the level of control and monitoring of its processes. Records are established and maintained that demonstrate conformance ISO 13485 and compliance to applicable regulatory requirements.

Changes that are made to Quality Management System processes are evaluated for their impact on the quality management system, evaluated for their impact on the medical devices produced under this quality management system, and controlled in accordance with ISO 13485 and applicable regulatory requirements.

An overview of the processes that make up the SILQ Quality Management System is shown in Figure 1. Detailed descriptions of each process including the inputs received, activities/analysis conducted and the resulting outputs are included in the subsections of the Quality Manual.

Management Responsibility

- Management Commitment

- Customer Focus

- Quality Policy

- Planning

- Responsibility, Authority, and Communication

- Management Review

Customer and Regulatory Activities/Feedback

Measurement, Analysis and Improvement

- Monitoring and measurement

- Control of Non-Conforming Product

- Analysis of data

- Improvement

Resource Management

- Provision of Resources;

- Human resources

- Infrastructure

- Work Environment

Product Realization

- Planning of product realization

- Risk Management

- Customer related Processes

- Design and Development

- Purchasing

- Production

- Control of Monitoring and Measuring and Devices

Product

Customer and Regulatory Activities

Figure 1: Quality System Process Interactions

Outsourced processes may include various contractors, consultants, and contract services including subassembly manufacturing and sterilization. All contract services and materials shall be conducted by approved suppliers. Suppliers’ processes and quality systems are monitored and evaluated to ensure conformance to ISO 13485 and customer requirements and compliance to applicable regulatory requirements. Controls are proportionate to the risk involved and include written quality agreements.

All component, subassembly, and sterilization suppliers are subjected to a risk-based classification and then approved per established requirements. Specifications (e.g., product drawings) are developed and provided to the suppliers. Incoming components and subassemblies are inspected and released per established procedures. Nonconforming materials are reviewed by designated individuals. Where applicable, feedback is provided to the suppliers and corrective actions are developed.

Computer software used in the quality management system is validated in accordance with QM.SLQ032 Software Validation. If changes are made to the computer software after initial validation, re-validation may be required. The specific approach and activities associated with the validation is proportionate to the risk associated with the use of the software. Records of validation activities are maintained and are considered documented evidence of the execution of the validation.

# Documentation Requirements and Data Control Process

## Control of Documents

QMS related documentation is appropriately controlled in accordance with QM.SLQ001 Document Control. These documents are reviewed for adequacy and approved by authorized personnel prior to being issued.

Controlled documents may be in a variety of formats including hard copy and electronic media. A master list is maintained identifying the current revision status of all controlled documents to preclude the use of invalid and/or obsolete documents.

Controlled documents are properly maintained to ensure that they remain legible.

The current revisions of controlled documents are available on the network and accessible at locations (point of use) where operations essential to the effective functioning of the Quality Management System are performed.

Documents of external origin are identified and their distribution is controlled.

Obsolete documents are promptly removed from issue and use. Obsolete document retention is defined. A history file of obsolete documents is retained.

Changes to controlled documents and data are reviewed and approved by the same functions that performed the original review and approval, unless specifically designated otherwise. The designated functions authorized to review and approve document changes have access to pertinent background information upon which to base their review and approval. The nature of the change is identified in the document or appropriate attachments.

Documented procedures are established to address personnel training following document changes.

Original data used to support the requirements of the Quality Management System are collected and recorded using appropriate methods. Data is retained in accordance with documented procedures.

All documents required by the Quality Management System are secured and retained electronically in appropriate locations and conditions to ensure their retrieval for the length of time required by regulations, standards, and other applicable requirements and to ensure prevention from damage or deterioration. The procedures to secure and control original records (including electronic records and data) include appropriate storage methods, access restriction, validation requirements, and information back-up.

The following is the document structure utilized to ensure the effective planning, operation and control of established processes:

| Level 1 | Quality Manual |
| --- | --- |
| Level 2 Documents | Quality Procedures |
| Level 3 Documents | Inspection Procedures, Manufacturing Routers, Manufacturing Procedures, Product Drawings, Equipment Drawing/Specifications, Tooling Drawings, etc. |
| Level 4 QMS Records | File Documents, Forms, Lists, Test Reports, Device History Records, Complaints, Management Reviews, Calibration Records, Internal Audits, Supplier Records, Risk Management Files, Design History Files, etc. |

## Control of Quality Records

Quality records are appropriately controlled in accordance with the requirement set forth below and QM.SLQ001 Document Control. Quality records are established and maintained to demonstrate the effective planning, operation, and control of the Quality Management System processes. Required records for QMS processes are defined in each applicable standard operating procedure.

Quality records shall be legible and readily identifiable.

Quality records are stored and maintained in such a way that they are readily retrievable and prevented from damage or deterioration. Storage facilities may be used for off-site storage.

Retention times of quality records have been established and recorded per applicable regulations.

Device history records are maintained for each product lot to provide traceability and identify the quantity manufactured and released.

Confidential health information is safeguarded in accordance with QM.SLQ028 Protection of Confidential Patient Information per applicable regulations.

# Management Responsibility

## Management Commitment

Senior Management is committed to communicate the importance of meeting customer as well as statutory and regulatory requirements.

The Management Representative is responsible for implementing this commitment by promoting awareness of customer and regulatory requirements throughout SILQ. The Management Representative is appointed by the CEO.

Senior Management defines the purpose, directions, and objectives for the quality system. The objectives are documented and communicated in the form of a Quality Policy and quality objectives.

Senior Management periodically reviews the quality system in Management Review meetings to ensure its continuing suitability, adequacy, and effectiveness. The review evaluates current status and performance of the quality system and initiates actions to further improve the system.

Senior Management is committed to providing resources necessary for establishing, implementing, and improving the quality system.

## Customer Focus

Senior Management ensures that user/customer requirements and applicable regulatory requirements are identified and met through design control activities. The organization quality policy focuses on meeting or exceeding customer expectations. Customer needs are identified at the beginning of the design and development process in accordance with QM.SLQ004 Design Control Program. Information from multiple resources is obtained and assessed to determine customer needs. The assessment may include data from various sources including:

- Identification of the organization strategic plan/business focus;

- Information regarding the disease conditions and the treatment challenges;

- Clinical information from published literature;

- Existing unmet customer needs;

- Feedback from customers (e.g., users, patients); and

- Competitive products.

## Quality Policy

Senior Management has established a quality policy that:

- Is appropriate to the purpose of the organization;

- Is customer focused; and

- Includes a commitment to comply with requirements and maintain the effectiveness of the Quality Management System.

The Quality Policy provides a framework for establishing and reviewing quality objectives. It is communicated and understood within the organization. The Quality Policy is periodically reviewed at Management Review meetings to ensure its continual relevance and suitability. It is documented in QM.SLQ035 and provided below:

Quality Policy

Excellent patient care is the focus of all SILQ activities. We endeavor to meet or exceed our customers’ satisfaction and expectations in providing safe and effective surface treated medical devices of the highest quality and that meets all applicable regulatory requirements.

Each SILQ employee is personally committed to achieving this goal and continually improving our ability to maintain customer focus by operating a highly effective quality system that meets the requirements of applicable global standards and regulatory requirements for medical devices.

## Quality Planning Process

Senior Management is responsible for defining and planning the tasks and resources needed to attain quality objectives and establish and maintain an effective Quality Management System. During Management Reviews, quality objectives, including those needed to meet applicable regulatory requirements and requirements for product, are established at relevant functions and levels within the organization. The quality objectives are measurable and consistent with the quality policy. The purpose of the quality system is to:

- Achieve the Quality Policy;

- Ensure and demonstrate the ability to provide products and related services that consistently meet customer requirements and applicable regulatory requirements;

- Ensure high level of customer satisfaction;

- Facilitate continual improvement of products;

- Document manufacturing and quality control processes;

- Maintain effectiveness, suitability, and adequacy of the quality system documentation; and

- Comply with required standards and applicable regulatory requirements for quality systems.

Changes to the quality system may be planned within Management Reviews, CAPA activities, customer feedback, or through responses to internal or external audits. These changes may be in response to changing circumstances, such as product, process, capacity, or other operational or organizational changes; or to improve the effectiveness, suitability, and adequacy of the quality system. Quality Planning is carried out in accordance with QM.SLQ025 Quality Planning.

## Responsibility, Authority and Communication

7.5.1 The general organizational structure, shown in Figure 2, provides the appropriate responsibility, authority, and interrelation of all personnel who manage, perform, and assess work affecting quality by providing the independence and authority necessary to perform the required tasks. SILQ’s organizational structure is detailed in the Organization Chart (QM.SLQ034).

Figure 2: General Organizational Structure

### 7.5.2 Management Representative

The V.P. of Regulatory/ Quality Assurance will serve as the Management Representative and has authority over and responsibility for:

- Ensuring that all processes needed for the Quality Management System are documented (When a requirement is required to be “documented”, it is also required to be established, implemented, and maintained);

- Reporting to Senior Management on the performance of the Quality Management System and any need for improvement;

- Ensuring the promotion of awareness of regulatory and customer requirements throughout the organization;

- Ensuring that appropriate communication processes are established within the organization;

- Ensuring that communication takes place regarding the effectiveness,  suitability and adequacy of the Quality Management System; and

- Maintaining a liaison with external regulatory agencies, notified and certification bodies on matters (e.g., audits) relating to the Quality Management System.

The Management Representative is responsible for monitoring the Quality Management System and its processes. Significant modifications to the Quality Management System, organizational structure and/or product technical documentation are reported in writing to the appropriate notified body, certification body and/or regulatory agency per the applicable procedures. Where operational events and quality policy conflict, the CEO (or delegate) has ultimate authority and responsibility to resolve the conflict. Records of such instances are retained.

### 7.5.3 Internal Communication

SILQ ensures effective communication between its various levels and functions regarding the processes of the quality management system and their effectiveness, suitability, and adequacy through periodic company meetings and Management Reviews. It is also the responsibility of each Manager or department head to ensure all relevant information is passed on to their functions.

## Management Review Process

SILQ schedules management review meetings on a regular basis (annually at a minimum). Reviews are performed to assess the status of the QMS against the set quality objectives and also to provide an opportunity to review these goals and objectives along with the continued suitability of the quality policy. An evaluation of the overall suitability and effectiveness, suitability, and adequacy of the QMS will be part of this review.

Quality performance indicators (input data) are established and reviewed as necessary. These indicators shall include (but are not limited to):

- Follow-up actions from previous management reviews;

- Status of quality policy and objectives;

- Monitoring and measurement of processes;

- Monitoring and measurement of product;

- Non-conformance reports;

- Post-market surveillance data; Customer complaints/feedback;

- Reporting to regulatory authorities;

- Internal audit; regulatory agency inspection and certification body audit results;

- Corrective and Preventive actions;

- Statistical techniques;

- Supplier performance;

- Training performance;

- Changes that could affect the Quality Management System;

- New or revised standards and regulatory requirements; and

- Recommendations for improvement.

During management reviews, input data is reviewed and analyzed to determine if improvements are required. The output from management review includes:

- Decisions and actions related to improvements needed to maintain the effectiveness, suitability, and adequacy of the Quality Management System and its processes;

- Improvement of product related to customer requirements;

- Changes needed to respond to applicable new or revised regulatory requirements;

- Resource needs; and

- Objectives (goals) established.

The dates and results of the management reviews are documented and maintained by the Management Representative. Corrective measures are documented and tracked to ensure closure of activities.

Management Review activities and responsibilities are further defined in QM.SLQ018 Management Review.

# Resource Management

## Provision of Resources

Senior Management has identified and provided the appropriate resources:

- To meet the established quality policy and objectives;

- To implement the Quality Management System requirements and maintain its effectiveness, suitability, and adequacy; and

- To meet applicable regulatory and user/customer needs.

## Human Resources

Senior Management strives to ensure personnel performing work affecting product quality are competent on the basis of appropriate education, training, skills, and experience.

Personnel performing work affecting product quality are selected based on required competencies, appropriate education, training (internal and/or external), skills, and experience.

Management identifies training needs and provides training for all personnel performing tasks affecting quality as required.  Training and effectiveness of training is conducted in accordance with established procedures by qualified individuals.  Appropriate records of education, experience, training, and skills are maintained.

The manufacture and testing of SILQ products may be contracted out to CMOs. These requirements and responsibilities are defined in the Quality Agreements established between SILQ and the CMOs and compliance is verified through the supplier monitoring process.

## Training Process

- Employee training is performed in accordance with QM.SLQ003 Employee Training. Responsibilities and requirements for training are established which includes the identification of training needs of all personnel performing activities affecting quality.

- Training is conducted to ensure that SILQ employees and representatives are adequately trained to perform their job assignments. Training programs are comprised of sessions conducted within the organization and appropriate outside educational programs, meetings, conferences, etc.

- Effectiveness of training may be evaluated and is proportionate to the risk associated with the work for which the training is being provided.

- Personnel performing specific tasks are qualified on the basis of appropriate education, training, and/or experience.

- Training requirements address both employees and contract personnel and include new employee orientation, job specific training, training in company-related issues, special qualification requirements, and training and/or supervision of personnel working under special environmental conditions or performing special processes.

- Field personnel are trained in the applicable complaint reporting procedures.

## Infrastructure

Senior Management identifies CMOs that can provide the appropriate infrastructure to meet user/customer, product, and regulatory requirements. This includes facilities, process equipment, and support services.

Requirements for maintenance activities and their frequency are established and documented. Records of maintenance activities are maintained.

## Work Environment and Contamination Control – Requirements for:

- Health, cleanliness, and clothing of personnel contacting product or the work environment are established;

- Work environment conditions and the procedure to monitor and control these conditions are established;

- Personnel who are required to work temporarily under special environmental conditions are trained or supervised by trained personnel; and

- Control of contaminated or potentially contaminated product to prevent contamination of other product, the work environment, or personnel are established.

## CMOs that provide product or services are selected based on their ability to fulfill applicable requirements for work environment and contamination control.

# Product Realization

## Planning Process

Processes are planned and developed for product realization consistent with the requirements of the other processes of the quality management system. These activities are conducted in accordance with QM.SLQ025 Quality Planning and QM.SLQ004 Design Control Program.

The following inputs and requirements are evaluated and determined in planning product realization, as appropriate:

- User/customer needs and requirements, including requirements not stated by the user/customer, but necessary for the specified or intended use;

- Statutory and regulatory requirements related to the product;

- Quality objectives and requirements for the product;

- The need to establish/provide specifications, processes, documents, resources, and facilities specific to the product;

- Required verification, validation, monitoring, inspection and test, handling, storage, distribution, and traceability activities specific to the product together with the criteria for product acceptance; and

- Records needed to provide evidence that the realization processes and resulting product meet specified requirements.

Documented requirements for risk management are established throughout the product realization process. Risk management activities are documented and records are maintained. Risk management activities and responsibilities are further defined in QM.SLQ012 Risk Management and QM.SLQ013 Risk Analysis.

## Customer Related Processes

SILQ determines the requirements related to product, which include:

- Customer requirement, including the requirements for delivery and post-delivery activities;

- Requirements not stated by the customer but necessary for specified or intended use, as known;

- \

- Any additional requirements determined by the organization.

SILQ reviews requirements related to product prior to the organization’s commitment to supply product to the customer and ensures that:

- Requirements are defined and documented;

- Any requirement differences between SILQ and its customers are resolved;

- Applicable regulatory requirements are met;

- Any user training identified is available or planned to be available;

- Affected organizational units are notified;

- The organization has the ability to meet the defined requirements; and

- Appropriate records of the review results and any actions arising from the review are maintained.

These activities are further defined in QM.SLQ036 Sales Order and QM.SLQ004 Design Control Program.

If product requirements change, relevant documents shall be amended and relevant personnel made aware of the changed requirement.

Any amendments to the contract / sales order is documented and communicated to the customer.

Sales orders are reviewed to ensure that applicable regulatory approval, language, and traceability requirements have been met prior to distribution/shipment.

Documented procedures have been established to ensure that the appropriate product information (e.g., brochures, instructions for use), feedback mechanisms (e.g., complaint handling), and safety information (e.g., advisory notices) are provided to the customer. Feedback mechanisms and safety information provided to the customer are further defined in QM.SLQ021 Product Complaint System and QM.SLQ030 Advisory Notices and Recalls, respectively.

Communication with customers relating to product information, inquiries, contracts, or order handling, including amendments, customer feedback, customer complaints, and advisory notices, are essential to the success of the business.  Appropriate customer communication requirements have been established.

# Design and Development Process

All design and development activities are planned and controlled in accordance with QM.SLQ004 Design Control Program. This procedure specifies the development of project plans that identify the responsibilities and resources for each design and development activity. The project plans are reviewed and updated as appropriate throughout the design and development process. This includes the design and development stages; the review, verification, validation, and design transfer activities at each stage; and the responsible individuals/departments and authority levels. Interfaces between different groups involved in the design and development process are managed to ensure effective communication and clear assignment of responsibility.

A Design History File (DHF) is established for each type of device. The file contains or references the records necessary to demonstrate that the design was developed in accordance with the approved project plan. The DHF includes, but not limited to:

- General description of the device, intended use, and labeling;

- Specifications for product;

- Specifications or processes for manufacturing, packaging, storage, handling, and distribution;

- Procedures for measuring and monitoring;

- As appropriate, requirements for installation; and

- As appropriate, procedures for servicing.

Design and development requirements (inputs) are defined and reviewed for adequacy and conformance to applicable standards, regulatory requirements and customer requirements. This includes:

- User/customer needs;

- Applicable regulatory requirements and standards;

- Functional, performance, usability and safety requirements;

- Information from previous similar designs; and

- Output(s) from risk management.

Requirements are complete, unambiguous, able to be verified or validated, and not in conflict with each other. These inputs are reviewed for adequacy, approved, and documented. Design input activities and requirements are further defined in QM.SLQ006 Design Input.

Design and development outputs are provided in a form that enables verification against the input requirements, acceptance criteria and applicable regulatory requirements prior to release. Design and development outputs provide the necessary information for purchasing and manufacturing, and specify the product characteristics essential for safe and proper use. Records of design and development outputs (e.g., specifications, engineering drawings, manufacturing procedures, etc.) are maintained. Design output activities and requirements are further defined in QM.SLQ007 Design Output.

Design outputs are the basis of the Device Master Record (DMR), also referred to as the Medical Device File. A DMR is established for every product at SILQ. The contents of the DMR/Medical Device File include, but are not limited to:

- A general description of the medical device, intended use, and labeling;

- Specifications for the device;

- Specifications or procedures for manufacturing, packaging, storage, handling, and distribution;

- Procedures for measuring and monitoring of the device;

- As appropriate, requirements for installation; and

- As appropriate, procedures for servicing.

Device Master Record requirements are further defined in QM.SLQ048 Device Master Record.

Design and development verification by designated, trained, and competent personnel is a planned and documented activity. Design and development verification is conducted to establish that the outputs meet the input requirements. Design and development validation is performed following successful verification under defined operating conditions on initial production units, lots, batches or their equivalents. Validation is planned and conducted to ensure that products conform to defined user needs and intended use(s) and include testing of products under actual or simulated use conditions.  The results of design and development verification and validation, including identification of the design, method(s), acceptance criteria, statistical techniques with rationale for sample size as appropriate, the date(s), and the individual(s) performing the verification and validation, are documented in the Design History File (DHF). Design verification and validation activities and requirements are further defined in QM.SLQ009 Design V&V.

Formal design and development reviews are planned and conducted at appropriate stages as defined in QM.SLQ008 Design Review and the project plan:

- To evaluate the ability of the design and development results to meet requirements;

- To identify any problems; and

- Propose necessary actions.

Participants in design and development reviews include representatives of functions involved with the design and development stage being reviewed. Records of the results of design and development reviews and any necessary actions are maintained and include the identification of the design under review, participants involved, and the date of the review.

The device design is translated into production specifications that are suitable for manufacturing in accordance with QM.SLQ010 Design Transfer. Results and conclusions of the design transfer are maintained.

When appropriate, clinical evaluations shall be conducted in compliance with applicable regulatory requirements. Clinical records are maintained as part of the design validation records.

A risk management file shall be prepared for each project as defined in QM.SLQ012 Risk Management. The file is updated as appropriate throughout product development from initial design to market release and while the product remains in distribution.

All design and development changes are performed in accordance with QM.SLQ001 Document Control, where the significance of the change to function, performance, usability, safety, and applicable regulatory requirements for the product and its intended use are evaluated. Prior to implementation, changes are identified, reviewed, verified, validated (as appropriate), and approved. Review of changes includes evaluation of the effect of the changes on constituent parts and product in process or already distributed, inputs or outputs of risk management, and product realization processes. Records of changes, their review and any necessary actions are maintained.

# Purchasing

## Purchasing Process

SILQ maintains procedures to ensure that purchased product conforms to specified requirements. Manufacturing and Quality Assurance evaluate and select suppliers based on established criteria. The level of control of suppliers is dependent upon the type of product or service provided. The monitoring and re-evaluation of supplier performance is planned. Results of the monitoring provides input into the supplier re-evaluation process. An Approved Supplier List and the records of supplier evaluation, monitoring, and re-evaluation are maintained.

Supplier evaluation, selection, monitoring and re-evaluation activities and requirements are further defined in QM.SLQ015 Supplier QA.

## Purchasing Information

Purchase orders and contracts for manufacturing materials and services may only be placed with suppliers listed on the Approved Supplier List. Purchase orders and contracts appropriately describe ordered products and services, including product identification, specification requirements for acceptance, and quality management system requirements.

Purchasing activities and requirements are further defined in QM.SLQ020 Purchasing Controls.

## Verification of Purchased Product

Purchased products are received in accordance with QM.SLQ045 Receiving and verified prior to use. Upon receipt, trained personnel perform inspections to verify that purchased product fulfills the specified requirements in accordance with incoming acceptance criteria.

Requirements for verification of purchased product are further defined in QM.SLQ039 Receiving Inspection.

Products or services that do not meet the expectations of SILQ will be treated as nonconforming.  Control of nonconforming product activities and requirements are defined in QM.SLQ040 Nonconforming Materials

# Production Process

## General Requirements

SILQ  produces the polymeric suspension used to apply the surface treatment internally and outsources finished device production activities (e.g., application of the surface treatment to the device, packaging, labeling, and sterilization) to applicable CMOs. Silq also performs quality control tests relevant to the surface treatment on all products and approves appropriate sections of the Device History Record prior to final release by applicable CMOs. All production-related activities performed in-house and by CMOs are carried out under controlled conditions. This includes:

- The availability of information and specifications that describe the characteristics of the product;

- The availability of documented procedures and requirements, and work instructions;

- The use of qualified personnel and suitable equipment/facilities;

- The availability and use of monitoring and measuring devices;

- The implementation of defined operations for labeling and packaging; and

- The implementation of release, storage, handling, and delivery requirements.

## Production Control

Device manufacturing and provision of associated services are carried out under controlled conditions as defined in the Device Master Record (DMR), also referred to as the Medical Device File. The controlled conditions include the control of, as applicable:

- Infrastructure: The requirements for work environment including, but not limited to, environmental and contamination controls are defined to ensure products produced meet specified requirements.

- Product process instruction: Instructions specifying product characteristics and manufacturing processes/servicing are communicated to operators in the form of manufacturing travelers, drawings, specifications, process instructions, and tooling.

- Equipment: Process equipment and tooling are selected on the basis of their capability to produce products that meet specified requirements.

- Cleanliness: Requirements for cleanliness of product are specified as appropriate.

- Product release and delivery: Products are released for shipping after all specified manufacturing, inspection and test activities have been satisfactorily completed and conformity of the product and of the associated Device History Record have been verified.

- Labeling and packaging: Device labeling and packaging protect the device from damage during the conditions of processing, storage, handling and distribution.

- Device History Record (DHR): A record is maintained for each lot of product manufactured. This record includes quantity manufactured and released for delivery. The DHR for each product or lot is approved prior to product release for delivery.

QM.SLQ049 Workstation Practices outlines good manufacturing practices in order to maintain controlled conditions in production environments.

Product is released by SILQ in accordance with QM.SLQ029 Device History Record (DHR) Review and Approval. Finished product is released by the CMO in accordance with established batch release procedures. Records of released product batches are maintained by SILQ.

## Production Process Validation is performed by SILQ in accordance with QM.SLQ047 Process Validation.  Process validations may also be performed by the applicable CMO. This requirement is reinforced between SILQ and the CMO using quality agreements in accordance with QM.SLQ020 Purchasing Controls and compliance is verified through the supplier monitoring process. General requirements for process validation includes:

- Documented procedures are established for the validation of production processes, including sterilization processes, where the resulting output cannot be verified by inspection, monitoring, or measurement. This included processes where deficiencies become apparent only after the product is in use.

- Validation is conducted and documented to demonstrate the ability of production processes to achieve planned results.

- Validation activities include defined criteria for review and approval of identified production processes, approval of equipment, qualification of personnel, use of specific methods and procedures, and revalidation for product/process changes.

- Computer software for production and inspection processes that affect the ability of the product to conform to specified requirements shall be validated.

## Identification and Traceability

- Identification

Raw materials, components, and finished devices are identified upon receipt and throughout storage. Identification can take the form of labels on the product or bins or other marking on the product during storage and manufacturing as well as package labels during delivery. The identification contains sufficient unique identifying information including product or part number.

The specific requirements for product identification or finished devices are included in the Device History Record.

If required by applicable regulatory requirements, finished devices are assigned a unique device identification.

When returned product is received, it is clearly identified and segregated from other product.

- Traceability

Traceability is maintained when specified either to facilitate corrective actions or manage risk. Documented procedures are provided for traceability requirements.

Traceability is based on identifying the product with unique control numbers. Traceability on product is identified through the inventory management system. Records of product traceability are maintained in the associated DHR.

- Acceptance status identification

Throughout production and storage, products are identified with respect to their acceptance status to avoid unintended use. Acceptance status identification can take the form of labels on the products, labeling of racks or areas within a manufacturing or warehouse location and/or electronic control of product via the inventory management system.

Identification and traceability activities and requirements are further defined in QM.SLQ019 Identification and Traceability.

## Control of Customer-Supplied Product / Customer Property

No physical customer property will be used or incorporated into the product. Patient data may be obtained for product complaint and failure investigation purposes and is controlled in accordance with QM.SLQ028 Protection of Confidential Patient Information.

## Handling, Storage, Packaging, Preservation and Delivery

- Handling

All products, at all stages of manufacturing and delivery, including returns, are handled in a manner such as to prevent damage, deterioration or contamination to the product packaging, personnel, or environment.

- Storage and Preservation

Designated storage areas are established to prevent damage, deterioration or contamination of the product pending delivery. This preservation includes identification, handling, packaging, storage and protection. Preservation shall also apply to the component level.

Finished product and components are used subject to FIFO (First-In-First-Out) practices. Prior to shipping, product shall be assessed to ensure expiration dates, if applicable, have not been exceeded.

Product known to be subject to deterioration over time shall be identified and assessed at defined intervals. If applicable, product with a limited shelf life is identified as such with an expiration date on receipt into stock.

Components shall be maintained according to any special storage conditions indicated.

- Delivery

Delivery services to be used are the same as is customary for medical device transportation (e.g., nationally recognized parcel delivery services).

Shipping documentation identifies the name and address of the customer, identification and quantity of product shipped, date shipped, and any lot or serial numbers used. Shipping activities and requirements are further defined in QM.SLQ046 Shipping.

In the manufacture and testing of SILQ products, activities that are contracted out to CMOs are defined and documented as requirements in the quality agreements established between SILQ and the CMO and compliance is verified through the supplier monitoring process.

## Control of Monitoring, Inspection, Measurement and Test Equipment

Appropriate measuring and monitoring devices are selected to ensure that measurement capability is consistent with measurement requirements. Equipment used for ensuring and verifying product conformity are calibrated and maintained in accordance with established requirements. The calibration status of equipment shall be readily evident. The requirements for the calibration and preventive maintenance system are further defined in QM.SLQ047 Calibration and Preventive Maintenance.

In the manufacture and testing of SILQ products, activities are contracted out to CMOs. Control of monitoring and measuring equipment used by CMOs is documented in the quality agreements, as applicable, established between SILQ and the CMO and compliance is verified through the supplier monitoring process.

# Measurement and Analysis Process

## General Requirements

Monitoring, measurement, analysis, and improvement processes including the determination of applicable methods (e.g., statistical methods) and the extent of their uses are planned and implemented to:

- Demonstrate conformity of the product;

- Ensure conformity of the Quality Management System; and

- Maintain the effectiveness, suitability, and adequacy of the Quality Management System.

## Feedback

SILQ monitors information on the performance of products and the Quality Management System to meet customer requirements in accordance with QM.SLQ021 Product Complaint System and QM.SLQ033 Post Market Surveillance. This includes a feedback system to provide early warning of quality problems and for input into the corrective and preventive action process and risk management process.

SILQ notifies the applicable regulatory authorities of complaints that meet specified reporting criteria of adverse events or issuance of advisory notices and recalls in accordance with QM.SLQ022 Medical Device Reporting and QM.SLQ030 Advisory Notices and Recalls.

## Internal Audits

QM.SLQ017 Internal Quality Audits defines the responsibilities and requirements for conducting internal quality assessments at planned intervals to determine whether the Quality Management System:

- Conforms to plans, applicable regulatory requirements and standards; and

- Is effectively implemented and maintained.

Internal audits are conducted in accordance with a planned program, taking into consideration the status and importance of the processes and areas to be audited, as well as the results of previous audits.

Appropriate corrective actions are assigned to personnel responsible for the areas where nonconforming processes and/or practices are identified.

Records of the internal audits and their results, including identification of the processes and areas audited and the conclusions, are maintained. Internal audit results are reviewed at Management Review meetings as defined in QM.SLQ018 Management Review.

## Monitoring and Measurement of Processes

Key and supporting processes are monitored by a variety of approaches and techniques, as appropriate for a particular process and its importance. These include:

- Conducting internal audits of the quality system; and

- Measuring product conformity and monitoring other quality assurance performance data and trends.

## Monitoring and Measurement of Product

The monitoring and measurement program for products is defined in the product’s Device Master Record and carried out by the applicable CMO. Statistical techniques are applied in accordance with QM.SLQ011 Statistical Techniques, where applicable, to control and verify the acceptability of product and process.

Statistically valid sampling plans are employed. When changes occur to the product or process, the effect on the validity to the plan should be considered.

Purchased product is subjected to incoming acceptance activities. In-process inspections are in the form of operator-based inspections. Acceptance activities are established and recorded to ensure product conformity. Finished products are subjected to a final inspection.

Products are released by Quality Assurance of applicable CMO for distribution after all specified activities have been satisfactorily completed and required documentation of the product has been verified.

## Control of Nonconforming Product is performed by SILQ and the applicable CMO, accordingly.

- Documented procedures have been established in the quality system of SILQ and applicable CMOs to define responsibilities and requirements to ensure product that does not conform to specified requirements is prevented from inadvertent use or distribution.

- Control requirements are defined for the identification, documentation, evaluation, classification, segregation, disposition, and prevention of recurrence of non-conforming product and for the notification of the functions concerned.

- Responsibility for review and authority for the disposition of nonconforming product is defined within SILQ and the CMO. SILQ Quality Assurance is responsible for the final disposition of nonconforming product.

- SILQ shall provide documented justification for accepting nonconforming product by concession and ensure that approval is obtained and applicable regulatory requirements are met. The justification should include a risk assessment and an evaluation of the concession on product safety and performance.

- Nonconforming product is segregated and scrapped or reworked/repaired per approved work instructions and re-inspected in accordance with written procedures to ensure that it meets the currently approved specifications and regulatory requirements. Rework instructions shall undergo the same review and approval as the original procedure. Evaluation of the potential adverse effect of the rework on the product shall be documented.

- Records are maintained of the evaluation, disposition, and justification for use of non-conforming product. These records become part of the Device History Record.

Control of nonconforming product activities and requirements are further defined in QM.SLQ040 Nonconforming Materials.

## Analysis of Data

Review and analysis of appropriate data is conducted to demonstrate the suitability, adequacy, and effectiveness of key and support processes, and to identify opportunities for improvement. Corrective and Preventive Actions (QM.SLQ016) or Management Reviews (QM.SLQ018) are used to drive and monitor improvement efforts.

Performance data provides information relating to:

- Customer feedback

- Conformity to product requirements

- Characteristics and trends of processes and product

- Internal audit results

- Supplier quality performance

- Corrective and preventive actions

Statistically valid methods are employed to determine and evaluate data as defined in QM.SLQ011 Statistical Techniques.

# Improvement Process

The QM.SLQ016 Corrective and Preventive Actions procedure defines responsibilities and activities to ensure the causes of nonconforming product, processes or systems are investigated and corrective action is implemented to prevent recurrence.

Changes necessary to ensure and maintain the continued suitability, adequacy, and effectiveness of the Quality Management System are identified and implemented through the use of quality policies, quality objectives, audit results, analysis of data, corrective and preventive actions, and management review.

Analysis of incoming raw materials and components, work operations, quality records, customer complaints, audit reports, and other sources of quality data is conducted to detect existing and potential causes of non-conforming product, processes, and systems.

Investigation of the root cause of the non-conforming product, processes, and systems and the identification of actions to correct and prevent recurrence is conducted and documented.

## Corrective Action

Actions are taken to eliminate the cause of nonconformities in order to prevent recurrence. Corrective actions are appropriate to the risks presented by the encountered issue. The QM.SLQ016 Corrective and Preventive Actions procedure has been established to define the requirements for:

- Reviewing nonconformities including customer complaints;

- Determining the cause of the nonconformities;

- Evaluating the need for action to ensure the nonconformities do not recur;

- Planning and documenting action needed and implementing such action, including, as appropriate, updating documentation;

- Verifying the corrective action does not adversely affect the ability to meet applicable regulatory requirements or the safety and performance of the device; and

- Reviewing the effectiveness of corrective actions taken.

## Preventive Action

Preventive actions are implemented to eliminate causes of potential nonconformities in order to prevent their occurrence. SILQ has established QM.SLQ016 Corrective and Preventive Actions procedure for reviewing potential nonconformities, determining causes, evaluating necessary actions, and implementing the preventive action. This procedure includes process, responsibility, authority, and documentation requirements.

Changes to documented procedures resulting from preventive action are reviewed, recorded, approved, and implemented. The effectiveness of preventive actions is verified.

# Appendix 1: ISO 13485 Clause Correlations with QMS Documentation

| ISO 13485 Clause | Corresponding QMS Documentation |
| --- | --- |
| 4 Quality Management System |  |
| 4.1 General Requirements | QM.SLQ027 Quality Manual |
| 4.2 Documentation Requirements | QM.SLQ027 Quality Manual QM.SLQ001 Document Control QM.SLQ002 Good Documentation Practices QM.SLQ028 Protection of Confidential Patient Information |
| 5 Management Responsibility | QM.SLQ027 Quality Manual |
| 5.1 Management Commitment | QM.SLQ018 Management Review |
| 5.2 Customer Focus | QM.SLQ004 Design Control Program |
| 5.3 Quality Policy | QM.SLQ018 Management Review QM.SLQ027 Quality Manual QM.SLQ035 Quality Policy |
| 5.4 Planning | QM.SLQ027 Quality Manual QM.SLQ018 Management Review QM.SLQ025 Quality Planning |
| 5.5 Responsibility, Authority and Communication | QM.SLQ027 Quality Manual |
| 5.6 Management Review | QM.SLQ018 Management Review |
| 6 Resource Management |  |
| 6.1 Provision of Resources | QM.SLQ027 Quality Manual |
| 6.2 Human Resources | QM.SLQ027 Quality Manual QM.SLQ003 Employee Training |
| 6.3 Infrastructure | QM.SLQ027 Quality Manual |
| 6.4 Work Environment and Contamination Control | QM.SLQ027 Quality Manual |
| 7 Product Realization |  |
| 7.1 Planning of Product Realization | QM.SLQ004 Design Control Program QM.SLQ012 Risk Management QM.SLQ013 Risk Analysis QM.SLQ025 Quality Planning |
| 7.2 Customer-Related Processes | QM.SLQ004 Design Control Program QM.SLQ020 Purchasing Controls QM.SLQ021 Product Complaint System QM.SLQ030 Advisory Notices and Recalls |
| 7.3 Design and Development | QM.SLQ001 Document Control QM.SLQ004 Design Control Program QM.SLQ006 Design Input QM.SLQ007 Design Output QM.SLQ008 Design Review QM.SLQ009 Design V&V QM.SLQ010 Design Transfer QM.SLQ012 Risk Management |
| 7.4 Purchasing | QM.SLQ015 Supplier QA QM.SLQ020 Purchasing Controls QM.SLQ039 Receiving Inspection |
| 7.5 Production and Service Provision | QM.SLQ019 Identification and Traceability QM.SLQ020 Purchasing Controls QM.SLQ027 Quality Manual QM.SLQ029 DHR Review and Approval QM.SLQ032 Software Validation QM.SLQ043 Work Order QM.SLQ045 Receiving QM.SLQ046 Shipping QM.SLQ048 Device Master Record QM.SLQ049 Workstation Practices |
| 7.6 Control of Monitoring and Measuring Equipment | QM.SLQ027 Quality Manual QM.SLQ047 Calibration and Preventive Maintenance |
| 8 Measurement, Analysis and Improvement |  |
| 8.1 General | QM.SLQ027 Quality Manual |
| 8.2 Monitoring and Measurement | QM.SLQ011 Statistical Techniques QM.SLQ017 Internal Quality Audits QM.SLQ018 Management Review QM.SLQ021 Product Complaint System QM.SLQ027 Quality Manual QM.SLQ022 Medical Device Reporting QM.SLQ030 Advisory Notices and Recalls QM.SLQ033 Post Market Surveillance |
| 8.3 Control of Nonconforming Product | QM.SLQ027 Quality Manual QM.SLQ040 Nonconforming Materials |
| 8.4 Analysis of Data | QM.SLQ011 Statistical Techniques QM.SLQ016 Corrective and Preventive Actions QM.SLQ018 Management Review |
| 8.5 Improvement | QM.SLQ016 Corrective and Preventive Actions QM.SLQ017 Internal Quality Audits QM.SLQ018 Management Review QM.SLQ027 Quality Manual |