Device History Record (DHR) Review and Approval

Standard Operating Procedure

# Purpose

## The purpose of this standard operating procedure is to define the review and approval process for the release of Device History Records (DHR) to ensure product is built and processed in accordance with their approved Device Master Record (DMR).

# Scope

## The following procedure applies to all products manufactured by SILQ.

## SILQ may utilize subcontractors to perform manufacturing activities which may encompass some of the DHR review and approval activities defined in this procedure to release product. Where specific requirements are stated within this document or other referenced SILQ procedures, they must be complied with by SILQ and any subcontractors engaged in any specific project.

# Documents

## Reference Documents

### QM.SLQ001		Document Control SOP

## Appendices / Associated Forms / Templates

### FM1-QM.SLQ029	DHR Review and Approval Form

# Definitions

## Abbreviations / Acronyms

### CMO:  Contract Manufacturing Organization

### DHR:  Device History Record

### DMR:  Device Master Record

## General Definitions

### Device History Record:  Compilation of records containing the production history of a finished assembly, product or device.

### Device Master Record:  The compilation of all the instructions, drawings and other documents required to produce a completed product.

### FileHold:  Software based document management system used to electronically store controlled documents.

# Responsibilities

## It is the responsibility of Quality Assurance personnel to perform the review, verify all required documentation, and approve the DHR for release of product.

## It is the responsibility of Quality Assurance management to ensure all aspects of DHR review and approval are performed in compliance with this procedure.

## It is the responsibility of Manufacturing/Operations (or CMO) to complete and collect all production documentation as stipulated in the DMR for inclusion in the DHR.

## It is the responsibility of shipping personnel to verify the DHR is complete and signed for release by Quality Assurance prior to shipping the finished product or device(s) for clinical or commercial distribution.

# Procedure: DHR Overview

## A DHR is to be created for all lots of product/devices that has been manufactured by or for SILQ that is to be released for internal use, investigational use, or commercial distribution.

## Each DHR generated for product/device(s) are to be reviewed and verified to be complete, contain correct and accurate content prior to approval for release.

## A DHR is to include at minimum, but is not limited to, the following:

### Manufacturing travelers or assembly records as defined in the DMR that include:

#### Part number

#### Product Lot/Batch number

#### Traceability to raw materials, components, and/or assemblies Lot/Batch numbers used in the build of finished product

#### List or record of completed processing steps and/or conditions

### Applicable Inspection/Testing forms as defined in the DMR

### Copies of all appropriate labeling used to identify the final product on the primary package or product itself

### Any applicable non-conforming material records and documentation on any rework

### Any applicable documentation in regards to any deviations in the specification or manufacturing procedure as defined by the DMR

### A completed DHR Review and Approval record (required for product release)

# Procedure: DHR Review and Approval

## Upon completion of all manufacturing and processing steps, all applicable records are to be compiled, verified to be complete by manufacturing, and submitted for review to the designated Quality Assurance Reviewer.

## The Quality Assurance reviewer is to review the submitted DHR and record the results in the DHR Review and Approval Form FM1-QM.SLQ029.

## The reviewer is to perform the following:

### Verify that all information recorded in the DHR complies with the following good documentation practices:

#### All initials, signatures, and dates are recorded and accurate.

#### All data entry fields or signature areas that are not required have been marked with “N/A” and all applicable sections have a straight diagonal line through them. No blank areas or spaces are to be left open in any documents or forms.

#### All corrections made have a single line through the incorrect information, have the corrected information recorded, and changes are initialed and dated.

### Verify all information listed in Section 6, as applicable, is present, complete, and correct in the DHR.

### Verify that lot/batch numbers are correct and have been recorded in all appropriate locations of the DHR.

### Verify that any non-conforming materials used in the manufacturing or processing of the finished product has been correctly documented. In addition, verify all non-conforming material reports are complete and their disposition approved for use in the final product. Document all applicable non-conforming material report numbers on the DHR Review and Approval Form. Ensure copies of the non-conforming material reports have been included with the DHR.

### Verify that any rework to the finished product has been correctly identified and properly documented.

### Verify any deviations associated with the production of the finished product has been properly documented and approved as defined in the Document Control process defined in SOP QM.001. Summarize the deviations on the DHR Review and Approval Form. Ensure copies of the deviation have been included with the DHR.

### Verify that all final inspection or testing has been completed and is in compliance with the defined acceptance criteria. Ensure all inspection or test records as stipulated in the DMR have been included in the DHR.

### Verify that copies of all product labeling have been included with the DHR. Ensure that all product labeling is complete and accurate.

# Procedure: DHR Approval

## Once the DHR and all associated documentation have been verified, the Quality Assurance reviewer is to finalize the DHR Review and Approval form and provide the approval signature/date where indicated on the form.

## Upon signature and date on the DHR Review and Approval form the finished product is considered released for its designated use.

## If any discrepancies in the DHR are observed, the DHR is not to be approved and the final product not released until discrepancies are corrected or any corrective actions are completed.

### The DHR may be approved for internal use (not for human use) with documented justification. Product released for internal use (not for human use) shall be labeled appropriately.

## All DHRs and completed DHR Review and Approval forms are to be archived in storage at SILQ.

### All DHRs and completed DHR Review and Approval forms are to be scanned and imported into FileHold and filed into the appropriate DHR folder.