Advisory Notices and Recalls

Standard Operating Procedure

# Purpose

## The purpose of this procedure is to define the process for issuing advisory notices and conducting recalls/field safety corrective actions of SILQ products.

## This procedure is intended to facilitate identification and completion of advisory notices and/or recalls/field safety corrective actions involving SILQ devices, using the FDA regulation. This procedure should be used in conjunction with the referenced regulatory documents in determining if events involving SILQ products should be communicated to customers via Advisory Notices or recalled and if necessary reported to the FDA.

## SILQ will issue advisory notices and recall finished devices from its consignees when risk analyses indicate significant risk to public health or deviations from established specifications.

# Scope

## The following procedure applies to all SILQ products.

## Information may come to the attention of SILQ by various means, including but not limited to: post-marketing surveillance, consumer complaint reports, Management Review, and/or contract manufacturer.

## Market withdrawal and stock recovery are not within the scope of this procedure. These activities pertain to the removal of distributed devices that are not in violation of current regulations and to the recovery of devices that have not left the direct control of SILQ.

# Documents

## Reference Documents

### QM.SLQ016	Corrective and Preventive Action SOP

### QM.SLQ021	Product Complaint System SOP

### QM.SLQ022	Medical Device Reporting SOP

### QM.SLQ027	Quality Manual

### 21 CFR 7	Voluntary Recalls

### 21 CFR 806	Medical Devices; Reports of Corrections and 					Removals

### 21 CFR 810	Medical Device Recall Authority

### 21 CFR 820.198	FDA Quality System Regulation – Complaint Files

### ISO 13485:2016	International Standards Organization – Medical Devices – Quality management systems – Requirements for regulatory purposes (section 8.3.3 – Actions in response to nonconforming product detected after delivery)

### EN ISO 14971:2012	Application of Risk Management to Medical Devices

### MEDDEV 2.12-1; Rev. 8; Guidelines on a Medical Device Vigilance System

### EU Directive for Medical Devices (MDD), 93/42/EEC (Annex II/V/VI Section 3.1)

## Appendices / Associated Forms / Templates

### FM1-QM.SLQ030	Recall Report Submitted to FDA

### FM2-QM.SLQ030	Field Safety Corrective Action Report (EU)

### FM3-QM.SLQ030	Field Safety Notice

### FM4-QM.SLQ030	Advisory and Recall Notices Distribution Log

# Definitions

## Abbreviations / Acronyms

### EU: European Union

## General Definitions

### Advisory Notice:  Notice issued by SILQ, subsequent to the delivery of a medical device, to provide information and/or advise what action should be taken in regards to the following items. NOTE: A field safety notice (a.k.a. safety alert) is a particular type of Advisory Notice.

#### The use of a medical device

#### The modification of a medical device

#### The return of a medical device to the organization that supplied it

#### The destruction of a medical device

#### The exchange of a medical device

#### The retrofit by purchaser of Manufacturer’s modification or design change

#### The advice given by Manufacturer regarding the use of the device and/or the follow up of patients, users or others (for example, where the device is no longer on the market or has been withdrawn but could still possibly be in use, such as implants, or change in analytical sensitivity or specificity for diagnostic devices).

### Complaint: Any statement (written, electronic, or oral) alleging deficiencies related to the identity, quality, durability, reliability, safety, effectiveness, or performance of product manufactured and/or sold by SILQ after it is released to its consignees.

### Complaint File:  The call forms, complaint forms, and SILQ (or contract manufacturer) evaluation of the returned part(s), if they were returned. The Complaint File will include, at a minimum:

#### The completed complaint form, with supporting documentation, stored in the complaint management system.

#### Any notes, communications, and other relevant information associated with the investigation of the complaint; stored in the complaint management system.

### Consignee:  Any person or firm that has received, purchased, or used a device that is subject to an advisory notice or recall. These entities typically comprise SILQ customers. Consignee does not mean lay individuals or patients, i.e., non-health professionals. (21 CFR 810)

### Correction:  The repair, modification, adjustment, relabeling, destruction or inspection (including patient monitoring) of a device without its physical removal from its point of use to some other location. (21 CFR 810)

### Corrective and Preventive Action (CAPA):   A process and designated team that evaluates and provides quality and performance feedback to determine appropriate corrective and preventive actions to be performed. (21 CFR 820)

### Corrective Action:  Measures taken to correct existing nonconformities or deficiencies and prevent future recurrence of the problem. (21 CFR 820)

### Distributor:  Any person who furthers the distribution of a device from the original place of manufacture to the person who makes delivery or sale to the ultimate user, i.e., the final or multiple distributors, but who does not repackage or otherwise change the container, wrapper, or labeling the device or device package. (21 CFR 820)

### Effectiveness Check:  Methods of verifying that the consignees have received Field Safety Notice/Advisory Notice and have taken the appropriate action. (MEDDEV 2.12-1)

### European Authorized Representative:  Formally designated person who acts and may be addressed by authorities and bodies in the Community instead of the Manufacturer with regard to the latter’s obligations under the directive.

### Field Safety Corrective Action (FSCA):  An action taken by a manufacturer to reduce a risk of death or serious deterioration in the state of health associated with the use of a medical device that is already placed on the market; equivalent to a recall or withdrawal. Such actions should be notified via a field safety notice. (MEDDEV 2.12-1)

### Field Safety Notice:  A communication to customers and/or users sent out by a manufacturer or its representative in relation to a Field Safety Corrective Action. Also known as a Safety Alert. (MEDDEV 2.12-1)

### Health Hazard Evaluation (HHE):  A form originated by the FDA’s Center for Devices and Radiologic Health (CDRH) and modified by Regulatory Affairs to assist in evaluating a hazard or risk presented by a medical device manufacturer by/for SILQ. This report is not to be considered a form or a template. (21 CFR 7)

### Investigation:  An inquiry or examination of the complaint which shall be initiated by Quality/Regulatory or designee when a trend or the potential development of hazardous conditions is suspected.

### Market Withdrawal:  Removal or correction of a distributed product which involves a minor violation that would not be subject to legal action by the FDA or which involves no violation (for example, normal stock rotation practices). (21 CFR 7)

### NCA:  National Competent Authority (MEDDEV 2.12-1)

### Recall:  The correction or removal of a device for human use where FDA finds that there is a reasonable probability that the device would cause serious, adverse health consequences or death. (21 CFR 810)

#### Or the removal or correction of marketed device that the FDA considers to be in violation of the laws it administers and against which the agency would initiate legal action (for example, seizure). (21 CFR 810)

#### Recall does not include market withdrawal or stock recovery. (21 CFR 7)

#### In addition, recalls can be initiated if the medical device fails to conform to any claims made by the manufacturer or importer relating to its effectiveness, benefits, performance characteristics, safety, or does not comply with appropriate government regulations. Recall activities require formal, documented attempts to account for all materials distributed. (21 CFR 810)

### Recall classification: The numerical designation assigned by the FDA (i.e., I, II, or III, as defined in 21 CFR 7.3 (m)) to a particular product recall to indicate the relative degree of health hazard presented by the product.

### Removal:  The physical removal of a device from its point of use to some other location for repair, modification, adjustment, relabeling, destruction, or inspection. (21 CFR 810)

### Risk to Health:  Either (1) there is reasonable probability that use of, or exposure to, the product will cause serious adverse health consequences or death, or (2) the use of, or exposure to, the product may cause either temporary or medically reversible adverse health consequences or an outcome where the probability of serious adverse health consequences is remote. (21 CFR 806)

### Root Cause:  The lowest-level correctable assembly, component or process that is determined to be the most likely cause or contributor of the reported event and should be supported by findings of the investigation.

### Serious, adverse health consequence:  Any significant adverse experience, including those that may be either life threatening or involve permanent or long-term injuries, but excluding injuries that are non-life-threatening and that are temporary and reasonably reversible. (21 CFR 810)

### Serious Public Health Threat:  Any event type which results in imminent risk of death, serious deterioration in state of health, or serious illness that requires prompt remedial action. (MEDDEV 2.12-1)

### Trend:  A series of data points that lead to a decision about a problem’s actual or potential existence or lack of existence. Typically, a trend is reflected by a run of more than three data points in the same direction or by a set of data that allows a linear regression to fit the data well.

# Responsibilities

## It is the responsibility of SILQ’s Quality Assurance department in conjunction with SILQ’s Executive Management to determine the need to issue an advisory notice or recall related to a SILQ product.

## It is the responsibility of the management review team to monitor quality assurance data available to the organization and determine based on these data, when advisory notices and/or recalls are warranted.

## It is the responsibility of the Quality Assurance Management to perform relevant investigations and analyses to determine the severity and extent of any problem being considered for an advisory notice or recall.

# Procedure

## Identification of issues that potentially suggest problems with SILQ manufactured products could come from any of the following sources:

### As part of ongoing manufacturing, post-market surveillance and/or complaint handling, a SILQ or contract manufacturer employee identifies an event or series of events which potentially suggests problems with SILQ manufactured products.

### As a part of the Management Review and CAPA processes, SILQ management will review quality assurance data to identify emerging issues and trends that potentially suggest problems with SILQ manufactured products.

## In the event that available data suggests that a problem may exist with product already sent to a SILQ customer, management will launch an investigation as a part of the CAPA process (see QM.SLQ016) related to the problem.

## If the investigation reveals that product shipped to a SILQ customer either fails to conform to specification, potentially fails to conform to specification, or contains other defects that could potentially result in failure of a finished medical device, SILQ management team will assign a representative to perform an analysis of the potential impact of the problem and report the results to the management review team. Results of the analysis will be documented in the CAPA file related to the issue.

## The representative will:

### Define the suspected problem

### Identify affected part(s) and lot number(s)

### Determine date(s) of manufacture

### Identify the number of devices and/or accessories manufactured

### Identify the number of devices and/or accessories distributed

### Identify the amount remaining in inventory, if any

### Amount of inventory in process, if any

### Identify the consignee list

## The management review team will review the results of the investigation and determine whether an advisory notice, field safety corrective action/product recall, or combination is warranted.

## In assessing the need for an FSCA, SILQ will take into account the methodology described in EN ISO 14971:2012. In case of doubt, there should be a predisposition to report and undertake a Field Safety Corrective Action.

## The results of the review to determine appropriate action will be documented in the CAPA file related to the issue. The strategy may include:

### Advisory Notices should be used to notify SILQ customers of product/use issues and advise them on any actions to be taken. The subject of an advisory notice may or may not be safety related. When used to notify customers of safety problems, these notices are referred to as Safety Alerts or Field Safety Notices.

### Field Safety Corrective Actions/Recalls are the method for taking action to correct a safety related problem with a medical device in the field.

## In the event the management team determines that an advisory notice is warranted, the following process is to be followed:

### The Management Team will assign a representative to draft the advisory notice in accordance with the regulations.

#### The Advisory Notice will, at a minimum, contain the following information related to the affected product;

##### Name

##### Part Number

##### Affected Lot Numbers

##### Description of the problem

##### Anticipated consequences

##### Likelihood of occurrence

##### Suggested action by customer

### Once drafted, the Advisory Notice will be reviewed by Quality management and the CEO for approval.

### Once a version of the Advisory Notice is approved by the Quality management and the CEO, the notice will be translated into the native country language where the product was distributed.

### The notice in English and in the native country language will be sent to all affected customers per the consignee list. Distribution of advisory notices may be electronic (email or fax) or via the mail. In all cases, a return receipt for the notice is required and will be stored in the CAPA file related to the issue.

### Use FM4-QM.SLQ030 Advisory and Recall Notices Distribution Log to track the status of the Advisory Notice distribution activities. Once complete, store the log in the appropriate CAPA file.

### If the Advisory Notice requests the customer to return affected products, the products will be handled per QM.SLQ021 Product Complaint System SOP and dispositioned appropriately.

## In the event that the Management Team determines that a Field Safety Corrective Action/Recall is to be initiated, the following process is to be followed:

### Quality management shall submit a report to the FDA as described in 21 CFR 806.10 within 10 working days of initiating the action/recall. See FM1-QM.SLQ030.

### Quality management shall submit an initial report to the relevant National Competent Authorities and SILQ’s notified body of the FSCA. See FM2-QM.SLQ030. Quality management shall notify the relevant distributor/customer with a Field Safety Notice. See FM3-QM.SLQ030. A copy of the Field Safety Notice per section 6.10 of this procedure should also be sent along with the FSCA notification to the relevant National Competent Authorities and SILQ’s notified body.

## The Management Team will assign a representative to draft the Field Safety/Recall Notice (sent to customers) to include the following:

### A uniform and consistent Field Safety/Recall Notice should be placed on company letterhead.

### The Field Safety Notice will, at a minimum, contain the following information related to the affected product (see FM3-QM.SLQ030 for a Field Safety Notice template):

#### A clear title with “URGENT FIELD SAFETY/RECALL NOTICE”, followed by the commercial name of the affected product, and FSCA-identifier and the type of action (for example, recall);

#### Specific details to enable the affected product to be easily identified, such as type of device, part number, part name, affected lot numbers;

#### A factual statement explaining the reasons for the FSCA including but not limited to a description of the problem, clarification of the hazards and the anticipated patient safety impact, likelihood of occurrence;

#### Advice on actions to be taken by the user including, as appropriate, identifying and quarantining the affected product; the method to return or dispose of affected product; recommended review of patients’ previous results or patient follow up (for example, implants).

#### A request to pass the Field Safety Notice to all those who need to be aware of it within the organization and to maintain awareness over an appropriate defined period;

#### If relevant, a request for the details of any affected devices that have been transferred to other organizations;

#### If relevant, a request that the recipient of the Field Safety Notice alerts other organizations;

#### Confirmation that the relevant National Competent Authorities have been advised of the FSCA;

#### If relevant, a request that the recipient maintain awareness of the Field Safety Notice and resulting action for an appropriate period to ensure effectiveness of the corrective action;

#### Any attempts to down play the risks in an inappropriate manner of advertised products or services should be omitted; and Contact point for customers.

### The Management team determines if phone calls to customers are required in addition to the Field Safety Notice/Recall Notice.

#### If phone calls are required, a script should be drafted to contain the appropriate information to communicate to the customer per Section 6.10.2. Alternately, the notice can suffice as the script.

### Once drafted, the Field Safety/Recall Notice and telephone script (if applicable) will be reviewed by the Quality Assurance management and CEO for approval.

### Once a version of the Field Safety Notice is approved by Quality Assurance management and the CEO, the notice will be translated into the native country language where the product was distributed.

### The notice in English and in the native country language will be sent immediately to all affected customers per the consignee list. Distribution of Field Safety Notices may be electronic (email or fax) or via the mail. In all cases, return receipt for the notice is required and will be stored in a separate Advisory Notice/Recall file related to the issue in appropriate folder within FileHold.

### If telephone calls are included in the recall strategy, calls should be initiated upon approval of the script and identification of the consignee list.

### Use FM4-QM.SLQ030 Advisory and Recall Notices Distribution Log to track the status of the notice distribution activities. Once complete, scan and import the log with the Advisory Notice/Recall file in appropriate folder within FileHold.

### Field Safety Notices will request the customer to either dispose of or return affected products. The products will be handled per QM.SLQ021 Product Complaint System and dispositioned appropriately. All product that are the subject of a recall will be segregated from all other returned product.

### The QA department will establish formal Effectiveness Check criteria for considering the corrective action complete. The Effectiveness Check criteria will specify the percentage of customers receiving the recall notice and the percentage of recall product accounted for (either returned or disposed of).

#### In the event that return receipts are not received on the first contact attempt from each customer, a representative from QA will attempt to contact the customer by any other feasible means.

#### At a minimum, three (3) documented attempts to contact each non-responsive customer must be made.

#### Recalls will be considered complete when established recall closure criteria is met or when three (3) documented attempts to contact each customer have been made without success.

#### Recall closure is approved by Quality Assurance management in writing and stored in the Advisory Notice/Recall file.

## Regulatory Agency Reporting

### FDA approved product

#### If it is determined that the Field Safety Corrective Action/Recall falls under the guidelines of FDA’s Corrections and Removals, then the FDA must be notified, as appropriate based on the recall classification, by following the instructions listed in 21 CFR 806.10 (reference FM1-QM.SLQ030).

#### If after submitting an initial report, it is determined that the same action must be extended to additional lots of the same device, an amended report must be submitted.

### CE marked product – notification to Competent Authorities and Notified Bodies

#### CE marked product should follow the guidelines of the Medical Device Vigilance system.

#### SILQ should issue a notification to the Competent Authorities of all countries affected at the same time and also to the NCA responsible for SILQ or authorized representative. Acknowledgment of receipt should be returned.

#### Content and requirements for notification to the Competent Authorities should follow MEDDEV 2.12-1 Rev. 8 Annex 4 (reference FM2-QM.SLQ030).

#### The European Authorized Representative will be notified in order to be able to report the FSCA to the relevant Competent Authorities at least 48 hours before the FSCA is implemented.

#### The Notified Body for SILQ will be notified of all advisory notices and recalls for CE marked products.

#### After the completion of the investigation and any remedial action, a final report shall be completed and should include a summary of the results of the recall, for instance the number of units recovered, not recovered, corrected, and/or destroyed. See FM2-QM.SLQ030. The final report shall be submitted to relevant National Competent Authorities.

## Quality Records

### The Advisory Notice/Recall report file will be numbered the same as the corresponding CAPA number.

### The complete Advisory Notice/Recall file will contain, or make reference to, all documents related to the recall including:

#### CAPA documentation

#### Risk Analyses (HHE or equivalent)

#### Recall report (sent to FDA)

#### Field Safety Correction Action report (initial and final reports sent to relevant National Competent Authorities and SILQ’s notified body)

#### Field Safety Notice/Recall Notice

#### Return receipt notifications

#### Documented contact attempts

#### Product disposition information

#### Non-conforming material reports

### Records of advisory notices/recalls shall be maintained for a period of five years or a period of time equivalent to the expected life of the device, whichever is greater.