Post-Market Surveillance

Standard Operating Procedure

# Purpose

## The purpose of this procedure is to outline the methods for gathering and reviewing post-market information as well as post-market production information to ensure customer requirements are met, to provide early warning signs of quality problems throughout the life-cycle of the medical device, and to ensure the long term safety and performance of devices after commercial release.

# Scope

## This procedure shall apply to all SILQ products approved for commercial release and distribution.

# Documents

## Reference Documents

### QM.SLQ001	Document Control SOP

### QM.SLQ016	CAPA SOP

### QM.SLQ018	Management Review SOP

### QM.SLQ021	Product Complaint System SOP

### QM.SLQ027	Quality Manual

### EU Directive for Medical Devices (MDDD) 93/42/EEC (Annex I/V/VI Section 3.1)

### MEDDEV 2.12/2 Rev. 2 (Jan 2012) Guidelines on Medical Devices – Post Market Clinical Follow-up Studies. A Guide for Manufacturers and Notified Bodies

### NB-MED 2.12-1 Rev 11 (Feb 2000) Post-Marketing Surveillance (PMS) post market/production

## Appendices / Associated Forms / Templates

### N/A

# Definitions

## Abbreviations / Acronyms

### CAPA: Corrective Action and Preventive Action

### PMCF: Post-Market Clinical Follow-up

### PMS: Post-Market Surveillance

### QA: Quality Assurance

### RA: Regulatory Assurance

## General Definitions

### Post-Market Clinical Follow-up Plan: The documented, proactive, organized methods and procedures set up by the manufacturer to collect clinical data based on the use of CE-marked device corresponding to its intended use. The objective is to confirm clinical performance and safety throughout the expected lifetime of the medical device, the acceptability of identified risks to detect emerging risks on the basis of factual evidence.

### Post-Market Clinical Follow-up Study: A study carried out following the CE Marking of a device and intended to answer specific questions relating to clinical safety or performance (e.g., residual risks) of a device when used in accordance with its approved labeling.

### Post-Market Surveillance: Systematic process to collect and analyze experience gained from medical devices that have been placed on the market.

### Residual Risk: Risk remaining after risk control measures have been taken.

# Responsibilities

## QA and RA are responsible for ensuring compliance with this procedure.

# Procedure

## A Post-Market Surveillance (PMS) Plan for each device shall be established. The requirements for the PMS Plan should be directly proportional to the risk associated with the device based on its intended use.

### PMS activities can be “reactive” – responding after an event; of which there are many types ranging from complaints to Field Safety Corrective Action/Recalls. These activities can be considered “passive” as they are largely data collection activities. See Table 1 for examples of reactive PMS data.

### PMS activities can be “proactive” – endeavors meant to anticipate and curtail events before they occur. In “proactive” PMS activities, information is actively sought to gain insight and data into the real-world performance of the device.

### Table 1. Examples of PMS data

| ## Proactive | ## Reactive |
| --- | --- |
| ## Post CE-market clinical trials, including PMCF ## User/customer questionnaires or surveys ## User reaction during training programs ## Expert user groups (focus groups) ## Bench-marking against competitors’ products ## Distributor experience and reports ## Manufacturer sponsored device tracking/implant registries ## MAUDE data of similar or competitor products ## EUDAMED database on medical devices | ## Customer Complaints ## Unsolicited customer feedback other than complaints ## Product returns and warranty claims ## In-house testing (routine) ## Failure analysis ## Literature reviews ## Field Safety Corrective Action/Recalls ## Public health advisories of safety notices from regulatory authorities |

### RA will evaluate the need for Post-Market Clinical Follow-up (PMCF) studies for each new product per MEDDEV 2.12/2 Rev 2 and establish a PMCF plan as appropriate. PMCF studies are one of several options available in post-market surveillance.  The decision to conduct PMCF studies must be based on the identification of possible residual risks and / or lack of clarity on long term clinical performance that may impact the benefit / risk ratio. This will be updated in the Clinical Evaluation Report section of the Technical File.

#### If PMCF studies are deemed not necessary, the justification must be documented in the PMS plan.

## All employees are responsible for ensuring any product-related feedback is forwarded to QA or RA. QA and RA will evaluate any feedback or complaints in accordance with the Product Complaint System procedure QM.SLQ021.

## Feedback information is analyzed and if necessary corrective and / or preventive actions are initiated per the CAPA procedure QM.SLQ016.

## Manufacturing and QA will provide an analysis of post-market production information that will be used to provide an early warning system for quality and service related issues.  Sources of post-market production information include, but are not limited to:

- Non-conforming Materials

- Device-related CAPAs

- Incoming Materials Quality Inspections

- Manufacturing Yields

- Analysis of processes

- Process validations

- Supplier issues

- Training issues

Feedback from these sources will be addressed during Management Reviews and other meetings as appropriate.

## QA and/or RA will prepare a summary of the feedback information obtained from PMS activities for Management Review meetings. Summary reports will be maintained as part of the Management Review document archives.

## Distributor agreements will include applicable quality and regulatory requirements (such as traceability, storage, feedback, complaint reporting, etc.). RA will be required to approve distributor agreements to ensure all quality and regulatory requirements are included in the agreement.

## Post-Market Surveillance Records shall be maintained per the requirements specified in QM.SLQ001 Document Control.