Receiving Inspection

Standard Operating Procedure

# Purpose

## The purpose of this standard operating procedure is to define the requirements for the receipt and inspection of incoming purchased product such as raw materials, components, parts, manufacturing materials and manufacturing supplies.

## This standard operating procedure ensures that all purchased and otherwise received product conforms to specified requirements and defines the process for incoming receiving acceptance activities, including, but not limited to, verification that received shipments:

## Are complete,

## Have proper identity,

## Are undamaged,

## Contain proper supporting documentation, as applicable

# Scope

## The following procedure applies to all products purchased by SILQ that impact the final product (commercial product or for use in clinical/investigational trials with human patients).  This includes, but is not limited to the following commercial and/ or fabricated items: manufacturing materials and supplies, raw materials, components and parts.

## This procedure does not apply to the receipt of purchased product intended solely for R&D use, refer to QM.SLQ020, Purchasing Controls SOP.  Note: while not required, this procedure may be used for receipt and inspection of these types of items.

## This procedure does not apply to the receipt of purchased items not used in the finished product or associated with the manufacturing process such as office supplies or items.

# Documents

## Reference Documents

### QM.SLQ001	Document Control SOP

### QM.SLQ015	Supplier Quality Assurance SOP

### QM.SLQ019	ID and Traceability SOP

### QM.SLQ020	Purchasing Controls SOP

### QM.SLQ040	Nonconforming Materials SOP

### QM.SLQ045	Receiving SOP

### 21 CFR Part 820, FDA Quality System Regulation (820.80 – Receiving, in-process, and finished device acceptance)

### ISO 13485:2016 – Medical Devices – Quality Management Systems – Requirements for Regulatory Purposes

## Appendices / Associated Forms / Templates

### Appendix 1		Chart of C=0 Sampling Plans

### FM1-QM.SLQ039	Basic Receiving Inspection Form

### FM2-QM.SLQ039	General Inspection Form

# Definitions

## Abbreviations / Acronyms

### AQL:  Acceptable Quality Limit per ANSI Z1.4

### ASL:  Approved Supplier List

### COA:  Certificate of Analysis

### COC:  Certificate of Conformance

### P.O.: Purchase Order

### RI:  Receiving Inspection

### SDS:  Safety Data Sheet

## General Definitions

### Acceptable Quality Limit:  Expressed as a percentage, the acceptable rejection level of product while maintaining lot acceptability (i.e., the largest percent defective in a lot that would be considered acceptable).

### Certificate of Analysis:   A signed document affirming that a certain product/material was tested for certain qualities and found to be acceptable.

### Certificate of Conformance: A signed document affirming that a product/ material has met the requirements of the relevant specifications, contract or regulation.

### Component: any material, substance, piece, part, software, firmware, labeling, or assembly, which is intended to be included in the finished, packaged, and labeled device.

### Fabricated Product:  Custom-manufactured raw materials, components, manufacturing materials, in-process devices, final product which are not available as an off-the-shelf item.

### Index Value: When used in zero acceptance sampling plans, the index value is the number associated with the AQL of ANSI Z1.4 to determine the appropriate sample size.  Note: per Squeglia “in all the zero acceptance sampling plans, the protection afforded to the consumer is equal to or greater than the ANSI Z1.4 plans.  This method of developing the plans provides for simple conversion from the ANSI Z1.4 plans to the c=0 plans.  The plans label these associated AQLs as “index values” because, of course, they are not AQLs.”

### Lot: product that consist of a single type, model, class, size, composition or software version that are manufactured under essentially the same conditions and that intended to have uniform characteristics and quality within specified limits.

### Manufacturing Material: any material or substance used in or used to facilitate the manufacturing process, a concomitant constituent, or a byproduct constituent produced during the manufacturing process, which is present in or on the final product as a residue or impurity not by design or intent of the manufacturer.

### Operating Supplies:  Product or materials that are not used in or used to facilitate the manufacturing process and are not in or on the final product, however are essential to the manufacturing processes and environmental controls such as filters, water separators, lubricants, etc.

### Product:  components, manufacturing materials, in-process devices, final product, and returned final product.

### Purchase Order:  a document issued by SILQ to a seller (supplier), indicating types, quantities, and agreed prices for products or services the seller will provide to SILQ. It is used to control the purchasing of products and services from external suppliers.   Issuing a purchase order to a supplier constitutes a legal offer to buy products or services. Acceptance of a purchase order by a seller usually forms a contract between SILQ and seller, so no contract exists until the purchase order is accepted.

### Quarantine: The process of physically segregating or isolating unknown or nonconforming materials from usable or accepted materials.

### Receiving Inspection:  the inspection and acceptance or rejection of all incoming product.

### Sample: A sample consists of one or more units of product randomly drawn from a lot.

### Source Controlled: Off-the-shelf raw materials or components purchased from a particular supplier based on that supplier’s part number and specification.

### Specification: any requirement with which a product, process, service or other activity must conform.

### Specification Controlled: Fabricated raw materials, components and assemblies purchased based on specifications generated by SILQ.

# Responsibilities

## It is the responsibility of the Operations management to ensure all aspects of this SOP are followed.

## It is the responsibility of authorized receiving inspection personnel to receive and inspect product in accordance with this procedure and to maintain applicable records.

## Quality Assurance management is responsible for ensuring compliance with the receiving / receiving inspection process and maintenance of associated records.

# Procedure: General Requirements

## All product shall be received in accordance with QM.SLQ045, Receiving SOP.

## All product received (with the exception of product intended for R&D / nonhuman use only) is to be from a supplier approved in accordance with QM.SLQ015, Supplier Quality Assurance SOP and listed on the Approved Supplier List (ASL).   Refer to Section 8.

## Fabricated product are to be inspected to an approved and released specification document (e.g. Engineering drawing).

## A packing list, copy of the Purchase Order (P.O.), or similar documentation, is to be provided by the supplier with the product and it is to contain sufficient information so that receiving personnel can match it to the correct SILQ P.O.

## Any inspection or measurement equipment required for Receiving Inspection (RI) of incoming product is to be appropriate for the task and within calibration (if applicable).

## Inspection of incoming product is to be performed per the inspection plans specified in this procedure as a minimum requirement, refer to Section 7, or the product specific inspection procedure / plan if available.

# Procedure: Inspection Plans (classification of Material Type)

## Inspection of product is to be performed based upon the assigned Material Type.   The materials are divided into distinct groups based upon their function, criticality and/or risk as it relates to final product.  The purpose for these groupings is to optimize efficiency in receiving /  inspection and to prevent or significantly reduce the use of low quality and/or defective parts that may impact the quality of the final product.

## The Inspection Plans listed in Table 1 are considered the minimum requirements; more stringent inspection plans may be employed.

## R&D / Engineering will determine the material type and inspectional attributes for individual parts.  The information will be specified in an inspection plan, source control sheet, or on the released drawing/specification.

### If no inspection plan, source control sheet, or on the released drawing/specification document exists, product is to be inspected to the requirements listed in Table 1.

Table 1	Table Classification of Material Type and Inspection Plans

| Material Type | Description | Inspection Plan |
| --- | --- | --- |
| A | - Products (raw materials, components, and/or assemblies) critical to the function or safety of a final product and where the product is manufactured such that dimensions may not be uniform across manufacturing lot (e.g., could have variability in key dimension(s) from part to part). - Products where yield and cost are critical. | 100% |
| B | - Products (raw materials, components and/or assemblies) included in the final product that could affect the safety or efficacy of the final product, but would be obvious to the end user (i.e. the device would not function or would be readily apparent to the operator). - Products manufactured such that dimensions are uniform across manufacturing lot (e.g., horizontal dimension of wafers, laser cutting, etc.). | AQL 1.0 |
| C | - Raw materials, components and/or assemblies which are available as “off the shelf” commercial product. | Basic |
| D | - Printed materials including product literature, packaging inserts, and labels, etc. | N = 1 100% proof of content; AQL 1.0 for legibility |
| E | - Other purchased products that do not affect product or other Quality Systems related to SILQ activities, such as plant and office supply items, machine shop raw materials and facilities materials. They do not require documented inspection. - R&D materials. | Record in receiving log only |

## The Zero Acceptance Number (C=0) Sampling Plans (from here on referred to as C=0 sampling plans) will be used for product inspected during the receiving inspection process unless specified differently on the product specific inspection plan, source control sheet, or on the released drawing/specification.  The acceptance number in all cases is zero with these sampling plans.  This means that to achieve a desired level of protection, a certain size sample is selected, the sample is inspected, and must have zero nonconforming (defective) samples to be considered accepted.

### As developed in the C=0 sampling plans, the “Index Value” is the number associated with the Acceptance Quality Limit (AQL) of ANSI Z1.4.  It is used to determine the sample size based on the lot size.  The AQLs are indicated for each material Type (refer to Table 1), and shall be used with the Index Value in Appendix 1, Chart of C=0 Sampling Plans.

# Procedure: Basic Inspection Plan

## This type of inspection plan generally applies to commercial product (off the shelf), source-controlled product or product related items (see Table 1).

## Product is to be inspected according the basic requirements below and results recorded on FM1-QM.SLQ039, Basic Receiving Inspection Form.

## Product is to be inspected to a current approved and released source control drawing/specification revision.  If none exists, inspect per the description on the P.O. and associated manufacturer’s specification

### Obtain the product with associated FM1-QM.SLQ039, Basic Receiving Inspection Form (as completed by the receiving designee) from the segregated area designated for ‘Pending Receiving Inspection’.

#### Review the Receiving section of the form for completeness.  Resolve any issues with the receiving designee

### Lot Number:  Enter the specific lot number for the product per requirements listed in QM.SLQ019, ID and Traceability SOP.

### Quantity Received: Enter the actual quantity of product received.

### Unit of Measure: Enter the unit of measure used for the product quantity (e.g., each, bottles, etc.)

### Inspection Plan: Indicate the material Type as ‘C’ for basic per Table 1.

### Quantity Inspected: For basic inspection enter ‘N/A’, as only a visual inspection of the primary packaging and associated documentation is being inspected.

### Verify that product delivered matches the part number or part description on the P.O:  Indicate results by checking applicable box.

### Verify quantity and unit of measure delivered of product matches the quantity and unit of measure ordered on the P.O.:  Indicate results by checking applicable box.

#### If the quantity does not match notify Purchasing.  Indicate if Purchasing reconciled P.O. (adjusted) to match the actual shipment or if not, indicate as ‘Fail’

### Inspect the primary package for damage:  Indicate results by checking applicable box.

### Review the COC/COA as applicable, to ensure information is correct per product requirements and attach to FM1-QM.SLQ039, Basic Receiving Inspection Form:  Indicate results by checking applicable box.

### Lot inspection: For basic inspection enter ‘N/A’, as only a visual inspection of the primary packaging and associated documentation is being inspected.

### Inspected By:  The performing the inspection signs their name here.

### Date: The employee enters the date they signed the record.

## Refer to Section 11 for completion of Lot Acceptance Results.

## Refer to Section 12 for labeling and release procedures

# Procedure: 100% Inspection Plan

## This type of inspection plan generally applies to product critical to the function or safety of a final product and where the product is manufactured such that dimensions may not be uniform across manufacturing lot (e.g., could have variability in key dimension(s) from part to part); (see Table 1).

## Product is to be inspected to a current approved and released drawing/specification revision with 100% inspection of all incoming product for the specified characteristics listed on the Receiving Inspection (RI) Procedure for the corresponding product.

### Obtain the product with associated Inspection Form (as completed by the receiving designee) from the segregated area designated for ‘Pending Receiving Inspection’.

#### Review the Receiving section of the form for completeness.  Resolve any issues with the receiving designee.

## Results shall be recorded on the associated product specific Inspection Form.  If a product specific Inspection Form is not in place for a given product use FM2-QM.SLQ039, General Inspection Form to record the inspection criteria and results and attach to FM1-QM.SLQ039, Basic Receiving Inspection Form.

## Refer to Section 11 for completion of Lot Acceptance Results.

## Refer to Section 12 for labeling and release procedures.

# Procedure: Sampling Inspection Plan, AQL 1.0

## This type of inspection plan generally applies to could affect the safety or efficacy of the final product, but would be obvious to the end user (i.e. the device would not function or would be readily apparent to the operator); or where the product is manufactured such that dimensions are uniform across manufacturing lot (low risk of variability of dimension(s) from part to part); (see Table 1).

## Product is to be inspected with sample sizes that are based on the size of the lot received per C=0, with an Index Value (approximate AQL) of 1.0 Sampling Plan specified in Appendix 1.  If a different AQL is specified in the product specific inspection plan, the Index Value and sample size is to be adjusted as necessary.

### Smaller individual receipts of the same part number from the same supplier and supplier lot may be combined into a single larger lot if the inspected sample is selected at random throughout all of the lots.

## Product is to be inspected to a current approved and released drawing/specification revision with sampling inspection of all incoming product for the specified characteristics listed on the Receiving Inspection (RI) Procedure for the corresponding product.

### Obtain the product with associated Inspection Form (as completed by the receiving designee) from the segregated area designated for ‘Pending Receiving Inspection’.

#### Review the Receiving section of the form for completeness.  Resolve any issues with the receiving designee.

## Results shall be recorded on the associated product specific Inspection Form.  If a product specific Inspection Form is not in place for a given product use FM2-QM.SLQ039, General Inspection Form to record the inspection criteria and results and attach to FM1-QM.SLQ039, Basic Receiving Inspection Form.

## If a sample from a lot fails inspection, the entire lot is to be rejected.  Refer to Section 11 for completion of Lot Acceptance Results.

## Refer to Section 12 for labeling and release procedures.

# Procedure: Lot Disposition

## The acceptance or rejection of the product lot must be documented.

## Lot Disposition:  Review the receiving and inspection sections of the associated inspection form(s).  If all results are indicated as ‘Pass’, indicate ‘Accepted’ for lot disposition.

### If any result was indicated as ‘Fail’, the lot is to be rejected;  indicate ‘Rejected’.  Initiate a Non-Conforming Material Report per QM.SLQ040 Nonconforming Materials SOP.  The product lot is to be removed and placed in quarantine pending the result of the Non-Conforming Material Report.

## Quantity Released from Quarantine:  Enter the quantity of accepted product to be released.

### If the lot disposition was ‘rejected’, enter ‘0’.

## Comments:  Enter any applicable comments.

## Disposition By:  Review the record for completeness and resolve and discrepancies.  If there were no comments, enter ‘N/A’ in the comments section.  Sign the form.

## Date:  Enter the date the disposition section was signed.

# Procedure: Product Labeling and Release

## Update product label with disposition of the product:

### Status: Description of the status. An example is: “Accepted”.

## Release of Dispositioned Product

### Once product has been dispositioned, the completed RI paperwork is to be filed in the associated Receiving Part Files in secure file storage.

### Accepted product is to be moved to appropriate storage location.

### Rejected product is to be moved to quarantine area per QM.SLQ040 Nonconforming Materials SOP.

# Procedure: Receiving and Receiving Inspection Records

## All records and applicable documentation are to be maintained in accordance with record retention policies defined in QM.SLQ001, Document Control SOP.

### Records and applicable documentation are to be scanned and imported into FileHold and filed within appropriate receiving inspection file folder.

Appendix 1 – Chart of C = 0   Sampling Plans

|  | Index Values (associated AQLs) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | 0.25 | 0.40 | 0.65 | 1.00 | 1.50 | 2.50 | 4.00 | 6.50 |
| Lot Size | Sample Size |  |  |  |  |  |  |  |
| 2 to 8 | * | * | * | * | * | 5 | 3 | 3 |
| 9 to 15 | * | * | * | 13 | 8 | 5 | 3 | 3 |
| 16 to 25 | * | * | 20 | 13 | 8 | 5 | 3 | 3 |
| 26 to 50 | * | 32 | 20 | 13 | 8 | 7 | 7 | 5 |
| 51 to 90 | 50 | 32 | 20 | 13 | 13 | 11 | 8 | 5 |
| 91 to 150 | 50 | 32 | 20 | 19 | 19 | 11 | 9 | 6 |
| 151 to 280 | 50 | 32 | 29 | 29 | 19 | 13 | 10 | 7 |
| 281 to 500 | 50 | 48 | 47 | 29 | 21 | 16 | 11 | 9 |
| 501 to 1,200 | 75 | 73 | 47 | 34 | 27 | 19 | 15 | 11 |
| 1,201 to 3,200 | 116 | 73 | 53 | 42 | 35 | 23 | 18 | 13 |
| 3,201 to 10,000 | 116 | 86 | 68 | 50 | 38 | 29 | 22 | 15 |
| 10,001 to 35,000 | 135 | 108 | 77 | 60 | 46 | 35 | 29 | 15 |
| 35,001 to 150,000 | 170 | 123 | 96 | 74 | 56 | 40 | 29 | 15 |
| 150,001 to 500,000 | 200 | 156 | 119 | 90 | 64 | 40 | 29 | 15 |
| 500,001 and over | 244 | 189 | 143 | 102 | 64 | 40 | 29 | 15 |
| * Indicates entire lot must be inspected |  |  |  |  |  |  |  |  |
| NOTE: The Acceptance Number in all cases is ZERO |  |  |  |  |  |  |  |  |

Squeglia, N. L. (2008).  Zero Acceptance Number Sampling Plans – 5th ed. Milwaukee: American Society for Quality, Quality Press. Page 6.

Squeglia, N. L. (2008).  Zero Acceptance Number Sampling Plans – 5th ed. Milwaukee: American Society for Quality, Quality Press.

Squeglia, N. L. (2008).  Zero Acceptance Number Sampling Plans – 5th ed. Milwaukee: American Society for Quality, Quality Press. Table 1b: Modifed C=0 Sampling Plans.