Receiving

Standard Operating Procedure

# Purpose

## The purpose of this standard operating procedure is to define the process for receiving incoming purchased product such as raw materials, components, parts, manufacturing materials, and manufacturing supplies.

# Scope

## The following procedure applies to all products purchased by SILQ that impact the final product (commercial product or for use in clinical/investigational trials with human patients).  This includes, but is not limited to the following commercial and/ or fabricated items: manufacturing materials and supplies, raw materials, components and parts.

## This procedure does not apply to the receipt of purchased items not used in the finished product or associated with the manufacturing process such as office supplies or items.

# Documents

## Reference Documents

### QM.SLQ001	Document Control SOP

### QM.SLQ019	ID and Traceability SOP

### QM.SLQ039	Receiving Inspection SOP

### QM.SLQ040	Nonconforming Materials SOP

### 21 CFR Part 820, FDA Quality System Regulation (820.80 – Receiving, in-process, and finished device acceptance)

### ISO 13485:2016 – Medical Devices – Quality Management Systems – Requirements for Regulatory Purposes

## Appendices / Associated Forms / Templates

### FM1-QM.SLQ039	Basic Receiving Inspection Form

### TMP1-QM.SLQ045	Quarantine Label Template

# Definitions

## Abbreviations / Acronyms

### ASL:  Approved Supplier List

### COA: Certificate of Analysis

### COC:  Certificate of Conformance

### P.O.: Purchase Order

## General Definitions

### Certificate of Analysis:   A signed document affirming that a certain product/material was tested for certain qualities and found to be acceptable

### Certificate of Conformance: A signed document affirming that a product/ material has met the requirements of the relevant specifications, contract or regulation.

### Component: any material, substance, piece, part, software, firmware, labeling, or assembly, which is intended to be included in the finished, packaged, and labeled device.

### Lot: product that consist of a single type, model, class, size, composition or software version that are manufactured under essentially the same conditions and that intended to have uniform characteristics and quality within specified limits.

### Manufacturing Material: any material or substance used in or used to facilitate the manufacturing process, a concomitant constituent, or a byproduct constituent produced during the manufacturing process, which is present in or on the final product as a residue or impurity not by design or intent of the manufacturer.

### Operating Supplies:  Product or materials that are not used in or used to facilitate the manufacturing process and are not in or on the final product, however are essential to the manufacturing processes and environmental controls such as filters, water separators, lubricants, etc.

### Product:  components, manufacturing materials, in-process devices, final product, and returned final product.

### Purchase Order:  a document issued by SILQ to a seller (supplier), indicating types, quantities, and agreed prices for products or services the seller will provide to SILQ. It is used to control the purchasing of products and services from external suppliers.   Issuing a purchase order to a supplier constitutes a legal offer to buy products or services. Acceptance of a purchase order by a seller usually forms a contract between SILQ and seller, so no contract exists until the purchase order is accepted.

### Quarantine: The process of physically segregating or isolating unknown or nonconforming materials from usable or accepted materials.

### Receiving Inspection:  the inspection and acceptance or rejection of all incoming product.

### Source Controlled: Off-the-shelf raw materials or components purchased from a particular supplier based on that supplier’s part number and specification.

### Specification: any requirement with which a product, process, service or other activity must conform.

### Specification Controlled: Fabricated raw materials, components and assemblies purchased based on specifications generated by SILQ.

# Responsibilities

## It is the responsibility of the Operations management to ensure all aspects of this SOP are followed.

## It is the responsibility of authorized receiving inspection personnel to receive and inspect product in accordance with this procedure and to maintain applicable records.

## Quality Assurance management is responsible for ensuring compliance with the receiving / receiving inspection process and maintenance of associated records.

# Safety Precautions

## Always perform proper lifting techniques, use pallet jack or forklift when lifting heavy packages.

## Always handle box cutter with caution, to prevent harm to self and product.

## Place hazardous materials in designated areas as soon as possible.

# Procedure: Receiving

## All product delivered to Silq shall be routed through the receiving dock.

## Sign for items received.

## All incoming shipments must be examined for external signs of damage.  If the shipment is damaged, immediately notify Purchasing and move to Hold area until disposition is decided by Purchasing.

## Obtain copy of P.O.

## Obtain packing slip and COC/COA from shipping package.

### Stamp or record receiving date on packing slip.

### Verify numbers of packages and identification of packages to packing slip.

### Attach packing slip to copy of P.O.

## For fabricated product, obtain product specific Receiving Inspection form.  If none exists, use FM1-QM.SLQ039 Basic Receiving Inspection form.

## For commercial (off the shelf) product, obtain FM1-QM.SLQ039 Basic Receiving Inspection form.

## Complete the ‘Receiving’ section of the respective inspection form as follows:

### Date Received: Enter the date the product was received.

### Material Description: Enter the title of the product from the associated specification document, if applicable, or general product description.

### SILQ Part Number and Revision: Enter the associated part number and revision level, if none, enter ‘N/A’.

### Purchase Order Number: Enter the associated P.O. number.

### Purchase Order Quantity: Enter the quantity of product ordered per the P.O.

### Supplier Lot Number:  Enter the supplier’s lot number for the product.  If none, enter ‘N/A’.

### Expiration Date: Enter the expiration date of the product. If none, enter ‘N/A’.

### Shipping Package Inspection: Indicate if the exterior shipping package is free of damage or deterioration from inspection per Section 7.3.

#### If Fail, initiate a Non-Conforming Material Report per QM.SLQ040 Nonconforming Materials SOP.

### Approved Supplier Verification: Indicate if the supplier is on the ASL.  If the supplier is not on the ASL, but the product is intended for R&D Use Only, indicate such on form.

#### If Fail, initiate a Non-Conforming Material Report per QM.SLQ040 Nonconforming Materials SOP.

### Completed By:  The receiving designee signs their name here

### Date: The employee enters the date they signed the record.

## Place product with Receiving Inspection form in segregated area designated for ‘Pending Receiving Inspection’.

## Notify person responsible for receiving inspection of product that product is awaiting inspection.

## Forward a copy of the P.O. and packing slip (from Section 7.5.3) to Purchasing.

## The receiving designee is to update the receiving log identifying the date product was received, who (supplier) the product was received from, supplier lot number, quantity, part number, revision, description of the product, the associated P.O. number, initials of who received the product and the name of who was notified of the incoming product for inspection.

# Procedure: Product Labeling and Quarantine

## Received product is to be labeled with the following information:

### SILQ Part Number (P/N): Alphanumeric code that identifies a particular assembly or component per QM.SLQ001 Document Control SOP, or a sample name per a particular DMR file.

### Rev: Alphanumeric code that identifies the revision level of the associated source control drawing or specification document.

### Description: A brief description of the product (e.g., polymeric dispersion solution)

### Lot number: Alphanumeric code that identifies which lot a particular assembly, component or test sample comes from.

### Status: Identification of the quarantine status.

## Label template TMP1-QM.SLQ045 may be used to generate the label with a Dymo labelmaker.

# Procedure: Receiving Records

## All records and applicable documentation are to be maintained in accordance with record retention policies defined in QM.SLQ001, Document Control SOP.

### Records and applicable documentation are to be scanned and imported into FileHold and filed within appropriate receiving inspection file folder.