Workstation Practices

Standard Operating Procedure

# Purpose

## The purpose of this procedure is to establish good manufacturing practices for manufacturing workstation practices, safety, housekeeping, documentation, equipment maintenance, personnel guidelines, and handling and storage of chemical containers in the production environment.  This will assure that components and materials are handled in a safe manner in order to prevent damage, mix up, contamination, and commingling or any other adverse effects to either the product or manufacturing personnel.

# Scope

## This procedure applies to all Silq personnel and subcontractors involved with manufacturing, inspection, service, and handling of product.  facilities and personnel operating in those environments.

# Documents

## Reference Documents

### QM.SLQ001		Document Control SOP

### QM.SLQ002		Good Documentation Practices SOP

### QM.SLQ040		Nonconforming Materials SOP

# Definitions

## Abbreviations / Acronyms

### CEA: Controlled Environment Area

### LHR: Lot History Record

### SDS: Safety Data Sheet

## General Definitions

### FileHold: Software-based document management system used to electronically store controlled documents.

# Responsibilities

## Trained Silq personnel are responsible for adhering to the requirements within this procedure.

## Operations – Compiles the lot history records for each production batch or lot. Performs a line clearance prior to beginning work on a new lot, new assembly, new component, or any new operation.

## Quality Assurance – Reviews lot history records prior to release of production lot or unit. Maintaining the lot history records for each product batch or lot.

# Procedure: Manufacturing Workstation Practices

## Line Clearance

### Perform a line clearance prior to beginning work on a new lot or any new operation. All materials not required in the operation must be cleared from the workstation (Does not include non-stock items that do not require lot traceability).  A complete workstation clearance shall be performed prior to beginning a new order.

### All labels and tags that do not apply to the new lot/ batch must be removed.

### All non-related paperwork must be removed.

### Leftover or excess materials from a previous work order must be identified properly with lot number and part number (traceability is required).

### Rejected material on the line shall be appropriately identified and segregated from the lot.  All reject material must be identified as reject/ scrap and disposed of immediately in scrap/reject bin or moved to a quarantine location.  Rework or re-performing of an operation can only be performed if it is explicitly specified in the manufacturing documentation.  All rejected material must be noted on the LHR.

### Note: Scrap/Reject label needs to be affixed to product or container that scrap is put into.

### All tooling/equipment and consumable materials shall be identified with an equipment number, tag, or labeled container.  If the size of the tool/equipment does not allow for identification, the tools/equipment or material must be traceable back to original labeled container.

### Every package or container of product must be clearly labeled with information sufficient to ensure traceability as required.  This information would include part number, revision, and lot number.  Label must be affixed in a way that it cannot be removed inadvertently.

### All processing of material shall be performed in the order specified in the manufacturing procedure and on the LHR.

### The LHR shall be available to the work area during the processing of the work order. Traceability must be maintained at all stages of the manufacturing process. Material and components being processed need to be identified with part number, revision, and lot number.  The LHR, label, or appropriately labeled tray/container/shelf is sufficient for this purpose.

### The lot number for all material used in the build of a product must be recorded in the LHR. If multiple lots are used, all lot numbers must be recorded.

### Consumables that require traceability within the LHR are to be documented on the applicable form as required by the process.

### All product and material shall be handled with care, taking precautions not to damage product.

### When Work Orders or any other unqualified production order is being processed in the production area, its status shall be clearly indicated. All materials unique to these lots shall be kept separated from standard production materials.

# Housekeeping

## General

### Production containers used for holding completed or in-process materials must be kept clean and free of old markings or tags.

### Rodents / pests found in the facility will be brought to the attention of the area supervisor.

### Workstations must be free of personal items.

### Any product that is being disposed of must be brought up to your supervisor/manager.

# LHR Documentation

## Note: Reference Good Documentation Practices SOP QM.SLQ002 for proper documentation practices.

## LHR’s shall be completed in real-time during manufacturing operations.

### Information requested on LHR must be provided.

## Crossed out or modified sequences must be approved by Operations and Quality management. Proper documentation shall be provided to justify the deviation.

## When working to a deviation, the deviation number must be documented on the LHR.

## Operators must verify current revisions of documentation and materials are being used.

## If information requested is not applicable an “NA” shall be written in the field, initialed and dated.  No blank spaces are to be left on the LHR.  Any pages or portions of pages that are intentionally left blank are to be crossed through, initialed and dated.

## Date format shall be either DD Mmm YYYY where DD is the Day, Mmm is the Month, and YYYY is the Year (example: 24 Aug 2022), or MM/DD/YYYY (example: 08/24/2022).

## Note: If information does not fit in space provided attach additional LHR page or write “see back” initial and date and record information on the back.  Any notations regarding work performed must provide all information required by LHR, such as, date, operator initial, quantities.  If needed for clarification an explanation of notation can be provided on the back of the LHR.

## All entries must be made per Good Documentation Practices (QM.SLQ002).

## All rejects must be documented in the space provided on the LHR, if applicable.

## If a discrepancy is noted in documentation, do not start production until production supervisor or QA has made an assessment of how to handle the situation.

## Operator must verify all information being filled out on LHR is accurate before turning in documentation and passing it on to the next operation.

# Work Instructions

## All commercial builds must be built using officially released documentation.

## Documents will be available in the area/work place where operation is performed.

## Operators must check for deviations, special instructions or any other document before starting an order.

## Operators must be trained to the operation being performed.  Operators in training must be working under instructor supervision or someone trained to the operation. Instructor or supervisor must also sign off on all operations performed by an operator in training next to the operator’s signature.

# Nonconformance Notice

## The NCR must be used for the following situations:

### Identify material with known or suspect discrepancies, e.g., product labeled incorrectly, product processed that has missed a production step, or if production step was not performed per work instructions.

### Remove material or product from production area and finished goods inventory and place in quarantine location.

## All nonconforming product or material will be dispositioned on the NCR by the appropriate personnel in accordance with Nonconforming Materials SOP (QM.SLQ040).

# Equipment Maintenance

## Operators must verify that equipment is acceptable for use.  Operators shall check the calibration sticker (if applicable) and verify the calibration status is current.

## If a discrepancy is found, operator must notify area supervisor.

## Out of service equipment will be segregated and/or clearly identified.

# Chemicals / Solvents / Water Containers

## All chemicals or solvents must be label properly.

## The SDS of the solvent or chemical must be made available.

## Flammable chemicals and solvents shall be stored in flammable cabinets.

### Operations will maintain access to chemical and flammable cabinets.

## Solvents and chemicals shall be identified and stored to prevent contamination, commingling, and unsafe conditions.  Chemical labels must contain the following information if applicable:

### Date filled

### Description

### Shelf life information

### Hazard warning label

### Part number

## Note: Standard Hazardous Material Identification System (HMIS) labels that are supplied with chemical are sufficient.

## Waste containers must be labeled with the word “WASTE “in large, easily readable letters.

## Solvent containers shall be dedicated to one solvent.

## Flammable solvents delivered to the production area must be stored in a sealed container, in a fireproof safety cabinet, or approved safety can.  (Approved safety can is a container constructed specifically for the purpose of transporting chemicals & solvents).

# Personnel Requirements for Production Areas

## Note: Visitors must comply with all requirements of this procedure.

## All employees shall practice good personal hygiene.

## Employees shall report to their supervisor any personal health condition such as skin infection or a contagious disease that might adversely affect product quality or the health of other employees.

## No smoking or tobacco products are allowed in production areas.

## No beverages or food, including gum and cough lozenges, are allowed in the production areas.

## All personnel shall abide by posted requirements and additional requirements specified in work instructions.

## Employees escorting visitors shall be responsible providing orientation and for assuring those visitors comply with these requirements.  If employee who is escorting visitors or contractors is not familiar or trained to the appropriate procedures, they must schedule their visit with the appropriate area supervisor.

## Footwear

### Open-toe / open-heel shoes, or “stiletto” (spiked) heels, are not allowed in production areas.

## Gloves

### Gloves shall be worn where required by work instructions.

### Individual employees may be required to wear gloves due to special working conditions or personal situation.

### Use gloves when handling chemicals and solvents.

## Controlled Environment Area (CEA)

### Cosmetic (make-up) of any type, including personal hand creams, shall not be brought into this area.

### Personnel that enter the CEA shall not wear cosmetics (make-up) of any type in this area. These includes nail enamel (including clear), false fingernails, false eyelashes, mascara, etc.  Fingernail length must be maintained so that the nails do not exceed the fingertips.  Visitors shall comply with this requirement.

### Full length pants must be worn to minimize skin exposure. When wearing skirts, dresses, etc. in the CEA, stockings, hose, or tights must be worn.