Calibration and Preventive Maintenance

Standard Operating Procedure

# Purpose

## The purpose of this procedure is to define the requirements and process for the calibration and preventive maintenance of all production test, inspection and measurement equipment used by Silq Technologies Corporation.

# Scope

## This procedure applies to equipment used in the manufacture of Silq product including equipment used for test, inspecting, and measurement at Silq’s manufacturing facilities. It does not apply to ancillary production items such as benches, chairs, general lighting, ventilation hoods, HVAC systems, etc.

# Documents

## Reference Documents

### QM.SLQ001		Document Control SOP

### QM.SLQ045		Receiving

### ISO 13485:2016 	International Standards Organization – Medical Devices – Quality Management Systems (section 7.6 – Control of monitoring and measuring equipment)

### 21 CFR 820 		Quality System Regulation (Subpart G: Production and Process Controls)

## Appendices / Associated Forms / Templates

### FM1-QM.SLQ050	Equipment Master List

### FM2-QM.SLQ050	Equipment Requirements Form

### FM3-QM.SLQ050	Equipment Service Log

### FM4-QM.SLQ050	Unscheduled Maintenance / CAL Out of 						Tolerance Investigation Form

### FM5-QM.SLQ050	Tag Out Information Form

### FM6-QM.SLQ050	Tag Out Tag

# Definitions

## Abbreviations / Acronyms

### CAL: Calibration

### ID: Identification

### PM:  Preventive Maintenance

## General Definitions

### Accuracy: The nearness of a result of the mean of a set of measurements to the true value.

### Cal Due Date: Date when equipment is due for calibration. Equipment is considered out of calibration one day after the calibration due date.

### Equipment file:  Compilation of all equipment documentation including, where appropriate, equipment identification form, user manual, service reports, records and calibration certificates.

### Equipment ID: A unique identification number assigned to Silq owned equipment or tool that has been identified as requiring calibration and/or preventive maintenance.

### Preventive Maintenance: Planned maintenance performed at a scheduled interval. Intent is to fix potential issues to keep equipment or a tool functioning to specification.

### Product: Components, manufacturing materials, in-process devices finished devices, and returned devices.

### Unscheduled Maintenance: Unplanned, unexpected maintenance to equipment or a tool.

# Responsibilities

## It is the responsibility of Silq employees associated with the use of measurement equipment to ensure that this procedure is complied with, to verify equipment is in good working condition, has been maintained and calibrated, and is within the calibration due date.

## It is the responsibility of the Manufacturing management to maintain status of calibration log, take appropriate action to ensure calibrations are completed at appropriate intervals, review all calibration data, maintain records and to ensure that any affected inventory is quarantined when a piece of equipment is found to be out of calibration tolerance.

## It is the responsibility of Engineering management and Quality Assurance to determine equipment calibration and preventive maintenance requirements.

# Procedure: General requirements

## Maintenance must be performed by a trained technician / personnel.

## Where a measurement is specified in a traveler for a product, the measurement equipment to be used shall be documented.

## Where needed, procedures shall define how to properly use inspection equipment.

## The user shall check the status of the equipment, as applicable, including the calibration (CAL) and preventative maintenance (PM) prior to use.  If the equipment CAL or PM is out of date, Operations and Quality departments shall be notified, and the equipment is not to be used unless documented and approved otherwise by these departments.

### If the out of CAL or PM issue affects product quality, the equipment is to be immediately removed from service, identified as ‘Out of Service’, and stored in a controlled area if equipment is mobile or portable.

### If the out of CAL or PM issue clearly does not affect product quality, the equipment may remain in service with documented approval by Operations and Quality.

# Procedure: Equipment Identification and Traceability

## Each piece of equipment shall have a unique identification number (Equipment ID) for tracking purposes. The Equipment ID will be assigned from and logged into the Equipment Master List, FM1-QM.SLQ050, or equivalent tracking software.

### The Equipment Master List, FM1-QM.SLQ050 (or equivalent), will be maintained which includes all equipment, preventative maintenance status, preventative maintenance due dates, calibration status, and calibration due dates.

## The Equipment ID shall be used only once. If the equipment is removed from service permanently, the Equipment ID cannot be reissued to another piece of equipment.

## The Equipment ID number shall be marked on the equipment, where practical. In situations where marking is not practical (e.g., some small equipment), the ID number may be placed near the equipment or on the storage container. Measures will be taken to protect ID marking from damage (e.g., solvent).

## The log shall be maintained in a digital format on the cloud server.

# Procedure: Calibration and PM Determination

## All new equipment shall be evaluated to determine if calibration and/or preventative maintenance is required and documented on the Equipment Requirements Form, FM2-QM.SLQ050, and approved by Quality and Engineering.

### Calibration and/or preventive maintenance requirements should be based off of manufacturer’s recommendations, FMEA, criticality of the product, or experience with specific equipment.

## An equipment file shall be established for each individual production equipment. The file shall include the completed Equipment Requirements Form (FM2-QM.SLQ050), an Equipment Service Log (FM3-QM.SLQ050),  completed maintenance forms (if applicable), calibration records (if applicable), and may include spare parts list and/or manufacturer supplied manuals.

## There is no set timeframe for initial calibration / preventive maintenance activities to occur, but new equipment is required to be entered into the calibration and/or preventive maintenance system(s) before IQ/OQ protocols can be executed (if required), and before equipment can be put into use unless a deviation, with justification, is documented and approved by both Operations and Quality.

## Equipment Calibration and PM Exclusions

### Any new or existing equipment such as non-measuring tooling, test fixtures, or servicing aids that does not require maintenance and calibration will implicitly be considered no calibration / preventative maintenance required.

#### Equipment not requiring calibration and preventive maintenance shall be initially inspected and released per standard receiving inspection procedure QM.SLQ045.

# Procedure: Preventative Maintenance

## The documented equipment PM schedule may be modified based on periodic review of maintenance files (i.e., performance history) and change in the frequency/amount of usage of the equipment.  The rationale for the change shall be documented on Equipment Requirements Form, FM2-QM.SLQ050.

## Product shall not be run on equipment during PM.

## Completion of a PM for a piece of equipment shall be documented on the associated Equipment Service Log (FM3-QM.SLQ050).  This document acts only as a service summary in the associated equipment file to ensure that the PM has been completed according to the set time interval by the Equipment Requirements Form and returned to service.

## Completed PM forms shall be filed in the associated equipment file.

## Update the Equipment Master List (FM1-QM.SLQ050), or equivalent tracking software, with the date the preventative maintenance was performed and the next due date.

## While every effort should be made to complete PM on schedule, a deviation of +2 days for weekly and monthly maintenance, +5 days for quarterly and semi-annual maintenance and +2 weeks for annual maintenance is allowed.

## A PM may require re-calibration after it is completed depending on the type and scope of maintenance performed. Recalibration shall be documented on the Equipment Service Log (FM3-QM.SLQ050) and calibration records included in the equipment file.  Examples of when re-calibration shall be considered include:

### There is evidence of mishandling or misuse of equipment, or removal of tamper proof seals;

### There is evidence of inaccurate readings;

### The equipment was moved to during the PM activities.

## If an out of specification condition is found during PM, the maintenance technician / personnel shall initiate an Unscheduled Maintenance / CAL Out of Tolerance Investigation Form (FM4-QM.SLQ050) and forward to the Quality department. The completed / approved form shall be filed in the associated equipment file.

### Quality Assurance shall investigate and assess the validity of the previous measuring results.  An evaluation shall be performed of any adverse impact on product quality and affected product shall be dispositioned appropriately.

## If the a non-repairable condition is found during PM, the equipment shall be removed from service permanently, identified as such, moved to a segregated location, or disposed of.  Update the Equipment Master List (FM1-QM.SLQ050), or equivalent tracking software, with the date of taken out of service and comment on the equipment disposition.

# Procedure: Unscheduled Maintenance

## Unscheduled maintenance is a single time maintenance activity to be performed as required based on individual equipment inspection results, any equipment failure, preventative action, damage to equipment based a specific event, etc.

## Unscheduled maintenance shall be documented on Unscheduled Maintenance / CAL Out of Tolerance Investigation Form (FM4-QM.SLQ050) and forwarded to the Quality department. The completed / approved form shall be filed in the associated equipment file.

### Where appropriate, action should be taken to prevent re-occurrence of unscheduled maintenance. This may include modification to PM schedule, modifications to existing equipment, or replacement with new equipment. This action should be assessed and documented on FM3-QM.SLQ050.

## Unscheduled maintenance shall be documented on the associated Equipment Service Log (FM3-QM.SLQ050).

# Procedure: Calibration

## All measurement equipment requiring calibration shall labeled with a calibration label. Calibration labels will include: calibration date, initials of person who performed the calibration, and calibration due date.

## The documented equipment calibration schedule and/or requirements may be modified based on periodic review of calibration records (e.g., performance history).  The rationale for the change shall be documented on Equipment Requirements Form, FM2-QM.SLQ050.

## Calibration shall be done to a NIST traceable standard. Calibrations to other than NIST traceable standards shall have a documented rational for use of that standard and be filed with the equipment file.

## Equipment shall not to be used during calibration.

## Internal calibration

### Internal calibrations shall be done via a documented procedure.

### Technicians / personnel performing internal calibrations shall be trained to the applicable procedure.

### A certificate of calibration will document equipment ID, date, who performed calibration, next calibration due date, standards used, “as found” condition, actions taken (if any) and actual measurements taken. This certificate shall be filed in the associated equipment file.

## External calibration

### The supplier performing calibration shall be on the Approved Supplier List.

### Supplier shall provide “as found” condition of measurement equipment.

### Supplier shall provide a certificate of calibration which will document equipment ID, date of calibration, who performed calibration, next calibration due date, standards used, “as found” condition, actions taken (if any), and “as left” condition (if taken). This certificate shall be filed in the associated equipment file.

## If a scheduled calibration is performed late, an explanation must be recorded with the associated Equipment Service Log (FM3-QM.SLQ050). A deviation may be written allowing its use, if applicable, which must be approved by Quality and Engineering.

## Upon successful completion of calibration for each piece of equipment, the person performing the calibration or in contact with the contracted calibration supplier shall:

### Update the calibration sticker on the equipment.

### Complete the associated Equipment Service Log (FM3-QM.SLQ050).

### Update the Equipment Master List (FM1-QM.SLQ050) or equivalent tracking software.

## If an out of specification condition is found during CAL, an Unscheduled Maintenance / CAL Out of Tolerance Investigation Form (FM4-QM.SLQ050) shall be completed and forwarded to the Quality department. The completed / approved form shall be filed in the associated equipment file.

### Quality Assurance shall investigate and assess the validity of the previous measuring results.  An evaluation shall be performed of any adverse impact on product quality and affected product shall be dispositioned appropriately.

# Procedure: Handling, Storage and Preservation of Measuring Equipment

## All measurement equipment shall be handled, stored, packaged, and transported (as applicable) in a manner that will not adversely affect the calibration or condition of the equipment.

### Where appropriate, procedures will define how to handle and store measurement equipment.

## All handheld and bench top calibrated equipment are considered portable and may be relocated without requiring re-calibration prior to use (e.g., calipers, micrometers, multimeters).

## Fixed calibrated equipment (e.g., coordinate measuring machines) are considered non-portable and require re-calibration prior to use after movement.

## Where appropriate, specific tooling or fixtures to properly secure measurement equipment may be used.

# Procedure:  Inactive Production Equipment

## Equipment that is performing per specification, but is no longer needed, shall be labeled as “Inactive.” If possible, it should be removed from point of use and moved to a storage area. Preventive maintenance and calibration is not required for any equipment designated as inactive. A memo should be written to the equipment file documenting inactive status and the Equipment Master List (FM1-QM.SLQ050) updated accordingly.

## To bring an inactive piece of equipment back into use, the engineer responsible for the equipment and Quality shall determine calibration, preventive maintenance, and/or revalidation items needed and to what extent prior to return to service.

## Inactive equipment shall be tagged out according to the following procedure.

### Assign a Tag Out ID.  The format three letter month (first three letters of the month), followed by the 4 digit year, a dash, and then three numbers, assigned in numeric order (example: Aug2022-001 is the first Tag Out ID in the month of August in 2022)

## Fill out a Tag Out Form (FM5-QM.SLQ050).  The first section is for used initially when equipment is being tagged out.  The second section is used only when equipment is being tagged in.

### Section 1: Tag Out Information

#### Tag Out ID: Assigned ID

#### Equipment Description: Identifies which equipment is being tagged out

#### Manufacturer: Identifies equipment manufacturer

#### Model #: Identifies equipment model number

#### Equipment ID: The equipment identification number as assigned by Silq

#### Tag Out Type: Identifies the reason for tagging out the equipment.  Check the type that is applicable.

##### If “Equipment Failure/Damage” is selected, briefly describe what type of failure/damage occurred.

#### Equipment Use During Tag Out: Identify to what extent the equipment may or may not be used during tag out.

#### Additional Comments: Include any additional comments that is applicable

#### Tag Out Information Complete By/Date: Signature of the person filling out the form and date signed

#### Quality Approval Sign/Date: The signature of Quality and the date signed. Once signature is attained, the equipment is officially tagged out

#### Tag Out Date: The date that Quality signed off, once signature is attained, the equipment is officially tagged out

### Section 2: Tag In Information

#### Result: Identifies why the equipment may be tagged back in (what occurred during tag out to change its status from tagged out to tagged back in).

##### If “(Date)”: is found in the option selected, add the date that the activity was completed

##### If “Equipment Repaired (Explain)” is selected, briefly describe the repair (example: change of air filter), and any work order/project number/ID/etc., as applicable.

#### Additional Comments: Include any additional comments that is applicable

#### Tag In Information Complete By/Date: Signature of the person filling the information for Section 2: Tag In Information, and the date signed

#### Quality Approval Sign/Date: The signature of Quality and the date signed.

#### Tag In Date: The date that Quality signed off, once the signature is attained, the equipment is officially tagged in

### Place the completed Tag Out form in Equipment File and note the tag out status in the Equipment Log.

## Place a Tag Out Tag (FM6-QM.SLQ050) on the equipment so that it is clearly visible to all users.  More than one identically filled tags may be placed.  The information on the Tag is as follows:

### Tag Out ID: The assigned ID

### Equipment Description: Identifies which equipment is being tagged out

### Equipment ID: Equipment identification number as assigned by Silq

### Tag Out Type: Identifies the reason for tagging out the equipment.  Check the type that is applicable.

### Equipment Use During Tag Out: Identify to what extent the equipment may or may not be used during tag out.

### Tag Out Date/Sign: Date equipment is tagged out, and initials of the person completing the information

## Tag In Procedure

### Fill out section 2 on the Tag Out Form.

### Once the equipment officially tagged in after Quality signs off, complete tag out tag, remove from equipment, and store with equipment file.

# Procedure: Records Retention

## All records and applicable documentation are filed and maintained by Quality Assurance in the associated equipment file as a Quality record in accordance with record retention policies defined in QM.SLQ001 Document Control SOP.

### Records are to be scanned and imported into FileHold and filed within appropriate equipment folder.