# Design Project Scope

| Project or DHF Designation: | Design Assessment, Valve Modification, ClearTract Catheter DC.SLQ001 |
| --- | --- |
| Device Name/Description: | SILQ ClearTract Catheters: Class II urological retention catheters intended for drainage and irrigation of the urinary tract, cleared under 510(k) K222118 and K233013. This project addresses the check valve assembly (Item No. 5 per device drawing), specifically a component supplier-initiated dimensional/geometry refinement to improve valve seating consistency and reduce dislodgement probability. No material, functional, or performance changes were made to the device. |
| Project Purpose: | Development of internal product design Change / Update to existing product design Corrective Action Other: |
| Project Endpoint: | Commercial Product Clinical Product Validated product design Design Output Only |
| Applicable Regulatory Markets: | US E.U. Canada Other OUS countries: |
| Reason for Project Initiation: | On November 6, 2024, Silq's contract manufacturer (Pathway MedTech) notified Silq that their component supplier (Ningbo Yingmed) had implemented a dimensional refinement to the check valve assembly used in the ClearTract Catheter product family. The modification refined the valve geometry to improve seating consistency and reduce dislodgement probability. No material or functional changes were made. Silq subsequently manufactured and distributed devices incorporating the modified valve without completing the required design control activities per QM.SLQ004 (Design Control Program SOP). This omission was cited as FDA 483 Observation 2 during the FDA inspection dated October 16, 2025. CAPA 003-2025 was initiated on October 20, 2025 as a corrective action to address this observation. This design control project (DC.SLQ001) is initiated as Corrective Action #2 of CAPA 003-2025 to retrospectively apply design controls to the supplier-initiated valve geometry modification, per QM.SLQ004 B Section 16 (Design Changes). The project will document the affected design inputs and outputs, confirm risk assessment coverage, complete a Letter-to-File regulatory determination, and close the design assessment through formal design review. |
| External Partner Constraints: | The supply chain for the affected component involves: Ningbo Yingmed (component supplier, manufacturer of check valve assembly) → Pathway MedTech (contract manufacturer, assembles finished catheters from supplied components) → SILQ Technologies (finished device manufacturer, device labeler and distributor). The valve geometry modification was initiated by Ningbo Yingmed and communicated to Silq through Pathway MedTech on November 6, 2024. No external partners are performing design control activities for this project; all assessment, documentation, and review activities are performed by Silq personnel. Pathway MedTech may be consulted for manufacturing-related data or confirmation as needed. |
| Regulatory Constraints: | 21 CFR 820.30 Design Controls (requires design control activities for changes to device design) ISO 13485:2016 clause 7.3 Design and Development (requires design change evaluation, verification, and validation as appropriate) QM.SLQ004 B Design Control Program SOP (Section 16: Design Changes — requires Design Project Scope Form for supplier-initiated component modifications) QM.SLQ005 B Design Project Planning SOP (governs project plan structure and deliverables) FDA Guidance: "Deciding When to Submit a 510(k) for a Change to an Existing Device" 21 CFR 876.5130 Urological Catheter and Accessories (device regulation) Product Code EZL Catheter, Retention Type, Balloon |
| Additional Risk / Hazard Considerations: | The valve geometry modification is a minor dimensional refinement only. No material, functional, or performance changes were made. The modification improves valve seating consistency and reduces the probability of a known minor malfunction (check valve dislodgement from the catheter neck connector). Per the CAPA 001-2025 risk/hazard analysis (Section IV), the non-conformance associated with valve dislodgement presents only a remote risk that the patient will need to seek a premature catheter exchange. No patient injuries as a result of this non-conformance have been reported. The risk management report was reviewed and determined not to require update no new failure modes, no increase in incidence, and no new hazards were identified. Per the Letter-to-File regulatory assessment (Flowchart B, decision node B5.2), a risk-based assessment of the changed device identified no new risks and no significantly modified existing risks. The modification reduces probability of occurrence for a known minor malfunction without introducing new hazards. The existing risk management file (RM-0018 Rev D Risk Management Plan; RM-0019 Rev D Hazard Analysis; RM-0020 Rev D DFMEA) adequately covers the affected component. Confirmation of continued risk assessment coverage is a deliverable of this project. |
| Additional Project Constraints: | This is a retrospective design assessment. The valve geometry modification was implemented by the supplier prior to Silq completing required design control activities. Devices incorporating the modified valve have already been manufactured and distributed. Therefore, this project formalizes the design control documentation after the fact. Other constraints: Containment activities have been completed under CAPA 001-2025, including 100% visual/dimensional inspection of affected finished goods lots (Lot SLQ-11202024, Lot SLQ-05012025), segregation and scrap of nonconforming units, and release of conforming units per established acceptance criteria. Device drawings have already been revised to reflect the dimensional refinement (completed under CAPA 001-2025 Corrective Action #2). QM.SLQ004 A→B revision (adding supplier change notification requirements) has been completed under DCO 089. The Letter-to-File regulatory assessment (draft) concludes that no new 510(k) submission is required. All deliverables for this project will be submitted under Document Change Orders |
| Conclusion: | This project (DC.SLQ001) retrospectively applies design controls to a supplier-initiated valve geometry modification of the check valve assembly in the SILQ ClearTract Catheter product family, per QM.SLQ004 B Section 16 (Design Changes) Upon completion, the design assessment deliverables will provide objective evidence that Silq has evaluated and documented the valve geometry modification under its design control program, supporting closure of CAPA 003-2025 Corrective Action #2 and addressing FDA 483 Observation 2. |

Applicable Design Control Phases

| DC Phase / Process | Applicable | Rationale (provide if ‘No’ checked or omission required) |
| --- | --- | --- |
| Project Planning | Yes No: |  |
| Design Inputs (+ User Needs) | Yes No: |  |
| Design Outputs | Yes No: |  |
| Risk Management | Yes No: |  |
| Design Verification | Yes No: |  |
| Design Validation | Yes No: | No change to intended use, user needs, or clinical performance. The Letter-to-File regulatory assessment (Flowchart B, decision node B5.3) determined that clinical data are not necessary to evaluate safety or effectiveness. Per QM.SLQ004 B Section 12, validation confirms the product meets user needs and intended use which are both unaffected by this component refinement |
| Clinical | Yes No: | Not applicable. No clinical investigation is required. |
| Design Transfer | Yes No: | Not applicable. No manufacturing process changes were made. The device continues production under existing manufacturing procedures at Pathway MedTech. The dimensional refinement affects the component received from the sub-supplier (Ningbo Yingmed) but does not change Pathway's assembly, testing, or packaging processes. |

| Approvals: |  |  |  |
| --- | --- | --- | --- |
| Printed Name | Department | Signature | Date |
| Brian McVerry | R&D/Engineering |  | 16 Mar 2026 |
| Ethan Rao | Quality Assurance |  | 16 Mar 2026 |
| Ethan Rao | Regulatory Affairs |  | 16 Mar 2026 |
| Chris Turner | Manufacturing |  | 16 Mar 2026 |