Risk Management Review Confirmation Memo — DC.SLQ001

Prepared by:

Ethan Rao

Director of R&D, QA, RA

Silq Technologies Corporation

Signature: 			Date: 13 Apr 2026

# Purpose

## The purpose of this memo is to formally document the review of existing risk management files for the SILQ ClearTract Catheter product family, conducted as part of DC.SLQ001 Phase 3 (Risk Management Review) per the approved Design Project Plan (DC.SLQ001 A).

## This review determines whether the existing risk management files adequately cover the valve geometry modification to the check valve assembly (Item No. 5 per device drawing), or whether revisions to the risk management files are required. The review is conducted in accordance with QM.SLQ012 B (Risk Management SOP) Section 8.7 and QM.SLQ013 B (Risk Analysis SOP).

# Background

## On November 6, 2024, Silq's contract manufacturer (Pathway MedTech) notified Silq that their component supplier (Ningbo Yingmed) had implemented a dimensional refinement to the check valve assembly used in the ClearTract Catheter product family. The modification refined the valve geometry to improve seating consistency and reduce dislodgement probability. No changes were made to materials, valve type, functional mechanism, or performance characteristics.

## CAPA 003-2025 was initiated on October 20, 2025 (FDA 483 Observation 2) to address Silq's failure to complete required design control activities prior to manufacturing with the modified component. Design control project DC.SLQ001 was established as Corrective Action #2 to retrospectively apply design controls to the change per QM.SLQ004 B Section 16 (Design Changes).

## The approved Design Project Scope Form (FM1-QM.SLQ004 DC.SLQ001) identifies Risk Management as an applicable design control phase. The project plan requires review of existing risk management files (RM-0018, RM-0019, RM-0020, RM-0021) to confirm they adequately cover the valve geometry change, and documentation of the risk assessment conclusion.

# Documents Reviewd

## The following risk management files were reviewed as part of this assessment:

### RM-0018: Risk Management Plan, SILQ Foley Catheter.

### RM-0019: Hazard Analysis (HA) Risk Management Report, SILQ Foley Catheter

### RM-0020: DFMEA Risk Management Report, HDX Foley Catheter

### RM-0021: Process FMEA Risk Management Report, SILQ ClearTract Catheter

### RM-0094: Risk Management, Production & Post-Production Review, SILQ (2020–2023)

### RM-0141: Risk Management File, Production & Post-Production Review, SILQ (2024)

### RM-0155: Risk Management File, Production & Post-Production Review, SILQ (2025)

# Risk rating definitions used in RM-0019 and RM-0020:

## Severity Ratings (per RM-0019 Table 3 / RM-0020 Table 2):

## 1: No Patient Effect

## 2: Minor injury that resolves without treatment

## 3: Moderate injury that may require treatment

## 4: Serious injury requiring medical intervention

## 5: Life threatening injury, disability, death

## Probability Ratings (per RM-0019 Table 4 / RM-0020 Table 3):

## 1: Improbable (<0.001%)

## 2: Remote (0.001% to 0.01%)

## 3: Occasional (0.01% to 0.1%)

## 4: Likely (0.1% to 1%)

## 5: Frequent (>1%)

## Risk Rating = Severity × Probability (for RM-0019) or Severity × Frequency where Frequency = √(Probability × Detection) (for RM-0020/RM-0021). All residual risk ratings in the current RM files are within acceptable limits (GREEN per Table 5 risk rating matrix).

# Review Methodology

## Identified all hazards (RM-0019) and failure modes (RM-0020, RM-0021) related to the check valve assembly and its functions (balloon inflation/deflation, backflow prevention, syringe connection, valve seating/retention)

## For each identified hazard or failure mode, evaluated whether the valve geometry modification changes:

### Severity of potential harm

### Likelihood of occurrence

### Introduces and new failure modes

### Invalidates or affects any existing risk controls

## Reviewed production and post-production risk reviews (RM-0094, RM-0141, RM-0155)

## Cross-referenced the CAPA 003-2025 Section IV risk determination and the Letter-to-File Section 6.2 (Flowchart B, B5.2) risk assessment conclusion.

## This review approach is consistent with QM.SLQ012 B Section 8.7, which states that a new formal Risk Management Plan is unnecessary when changes to a product line do not significantly impact safety or efficacy, and that the DCO process is used to identify required risk management activities.

# Findings

## Hazard Category 10: Function (Operational)

| ## PH # | ## Potential Hazard | ## Potential Harm | ## Potential Failure Mode | ## S | ## P1 | ## R1 | ## Impact |
| --- | --- | --- | --- | --- | --- | --- | --- |
| ## 10.1 | ## Procedure Complication | ## Delay in treatment | ## Balloon deflates or does not inflate (wall thickness issue) | ## 1 | ## 1 | ## 1 | ## No Change: The modification does not alter wall thickness |
| ## 10.1 | ## Procedure Complication | ## Delay in treatment | ## Balloon does not inflate (Valve Malfunctions) | ## 1 | ## 2 | ## 2 | ## No Change: The modification does not increase likelihood of malfunction |
| ## 10.3 | ## Blunt / rough surface | ## Urinary tract damage | ## Irrigation pressure too high Insufficient hole geometry | ## 3 | ## 1 | ## 3 | ## No Change: The modification does not affect hole geometry, cracking pressure, or flow characteristics. |
| ## 10.4 | ## Procedure Complication | ## Extended procedure / Urinary tract damage | ## Balloon will not deflate — Valve malfunctions | ## 3 | ## 1 | ## 3 | ## No Change: The modification does not increase likelihood of malfunction. deflation mechanism is not affected by the dimensional refinement. |

## All valve-related hazards in RM-0019 are adequately covered. The valve geometry modification does not change the severity or probability ratings for any identified hazard. No new hazards are introduced. Existing risk controls (Valve selection — Design; Dimensional Assessment — Design) and verification references (TR-0072, TR-HDX-0001) remain valid.

## DFMEA (RM-0020 Rev D)

## The following failure modes in RM-0020 Rev D (Table 6, Component #4) are related to the check valve assembly:

| ## FM | ## Potential Failure Mode | ## Potential Cause | ## S | ## P1 | ## D1 | ## F1 | ## R1 | ## Impact |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| ## 4.1 | ## Leaking Valve (fail to seal) | ## Valve malfunction | ## 2 | ## 1 | ## 2 | ## 2 | ## 4 | ## No Change: The modification improves seating consistency; it does not degrade the seal mechanism. No material change. Severity and frequency unchanged. |
| ## 4.2 | ## Leaking Valve (fail to seal) | ## Performance degrades over time | ## 2 | ## 1 | ## 2 | ## 2 | ## 4 | ## No Change: No material change was made; degradation profile is unchanged. |
| ## 4.3 | ## Balloon won't deflate (valve fails to open) | ## Valve malfunction | ## 3 | ## 1 | ## 2 | ## 2 | ## 4 | ## No Change: The deflation mechanism (valve opening) is not affected by the dimensional refinement. |
| ## 4.4 | ## Balloon won't deflate (valve fails to open) | ## Performance degrades over time | ## 3 | ## 1 | ## 2 | ## 2 | ## 4 | ## No Change: The deflation mechanism (valve opening) is not affected by the dimensional refinement. |
| ## 4.5 | ## Syringe connection failure | ## Not designed to Luer standards | ## 1 | ## 1 | ## 1 | ## 1 | ## 1 | ## No Change: The deflation mechanism (valve opening) is not affected by the dimensional refinement. |
| ## 4.6 | ## Valve disconnects from catheter | ## Inadequate design | ## 1 | ## 1 | ## 2 | ## 2 | ## 2 | ## Reduced Probability: This is the failure mode most directly relevant to the valve geometry modification. The dimensional refinement improves valve seating consistency and reduces dislodgement probability. Severity remains S=1 (No patient effect). |

## All valve-related failure modes in RM-0020 are adequately covered. Failure mode 4-6 ("Valve disconnects from catheter") is the most directly relevant. The valve geometry modification was specifically intended to reduce the probability of this failure mode. The modification improves the risk profile for this failure mode without changing the severity rating or introducing new failure modes. All existing risk controls and verification references remain valid.

## PFMEA (RM-0021 Rev F)

## The following process steps in RM-0021 Rev F (Table 6) are related to the check valve assembly:

| ## # | ## Process Description/Function | ## Potential Failure Mode | ## Potential Cause | ## S | ## P1 | ## D1 | ## F1 | ## R1 | ## Impact |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| ## 6 | ## Receiving Inspection / Quality control | ## Accept Non-Conforming Material | ## Operator Error | ## 2 | ## 1 | ## 1 | ## 1 | ## 2 | ## The receiving inspection process is unchanged. The revised drawings (SLQ-6000/6002/6004 Rev C) added Note 5.5 (check valve verification, ≤26 mm criterion), which enhances detectability |
| ## 15 | ## Valve installation | ## Loose valve / Valve disconnects | ## Operator Error | ## 1 | ## 1 | ## 1 | ## 1 | ## 1 | ## Reduced Probability: The geometry refinement improves valve seating consistency during assembly, reducing the likelihood of loose valve installation. Existing risk controls (Supplier Qualification, Receiving Inspection, In-Process Inspection) remain valid. |

## Relevant process failure modes in RM-0021 are adequately covered. The valve geometry modification improves (does not worsen) the risk profile for process step #15 (Valve installation). The addition of check valve verification criteria in the revised device drawings (Note 5.5 in SLQ-6000/6002/6004 Rev C) enhances detection capability for receiving inspection (step #6). No manufacturing process changes were made. Pathway MedTech's assembly processes are unaffected.

# Production and Post-Production Review

## The five product complaint files (PCF 2022002, 2023002, 2024002, 2024003, 2025001) referenced in CAPA 001-2025 involved valve dislodgement, a previously known failure mode already covered by RM-0020 #4-6 ("Valve disconnects from catheter") and RM-0021 #15 ("Loose valve / Valve disconnects"). The valve geometry modification was specifically intended to address this known issue. Since the modification, no further product complaints have been received. Effectiveness verification of CAPA 001-2025 will further confirm the modification provides the intended effect.

# Conclusion

## The existing risk management files (RM-0018 Rev D, RM-0019 Rev D, RM-0020 Rev D, RM-0021 Rev F) adequately cover the valve geometry modification. applicability to the valve geometry change. No new failure modes are introduced by the valve geometry modification. No increase in severity or probability of existing hazards and no new hazards are identified. No revision to the Risk Management File is required. The revised device drawings (SLQ-6000/6002/6004 Rev C) enhance detection capability. By adding Note 5.5 (check valve verification: properly secured, ≤26 mm from sleeve to valve tip) and updating the check valve color specification to match vendor output (Rev C description). These drawing changes support, but do not require revision of, the risk management files.

## This memo is a DC.SLQ001 project record. It will be filed in the DC.SLQ001 project folder and submitted under DCO 089. As a risk assessment confirmation memo (not a revision to a controlled RM document), it requires author signature only per QM.SLQ002 B good documentation practices