Meeting Date / Time: 16 Mar 2026 9:00AM PT

Timeframe under review: 2025

## Agenda Items: (check all that apply to this meeting)

## Review of minutes and action items from the previous meeting (required at each meeting)

### Review of Quality Policy

### Review of general quality goals and objectives

### Results, actions and status of internal audits, regulatory agency inspections and notified body audits

### Status of the Corrective and Preventive Action System

### Review of Product Complaints trends

### Review of Production Data including nonconformities

### Review of training system performance

### Review of Post Market Surveillance / Customer Feedback Data

### Review of reportings to regulatory authorities

### Review of statistical techniques (sampling plans)

### Review of supplier performance

### Changes that could affect the quality management system

### Recommendations for improvement

### New or revised regulatory requirements

### Review of Risk Management Activities

### Review of Design Control Activities

### Review of Resource Levels

Attendees:	Signature	Date

CEO:							Date:

QA Management:							Date:

R&D/Engineering Management:						Date:

Manufacturing Management:							Date:

Other:							Date:

## (This page to be used to demonstrate execution of MR for regulatory purposes)

Data Attachments: Contained within “2024 Management Review Slides 16 Mar 2026.pdf”

CAPA System Metric

Audit Reports / Audit System Metrics

Complaint Trend Data

Post Market Surveillance / Customer Feedback Data

Non-Conforming Material Report Data

Training System Metrics

Supplier Quality Data

List Other Attachments:

## Discussion Topics and Minutes:

| Topic: Review of minutes and action items from the previous meeting |  |
| --- | --- |
| Topic Minutes/Notes |  |
| The management review team reviewed the 2024 Management Review Meeting Minutes. The meeting minutes are maintained Silq’s QMS system. There were 6 action items from 2024. 3 action items have been completed. 1 action item was revised. 2 action items postponed until next receipt of M.SLQ001/production of C.SLQ001. No further action required. |  |
| Action Items |  |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |

| Topic: Review of the Quality Policy |  |
| --- | --- |
| Topic Minutes/Notes |  |
| The team reviewed the quality policy established in QM.SLQ027 Rev D and QM.SLQ035 Rev C. The team decided to revise the company quality policy to more accurately reflect ongoing company activities. Revised Quality Policy: Excellent patient care is the focus of all SILQ activities. We endeavor to meet or exceed our customers’ satisfaction and expectations in providing safe and effective surface treated medical devices of the highest quality and that meets all applicable regulatory requirements. Each SILQ employee is personally committed to achieving this goal and continually improving our ability to maintain customer focus by operating a highly effective quality system that meets the requirements of applicable global standards and regulatory requirements for medical devices. Action Item: Prepare Document Change Order regarding quality policy |  |
| Action Items |  |
| - Action Item: Prepare Document Change Order regarding quality policy |  |
| Assigned to: Ethan Rao | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |

| Topic: Review of general quality goals and objectives |  |
| --- | --- |
| Topic Minutes/Notes |  |
| There were two quality objectives established previously to review. The team reviewed the 2 quality objectives for 2025 involving product quality. Existing Objective: Incoming quality Current Goal: The lot acceptance rate of incoming material used in finished product shall be greater than 80%. Existing Objective: Finished Product Quality Current Goal: The complaint rate of failed catheters shall be less than 5% of the overall distributed catheters. The team agreed that the existing quality objectives are not sufficiently detailed and establish thresholds that are too lenient. The team agreed that three new objectives would be added and that the two existing objectives would be reviewed using historical data and revised to more accurately reflect quality operations. New Objective Recommendation: Active Post-Market Surveillance New Objective Recommendation: Dynamic Employee Training Program New Objective Recommendation: Quarterly Execution of Quality Plan Objectives - Action Items: Assemble historical data regarding the existing Incoming Quality and Finished Product Quality objectives and subsequently propose revision to both objectives for review by the management team. Propose specific and quantifiable goals of an Active Post-Market Surveillance Program, Dynamic Employee Training Program, and Quarterly Execution of Quality Plan Objectives to be reviewed by the management team. Once management team is aligned on all revised quality objectives, a document change order will be prepared and executed. |  |
| Action Items |  |
| - Action Item: Assemble historical data regarding the existing Incoming Quality and Finished Product Quality objectives. Propose revision to both objectives for review by the management team |  |
| Assigned to: Ethan Rao | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: Propose specific and quantifiable goals of an Active Post-Market Surveillance Program, Dynamic Employee Training Program, and Quarterly Execution of Quality Plan Objectives to be reviewed by the management team |  |
| Assigned to: Chuck Greiner, Ethan Rao | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: Review and agree on revised Quality Objectives |  |
| Assigned to: Management Team | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: Prepare Document Change Order regarding quality objectives |  |
| Assigned to: Ethan Rao | CAPA Required Opportunity for Improvement Simple Action |

| Topic: Results, actions and status of internal audits, regulatory agency inspections and notified body audits |  |
| --- | --- |
| Topic Minutes/Notes |  |
| The team reviewed summary of audits in 2025 which included 1 internal audit and 1 external FDA inspection. 1) The internal audit was performed by Dr. Christopher Devine on January 16th through 17th, 2025. The scope of the audit covered ISO 13485:2016 and 21 CFR Parts 803, 806 and 820. The audit was completed with 0 findings, 0 recommendations and 0 concerns. Reference audit report # 24-01. 2) The external FDA inspection was performed October 14th-16th. The scope of the audit was a Level-2 Surveillance Inspection of Class II Medical Device Manufacturer The audit was completed with 2 observations noted in FDA Inspection FEI: 3014650073 Form 483 Team discussed the upcoming 2025 audit schedule which included: - 2025 Internal Audit schedule Team agreed that a Supplier Audit of Class I supplier Pathway MedTech was appropriate to be performed in Q3, 2026. |  |
| Action Items |  |
| - Action Item: Update 2026 Supplier Audit schedule to include on-site audit of Pathway MedTech in Q3, 2026 |  |
| Assigned to: Ethan Rao | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |

| Topic: Status of the Corrective and Preventive Action System |  |
| --- | --- |
| Topic Minutes/Notes |  |
| The team reviewed the CAPA system status. Zero CAPAs were initiated prior to 2025. Three CAPAs were initiated in 2025. CAPA 2025-001: Initiated by Ethan Rao 30 July 2025 Cause: (PCF 2025001) “In 5 occasions over a period of 3 years, it was reported to Silq that the port allowing connection of the saline syringe to the catheter became unseated from the catheter when the syringe was being connected” Preventative Action: 100% visual/dimensional inspection of the SLQ-211610SPT finished goods (Performed 17 Nov 2025) Effectiveness Confirmation 17 Nov 2026: All product complaints regarding inspected devices to be monitored. CAPA will be closed if no reoccurrences of valve malfunction are observed CAPA 2025-002: Initiated by Ethan Rao 20 Oct 2025 Cause: (FEI: 3014650073 483 Form Observation 1) An MDR report was not submitted within 30 days of receiving or otherwise becoming aware of information that reasonably suggests that a marketed device has malfunctioned in two cases since January 2023 where patients were required to seek medical intervention. Specifically, these events were documented in PCF 2023001 and PCF 2023004. Corrective Action: All product complaints were reviewed in detail to capture any additional issues that should have triggered an MDR. Events described in PCF 2023001 and PCF 2023004 were reported on 15 Dec 2025. Revised QM.SLQ022 Rev B (Silq’s internal Medical Device Reporting procedure) to align Silq quality policy procedures with respect to the FDA definition of “Serious Injury” Effectiveness Confirmation 01 Dec 2026: All product complaints to be evaluated 1 year from effective date of QM.SLQ022 Rev B (01 Dec 2025). CAPA will be closed if all events that meet reportability criteria have been reported appropriately. CAPA 2025-003: Initiated by Ethan Rao 20 Oct 2025 Cause: (FEI: 3014650073 483 Form Observation 2) An unauthorized design change was implemented by a supplier of Silq’s contract manufacturer. On November 6, 2024, contract manufacturer informed Silq via written correspondence that a supplier has "implemented an improved valve that ensures secure placement and prevents dislodgement." Silq proceeded to manufacture and distribute the devices without completing the required regulatory processes per "Design Control Program" (QM.SLQ004, Revision A, effective September 25, 2020). . Corrective Action: Design review of the check valve modification performed. Pathway personnel to update device drawings to reflect modification appropriately (Received 13 Mar 2026). Silq will subsequently document a letter-to-file and revise QM.SLQ004, QM.SLQ005, and QM.SLQ006 to mandate that any notification of supplier modification trigger a mandatory design review, risk assessment, and regulatory evaluation. All Silq employees who require training for Design Control will be retrained Effectiveness Confirmation 01 Jul 2026: Successful completion of employee retraining No action required. |  |
| Action Items |  |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |

| Topic: Review of Product Complaints trends |  |
| --- | --- |
| Topic Minutes/Notes |  |
| Team reviewed 1 Complaint received in 2025. PCF# 2025001 Complaint status: Closed Complainant: Olusegun Adelaja, Lead Senior Clinical Research Coordinator Location: Charlottesville, VA Date reported: July 25th, 2025 Number of units: 2 (16Fr x 10mL) Complaint description: The valve of two (2) ClearTract 16Fr catheters (Lot SLQ-81040413231) malfunctioned resulting in the nurses being unable to inflate the balloon. Patient injury: None Reportability determination: Not a reportable complaint Failure investigation: While the catheters were unable to be returned to Silq, there is high confidence the root cause is the same failure mode observed in June 2024 during investigation into PCF 2024003 Corrective action: CAPA 2025-001 Team reviewed historical trend of product complaints. While the most common cause of product complaints has been addressed (Valve Malfunction 38%), the significant decrease in the quantity of product complaints indicates that there is a gap in reporting as a result of the reduction in post-market clinical study activities. The team agreed that an Active Post Market Surveillance Procedure should be established Action Item: Establish Active Post Market Surveillance procedure. After the specific content and objective procedures are established and agreed upon by the management team, a CAPA will be opened to detail revisions to quality management procedures. |  |
| Action Items |  |
| - Action Item: Establish Active Post Market Surveillance objectives and procedure |  |
| Assigned to: Chuck Greiner | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: Review and align on Active Post Market Surveillance objectives and procedure |  |
| Assigned to: Management Team | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: Open CAPA to detail revisions to relevant quality management procedures. |  |
| Assigned to: Ethan Rao | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |

| Topic: Status of the Production Data including nonconformities |  |
| --- | --- |
| Topic Minutes/Notes |  |
| Action Items |  |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |

| Topic: Review of training system performance |  |
| --- | --- |
| Topic Minutes/Notes |  |
| Action Items |  |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |

| Topic: Review of Post Market Surveillance / Customer Feedback |  |
| --- | --- |
| Topic Minutes/Notes |  |
| Action Items |  |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |

| Topic: Review of reportings to regulatory authorities |  |
| --- | --- |
| Topic Minutes/Notes |  |
| Action Items |  |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |

| Topic: Review of statistical techniques (sampling plans) used to evaluate the quality of components, materials and finished goods to assure continued applicability |  |
| --- | --- |
| Topic Minutes/Notes |  |
| Action Items |  |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |

| Topic: Review of supplier performance |  |
| --- | --- |
| Topic Minutes/Notes |  |
| Action Items |  |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |

| Topic: Changes that could affect the quality management system |  |
| --- | --- |
| Topic Minutes/Notes |  |
| Action Items |  |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |

| Topic: General recommendations for improvement |  |
| --- | --- |
| Topic Minutes/Notes |  |
| Action Items |  |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |

| Topic: New or revised regulatory requirements |  |
| --- | --- |
| Topic Minutes/Notes |  |
| Action Items |  |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |

| Topic: Review of Risk Management Activities |  |
| --- | --- |
| Topic Minutes/Notes |  |
| Action Items |  |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |

| Topic: Review of Design Control Activities |  |
| --- | --- |
| Topic Minutes/Notes |  |
| Action Items |  |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |

| Topic: Review of Resources |  |
| --- | --- |
| Topic Minutes/Notes |  |
| Action Items |  |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |

| Additional Topic: |  |
| --- | --- |
| Topic Minutes/Notes |  |
| Action Items |  |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |
| - Action Item: |  |
| Assigned to: | CAPA Required Opportunity for Improvement Simple Action |

Meeting Date / Time: _____________________________________

Timeframe under review: ________________________________

Conclusions:  (check the appropriate conclusion based on group consensus)

The quality system has been determined effective, suitable and adequate without changes required at this time

The quality system has been determined effective, suitable and adequate. However, actions are required to maintain the system.

The quality system has been determined effective, suitable and adequate with opportunities for improvement identified at this meeting

The quality system has significant deficiencies and is currently either not effective, not suitable or not adequate.  Immediate corrective actions are required.

Changes:  (check the appropriate conclusion based on group consensus)

Proposed changes made to the Quality System are not substantive.

Proposed changes made to the Quality System are substantive and require notification of the Notified Body.

There are no proposed changes to the Quality System at this time.

Approximate Next Meeting Date: 		_____________________________