# Internal Quality Audit Checklist

| Audit #: _____________________ |  |  |
| --- | --- | --- |
| Quality Management System (ISO 13485:2016 Section 4; 21 CFR 820 Subpart A, B, D, M; MDD 93/42/EEC Annex II) |  |  |
| # | Audit Question / Item | Evidence / Comments |
|  | General Requirements – ISO 13485:2016 Section 4.1; 21 CFR 820.5; MDD 93/42/EEC Annex II (1, 3.2, 6.1) |  |
|  | Review previous audit results and audit action effectiveness for this section. |  |
|  | Using a risk-based approach, has the Organization established, documented, implemented and maintained a QMS and maintained its effectiveness in accordance with ISO 13485 and applicable regulatory requirements? (Questions in section 4.1 are verified throughout the audit) |  |
|  | Where has the Organization identified the processes needed for the QMS or regulatory requirements and their application throughout the organization? |  |
|  | Has the organization identified the criteria and methods needed to ensure the operation and control of these processes are effective? |  |
|  | How has the Organization implemented actions necessary to achieve planned results and maintain the effectiveness of processes needed for the QMS or regulatory requirements? |  |
|  | Are processes needed for the QMS managed by the Organization in accordance with the requirements of ISO 13485 and/or regulatory requirements? |  |
|  | How does the Organization ensure the availability of resources and information necessary to support the operation and monitoring of these processes? |  |
|  | How does the organization monitor, measure and analyze the identified processes? |  |
|  | Documentation Requirements – ISO 13485:2016 Section 4.2; 21 CFR 820.40; MDD 93/42/EEC Annex II |  |
|  | General – ISO 13485:2016 Section 4.2.1; 21 CFR 820.40, 820.181; MDD 93/42/EEC Annex II (3.1, 3.2, 4, 6.1) |  |
|  | Does the Organization have documented statements of a quality policy and quality objectives? (See 5.3, 5.4.1) |  |
|  | Does the Organization have a quality manual? |  |
|  | Are adequate documents, including records. in place to ensure the effective planning, operation and control of the Organization’s processes? |  |
|  | Does documentation include the records required by 13485:2016? (See 4.2.5) |  |
|  | Are there provisions to include records for other documentation specified by applicable regulatory requirements? |  |
|  | General – ISO 13485:2016 Section 4.2.3; 21 CFR 820.40, 820.181; MDD 93/42/EEC Annex II (3.2, 6.1) |  |
|  | Does the Organization maintain a file / Device Master Record (DMR) for each type or model of device containing or identifying documents with product specifications, QMS requirements, and regulatory requirements (See 4.2.3)? |  |
|  | Review a file / DMR for one or more of the Organization’s devices to verify it contains or references the following: - Device specifications (i.e., drawings, component specifications, source control drawings, etc.)? - Production process specifications (i.e., process work inspection, identification of equip., production methods, etc.)? - Quality control test procedures (i.e., identification of test methods, test equipment, acceptance criteria, etc.)? - Manufacturing, packaging, labeling and distribution specifications (i.e., labeling specification / drawings, identification of methods, materials, and processes)? - Equipment maintenance procedures? - Procedures for measuring and monitoring? - As appropriate, requirements for installation and servicing? |  |
|  | Is each File / DMR prepared and approved in accordance with FDA and ISO 13485 requirements? |  |
|  | Quality Manual – ISO 13485:2016 Section 4.2.2; 21 CFR 820.186; MDD 93/42/EEC Annex II (3.2) |  |
|  | Does the quality manual define the scope of the QMS, including details of and justification for exclusions and/or requirements that don’t apply? |  |
|  | Does the quality manual contain or reference the documented procedures established and include a description of the interaction between the processes for the QMS? |  |
|  | Does the quality manual outline the documentation structure of the QMS? |  |
|  | Control of Documents – ISO 13485:2016 Section 4.2.4; 21 CFR 820.40; MDD 93/42/EEC Annex II (3.2, 6.1) |  |
|  | Is there a documented procedure that defines the controls for all documents? |  |
|  | Does the above procedure include controls for documents of external origin, such as standards, customer drawings and data related to quality system requirements? Does it ensure that these documents are identified and their distribution controlled? |  |
|  | Are controlled documents reviewed for adequacy and approved prior to use? |  |
|  | Are controls in place to review and update as necessary and re-approve documents? |  |
|  | Have authorized functional department and associated personnel been identified to review and approve changes? |  |
|  | Are records of changes maintained? Do they include description of change, affected documents, authorized approval signature(s) and date and effective date? |  |
|  | Do controls ensure that changes and the current revision status of documents are identified? |  |
|  | Are controls in place to ensure only the most current revision of applicable documents are available at points of use? |  |
|  | Are controls in place to remove obsolete documents from all points of use or otherwise controlled to prevent unintended use? Are obsolete documents readily available in the event they needed for any purpose? |  |
|  | Is a system to retain obsolete documents for specified retention times (in accordance with applicable regulatory requirements) in place? |  |
|  | Is there a master list or equivalent document control procedure identifying the current revision of documents? |  |
|  | Are document changes communicated to affected personnel in a timely manner? |  |
|  | Are controls in place to ensure that documents remain legible and readily identifiable? |  |
|  | Are controls in place to prevent documents from deterioration and loss? Including electronic files, where applicable? |  |
|  | Control of Records – ISO 13485:2016 Section 4.2.5; 21 CFR 820.180, 820.184; MDD 93/42/EEC Annex II (3.2, 6.1) |  |
|  | Does the Organization have a documented procedure defining the controls needed for the identification, storage, security, integrity and protection, retrieval, retention time and disposition of records? |  |
|  | Has the organization established documented procedures to define methods for protecting confidential health information contained in records in accordance with applicable regulatory requirements? |  |
|  | Does the above procedure define the following with respect to quality records: - Identification - Preparation - Location - Retention periods - Maintenance - Disposition - Storage requirements |  |
|  | What records exist that provide evidence of conformity to requirements and of the effective operation of the QMS? (Should be reviewed throughout the audit) |  |
|  | Are relevant subcontractor records considered to be quality data? |  |
|  | Does a Device History Record (DHR) exist for each work order which demonstrates the device was manufactured in accordance with the DMR and applicable regulatory requirements? |  |
|  | Does the DHR contain (at a minimum) or reference the location of the following information: - Work order number used? - Date(s) of manufacture? - Initials or signature of person(s) responsible? - Quantity manufactured? - Quantity released for distribution? - Acceptance records (which demonstrate the device was manufactured in accordance with the DMR)? - Primary product identification label and labeling used? |  |
|  | Is a Quality System Record (QSR) maintained and does it include or refer to the location of all the procedures and the documentation of activities required but are not specific to a particular device? |  |
|  | Are records legible, readily identifiable, retrievable and stored in a manner to prevent deterioration or damage? (Should be reviewed throughout the audit) |  |
|  | Where applicable, are quality records made available for evaluation by the customer or the customer’s representative for an agreed period? |  |
|  | Records – General Requirements – No ISO Equivalent; 21 CFR 820.180; No MDD Equivalent |  |
|  | Are confidential records marked as such? |  |
|  | Does the organization have a policy that defines which QMS records are not to be shared with an FDA investigator? |  |
| Additional Observations / Comments: |  |  |
| Auditor Signature / Date: |  |  |

| Audit #: _____________________ |  |  |
| --- | --- | --- |
| Management Responsibility (ISO 13485:2016 Section 5; 21 CFR 820 Subpart B; MDD 93/42/EEC Annex II) |  |  |
| # | Audit Question / Item | Evidence / Comments |
|  | Management Commitment – ISO 13485:2016 Section 5.1; 21 CFR 820.20; MDD 93/42/EEC Annex II (2, 3.1) |  |
|  | Review previous audit results and audit action effectiveness for this section. |  |
|  | How does top management communicate the importance of meeting customer, applicable regulatory requirements and legal requirements to the organization? |  |
|  | Has a company quality policy been established? (See 5.3) |  |
|  | What are the quality objectives established by top management? (See 5.4.1) |  |
|  | Does top management conduct management reviews? (See 5.6) |  |
|  | How does top management ensure the availability of resources to support and continually improve the QMS? |  |
|  | Customer Focus – ISO 13485:2016 Section 5.2; No QSR Equivalent; No MDD Equivalent |  |
|  | How does top management ensure that customer requirements are determined and met? |  |
|  | Quality Policy – ISO 13485:2016 Section 5.3; 21 CFR 820.20; MDD 93/42/EEC Annex II (3.2) |  |
|  | Is the quality policy is appropriate to the goals of the Organization and needs of customers? |  |
|  | Does the quality policy include a commitment to comply with requirements and to maintain QMS effectiveness? |  |
|  | Does quality policy provide a framework for the quality objectives? |  |
|  | Is the quality policy communicated and understood within the organization? |  |
|  | Is there an established process to review the quality policy for continuing suitability? |  |
|  | Planning – ISO 13485:2016 Section 5.4; No QSR Equivalent; MDD 93/42/EEC Annex II |  |
|  | Quality Objectives – ISO 13485:2016 Section 5.4.1; No QSR Equivalent; MDD 93/42/EEC Annex II (3.2) |  |
|  | Has top management established quality objectives (including those needed to meet applicable regulatory and product requirements) at relevant functions and levels within the organization? |  |
|  | Are the quality objectives measurable and consistent with the quality policy? |  |
|  | What are the objectives and how is progress tracked? |  |
|  | Quality Management System Planning – ISO 13485:2016 Section 5.4.2; No QSR Equivalent; MDD 93/42/EEC Annex II (3.2) |  |
|  | How does the Organization ensure that the planning of the QMS is carried out in order to meet the requirements given in ISO 13485:2016, as well as the quality objectives? |  |
|  | How does the Organization ensure that the integrity of the QMS is maintained when changes to the QMS are planned and implemented? |  |
|  | Responsibility, Authority and Communication – ISO 13485:2016 Section 5.5; 21 CFR 820.20; MDD 93/42/EEC Annex II |  |
|  | Responsibility and Authority – ISO 13485:2016 Section 5.5.1; 21 CFR 820.20; MDD 93/42/EEC Annex II (3.2) |  |
|  | Does the Organization have an organizational chart which describes the interrelation of personnel? |  |
|  | Are personnel who manage, perform and verify work affecting quality have the independence and authority to perform these tasks? |  |
|  | How is the organizational chart controlled? |  |
|  | Does the Organization have job descriptions which describe the functions of the position, responsibilities and authorities? |  |
|  | Management Representative – ISO 13485:2016 Section 5.5.2; 21 CFR 820.20; No MDD Equivalent |  |
|  | Has the Organization appointed a management representative? Is this appointment documented? |  |
|  | Does the management representative have responsibility and authority to: - ensure that processes needed for the QMS are effectively established, implemented and maintained? - report to top management on the performance of the QMS and any need for improvement? - ensure the promotion of awareness of applicable regulatory and customer requirements and quality management system requirements throughout the organization? - ensure the quality policy and quality objectives are understood, implemented and maintained at all levels? |  |
|  | Internal Communication – ISO 13485:2016 Section 5.5.3; No QSR Equivalent; No MDD Equivalent |  |
|  | How is information regarding the effectiveness of the QMS communicated within the organization? |  |
|  | Is the communication documented? |  |
|  | Management Review – ISO 13485:2016 Section 5.6; 21 CFR 820.20; MDD 93/42/EEC Annex II |  |
|  | General – ISO 13485:2016 Section 5.6.1; 21 CFR 820.20; MDD 93/42/EEC Annex II (3.2) |  |
|  | Does the Organization have a documented procedure requiring top management to review the quality system at defined intervals and sufficient frequency? |  |
|  | What is the frequency that top management performs management reviews? |  |
|  | Are management reviews conducted to assess the suitability, adequacy and effectiveness of the quality system with regards to requirements of applicable external regulations and internal quality procedures? |  |
|  | Do the reviews include assessing opportunities for improvement and the need for changes to the quality management system including quality policy and quality objectives? |  |
|  | Are records of management reviews maintained? |  |
|  | Review Input – ISO 13485:2016 Section 5.6.2; No QSR Equivalent; MDD 93/42/EEC Annex II (3.2) |  |
|  | Verify that each of the following were included in review(s), at a minimum: - feedback - complaint handling - reporting to regulatory authorities - results of audits, - monitoring and measurement processes - monitoring and measurement of product - status of preventive and corrective actions, - follow-up actions from previous management reviews, - changes that could affect the quality management system, - recommendations for improvement, and - new or revised regulatory requirements. |  |
|  | Review Output – ISO 13485:2016 Section 5.6.3; No QSR Equivalent; MDD 93/42/EEC Annex II (3.2) |  |
|  | Are records from management reviews maintained regarding any decisions and actions related to; and do they include, at a minimum: - the dates of the review, - attendees, - improvement needed to maintain the suitability, adequacy and effectiveness of the quality management system and its processes, - improvement of product related to customer requirements, and - Changes needed to respond to revised regulatory requirements - resource needs. |  |
| Additional Observations / Comments: |  |  |
| Auditor Signature / Date: |  |  |

| Audit #: _____________________ |  |  |
| --- | --- | --- |
| Resource Management (ISO 13485:2016 Section 6; 21 CFR 820 Subpart B, G; MDD 93/42/EEC Annex I, II) |  |  |
| # | Audit Question / Item | Evidence / Comments |
|  | Provision of Resources – ISO 13485:2016 Section 6.1; 21 CFR 820.20; MDD 93/42/EEC Annex II (3.2) |  |
|  | Review previous audit results and audit action effectiveness for this section. |  |
|  | What resources has the Organization provided to implement and maintain the QMS effectiveness? |  |
|  | What resources has the Organization provided to ensure that customer and regulatory requirements are met? (See 6.2, 6.3, 6.4) |  |
|  | Human Resources – ISO 13485:2016 Section 6.2; 21 CFR 820.25; MDD 93/42/EEC Annex II (5.2) |  |
|  | How are the education, training, skills and experience required for a job/task defined and documented? |  |
|  | How does the organization document the processes for establishing competence, provide needed training and ensure awareness of personnel? |  |
|  | Review sampling of records. How does person meet identified qualifications? |  |
|  | Does the Organization have a documented procedure for ensuring all personnel are trained to adequately perform their assigned responsibilities? |  |
|  | Have training requirements been determined for the general employee population, as well as, for various job functions? Including those personnel performing work that affect product quality? |  |
|  | Is there a provision for training temporary and new employees? |  |
|  | Does re-training occur after changes to job functions and/or procedure revisions? |  |
|  | Does the training include awareness of device defects resulting from improper job performance? |  |
|  | Are there provisions for all personnel who are required to work under special environmental conditions or who perform special processes or functions are appropriately trained or supervised by a trained person? (See 6.4.1) |  |
|  | Are training and qualification records maintained for all personnel? |  |
|  | Is the effectiveness of training evaluated or other action taken in order to ensure competency is evaluated? Is the methodology used to check effectiveness proportionate with the work for which the training or other action is being provided? |  |
|  | How do individual activities contribute to the achievement of the quality policy and quality objectives? |  |
|  | Infrastructure – ISO 13485:2016 Section 6.3; 21 CFR 820.70; MDD 93/42/EEC Annex I (1) |  |
|  | Are the buildings, workspace, and utilities appropriate to meet product requirements, prevent mix-ups, maintain traceability, and assure orderly handling? How are they maintained? |  |
|  | Has a schedule been established and maintained for the periodic adjustment, cleaning, and other preventive maintenance of equipment to ensure manufacturing and quality specifications are met? Including hardware and software? |  |
|  | Are maintenance activities, including date and person(s) performing the activity, documented? |  |
|  | Is unscheduled maintenance documented, including date and person(s) performing the maintenance? |  |
|  | Are the maintenance activities being conducted as per the procedures? And are records being maintained? |  |
|  | What supporting services (such as transport or communication or information systems) are needed to ensure that product meets requirements? How are they maintained? |  |
|  | Work Environment– ISO 13485:2016 Section 6.4.1; 21 CFR 820.70; MDD 93/42/EEC Annex I (8) |  |
|  | Does the Organization have a documented procedure defining the requirements for the work environment? |  |
|  | Where environmental conditions may have a potential adverse effect on product quality, have appropriate measures been put in place to ensure the control of those conditions? Check controls for: - Temperature/humidity - Cleanliness/contamination - Personnel health, cleanliness and clothing |  |
|  | Where environmental monitoring has been implemented, is the output regularly reviewed? Are appropriate corrective measures taken when necessary? Are reviews and corrective measures documented? |  |
|  | Is there a schedule for inspection of environmental controls? Is the schedule adhered to and is the equipment functioning properly? Are all activities documented? |  |
|  | Contamination Control – ISO 13485:2016 Section 6.4.2; 21 CFR 820.70; MDD 93/42/EEC Annex I (8) |  |
|  | Does the organization have established procedures for the control of contaminated or potentially contaminated product in order to prevent contamination of the work environment, personnel, or product? (i.e. return of product from the clinical environment) |  |
|  | For sterile medical devices, has the organization established procedures documenting the requirements for control of contamination with microorganisms or particulate matter and maintain the required cleanliness during assembly or packaging processes? |  |
|  | Review environmental monitoring records |  |
| Additional Observations / Comments: |  |  |
| Auditor Signature / Date: |  |  |

| Audit #: _____________________ |  |  |
| --- | --- | --- |
| Product Realization (ISO 13485:2016 Section 7; 21 CFR 820 Subpart B, C, E, F, G, H, L, N; MDD 93/42/EEC Annex I, II) |  |  |
| # | Audit Question / Item | Evidence / Comments |
|  | Planning of Product Realization – ISO 13485:2016 Section 7.1; 21 CFR 820.20; MDD 93/42/EEC Annex I (2), Annex II (3.2) |  |
|  | Review previous audit results and audit action effectiveness for this section. |  |
|  | Does the Organization have a documented procedure defining the processes needed for product realization? |  |
|  | Is the planning of product realization consistent with the requirements of the other processes of the QMS? (Verify there are no inconsistencies or conflicts between quality system procedures) |  |
|  | Where in the product realization process are the quality objectives and requirements for products determined? |  |
|  | When planning for product realization, how processes, documents, and provide resources specific to the product established? |  |
|  | How are processes established to provide resources specific to the product including infrastructure and work environment? |  |
|  | How are verification, validation, monitoring, inspection, test activities, handling, storage, distribution and traceability activities specific to the product and the criteria for product acceptance determined? |  |
|  | What records exist showing that both the realization processes and the product meets these requirements? |  |
|  | What are the outputs of product realization planning? |  |
|  | Does the Organization have a documented procedure defining the requirements for risk management in the product realization process? |  |
|  | Review risk management file for sampling of products. |  |
|  | Customer-related Processes: Determination of Requirements Related to the product – ISO 13485:2016 Section 7.2.1; 21 CFR 820.30, 820.50; MDD 93/42/EEC Annex I, Annex II (3.2) |  |
|  | How does the Organization determine each of the following requirements: a) requirements specified by the customer, including the requirements for delivery and post-delivery activities, b) requirements not stated by the customer but necessary for specified or intended use, where known, c) statutory and regulatory requirements related to the product, and d) any user training needed to ensure specified performance and safe use of the medical device; e) any additional requirements determined by the organization. |  |
|  | Customer-related Processes: Review of Requirements Related to the Product – ISO 13485:2016 Section 7.2.2; 21 CFR 820.50, 820.160; No MDD Equivalent |  |
|  | Does the Organization have a documented procedure defining the processes / reviews to ensure that the organization has the ability to meet requirements, including regulatory requirements, before committing to supply product? |  |
|  | How does the Organization ensure that contract or order requirements differing from those previously expressed are resolved before committing to supply product? |  |
|  | Are records of the product requirement review results and actions resulting from them maintained? |  |
|  | When customers don’t have documented requirements, how does the Organization confirm their requirements before accepting orders? |  |
|  | When product requirements are changed, how does the Organization ensure that relevant documents are changed and that relevant personnel are made aware of the changes? |  |
|  | Is user training identified in accordance with 7.2.1 above and is available or planned to be available? |  |
|  | Customer-related Processes: Customer Communication – ISO 13485:2016 Section 7.2.3; No QSR Equivalent; MDD 93/42/EEC Annex II (3.1) |  |
|  | What method(s) are used to communicate with customers regarding: - product information? - enquiries, contracts, or order handling, including amendments? - feedback, including customer complaints? (See section 8.2.1) - advisory notices? (See section 8.5.1) |  |
|  | When customer communication involves recalls, Field Safety Notices and actions, does the organization have established procedures to communicate with regulatory authorities in accordance with applicable regulatory requirements? |  |
|  | Design and Development – ISO 13485:2016 Section 7.3.1; 21 CFR 820.30; MDD 93/42/EEC Annex II (3.2) |  |
|  | Does the organization have established procedures for design and development? |  |
|  | Design and Development Planning – ISO 13485:2016 Section 7.3.2; 21 CFR 820.30; MDD 93/42/EEC Annex II (3.2) |  |
|  | Does the Organization have a documented procedure and plans for design and development? |  |
|  | What are the stages in the design and development process and where are they documented? |  |
|  | How are the review, verification, validation, and design transfer activities appropriate to each design and development stage determined? |  |
|  | How/where are design and development responsibilities, authorities and interfaces defined? |  |
|  | How does the Organization ensure effective communication and clear assignment of responsibility between different groups involved in design and development? |  |
|  | How does the organization ensure traceability of design and development outputs to design and development inputs? Is this documented? |  |
|  | Are resource needs, including necessary competence of personnel defined by the organization during design and development planning? |  |
|  | As product design and development progresses, how are the planning outputs documented and updated? |  |
|  | Design and Development Inputs – ISO 13485:2016 Section 7.3.3; 21 CFR 820.30; MDD 93/42/EEC Annex II (3.2) |  |
|  | Are design input procedures established that ensure that design requirements relating to a device are appropriate and include functional, performance, usability and safety requirements that address the intended use of the device, including the needs of the user and patient? |  |
|  | Do design input procedures include provisions for applicable regulatory requirements and standards? |  |
|  | Do design inputs include applicable output(s) of risk management? |  |
|  | Do design inputs include provisions for information derived from previous similar designs? |  |
|  | Are other requirements essential for design and development of product and processes considered as part of the design input process? |  |
|  | Are the design inputs reviewed for adequacy and approved? |  |
|  | How does the Organization ensure that requirements are complete, unambiguous able to be verified and validated and don’t conflict with each other? |  |
|  | Design and Development Outputs – ISO 13485:2016 Section 7.3.4; 21 CFR 820.30; MDD 93/42/EEC Annex II (3.2) |  |
|  | Are design output procedures established that make reference to acceptance criteria and ensure that the outputs that are essential for the proper functioning of the device are identified? |  |
|  | Do the design and development outputs contain appropriate information for purchasing, production and service provisions where appropriate? |  |
|  | Is the design outputs documented, reviewed, and approved before release? |  |
|  | Is there a process to ensure the design outputs meet the input requirements? |  |
|  | Are the design outputs documented in the DHF? |  |
|  | Design and Development Review – ISO 13485:2016 Section 7.3.5; 21 CFR 820.30; MDD 93/42/EEC Annex II (3.2) |  |
|  | Are there procedures that ensure that formal documented design reviews of the design and development / design change are planned and conducted at the appropriate stages of the device’s design development? |  |
|  | Do the design reviews evaluate the ability of the results of the design and development to meet requirements? |  |
|  | Do the design reviews include participants representative of functions concerned with the design and development stage being reviewed? including other specialist personnel? |  |
|  | Do the design reviews document the identification of the design, the date, actions, and the individuals performing the review? |  |
|  | Are design review documented in the design history file (DHF)? |  |
|  | Review records of the results of the reviews and verify any necessary actions have been taken. |  |
|  | Design and Development Verification – ISO 13485:2016 Section 7.3.6; 21 CFR 820.30; MDD 93/42/EEC Annex II (3.2) |  |
|  | Are design verification requirements established in procedures that confirms how design output meets the design input requirements? |  |
|  | Are plans for verification, including methods, acceptance criteria and as appropriate, statistical techniques with rationale for sample size documented? |  |
|  | In cases where the intended use requires the medical device to be connected to, or to have an interface with, other medical device(s), does verification include confirmation that the design outputs meet design inputs when connected or interfaced? |  |
|  | Are the results of the design verification, including the identification of the design, method(s), the date, and the individuals performing the verification, documented in the DHF? |  |
|  | Review records of the results of the verification activities and resulting actions. |  |
|  | Design and Development Validation – ISO 13485:2016 Section 7.3.7; 21 CFR 820.30; MDD 93/42/EEC Annex II (3.2) |  |
|  | Are design validation requirements established in procedures that outline how design validation is performed under defined operating conditions on production equivalent product? Is the rationale for the choice of product used for validation documented? |  |
|  | Are plans for validation, including methods, acceptance criteria and as appropriate, statistical techniques with rationale for sample size documented? |  |
|  | In cases where the intended use requires the medical device to be connected to, or to have an interface with, other medical device(s), does verification include confirmation that the design outputs meet design inputs when connected or interfaced? |  |
|  | Does the design validation confirm that devices conform to defined user needs and intended uses and include testing of production units under actual or simulated use conditions? |  |
|  | Is software validation considered where appropriate? |  |
|  | Are the results of design validation documented in the DHF? |  |
|  | Does the design validation include clinical evaluations, if required by national or regional regulations? |  |
|  | Has the design validation been completed prior to delivery to the customer or implementation of the product? |  |
|  | Review records of the validation activity results and any follow-up actions. |  |
|  | Design and Development Transfer – ISO 13485:2016 Section 7.3.8; 21 CFR 820.30; No MDD Equivalent |  |
|  | Has the device design been translated into production specifications as part of design transfer? |  |
|  | Are there procedures describing design transfer? Do these procedures ensure the design and development outputs are verified as suitable for manufacturing before becoming final production specifications and that production capability can meet production requirements? |  |
|  | Control of Design and Development changes – ISO 13485:2016 Section 7.3.9; 21 CFR 820.30; MDD 93/42/EEC Annex II (3.2, 3.4) |  |
|  | Does the Organization have a documented procedure describing how design and development changes are identified and documented? Do these procedures ensure the organization determines the significance of the change to function, performance, usability, safety and applicable regulatory requirements for the medical device and its intended use? |  |
|  | Are changes reviewed, verified, validated as appropriate, and approved before implementation? |  |
|  | Does the review of design and development changes include evaluation of the effect on component parts and products in the field, inputs or outputs of risk management and product realization processes? |  |
|  | Are records of design changes, their review and necessary action(s) maintained? |  |
|  | Design and Development Files, Design History File – ISO 13485:2016 Section 7.3.10; 21 CFR 820.30; MDD 93/42/EEC Annex II (3.2) |  |
|  | Has a design history file been established for each type of device? Does the DHF contain or reference the records necessary to demonstrate that the design was developed in accordance with the approved design plan and the requirements of design control? |  |
|  | Purchasing – ISO 13485:2016 Section 7.4; 21 CFR 820.50; Annex II |  |
|  | Purchasing Process – ISO 13485:2016 Section 7.4.1; 21 CFR 820.50; MDD 93/42/EEC Annex II (3.2) |  |
|  | Does the Organization have a documented procedure in place to ensure purchased material conforms to specified requirements? |  |
|  | How is the type and extent of control applied to the supplier and the purchased product determined? |  |
|  | Does the Organization have a documented procedure defining the process to evaluate, select and rate suppliers, contractors, and consultants on their ability to meet specified requirements including quality requirements? |  |
|  | Are the results of supplier selections / qualification documented to include the following: - Based on the supplier’s ability to provide product that meets the organizations requirements - Based on the performance of the supplier - Based on the effect of the purchased product on the quality of the medical device - Proportionate to the risk associated with the medical device |  |
|  | Has a list of approved suppliers, contractors, and consultants been established and is it current? |  |
|  | Is a system in place to ensure purchase orders are only placed with approved suppliers? |  |
|  | How is on-going monitoring and re-evaluation of supplier performance performed? Are results of supplier monitoring data included in the re-evaluation process? |  |
|  | Are suppliers notified when their product or service is found nonconforming to the specified requirements? Are issues addressed proportionate to the risk associated with the purchased product and compliance to applicable regulatory requirements? |  |
|  | Where appropriate, are suppliers requested to investigate the cause of the nonconformance and submit CAPA plans to prevent recurrences? |  |
|  | Are supplier CAPA plans reviewed for adequacy? |  |
|  | Is there a system in place to ensure suppliers provide notification of any changes that may affect the quality or performance of the product? |  |
|  | Are records of the evaluation, selection monitoring and re-evaluation of supplier capability or performance and any necessary actions arising from these activities maintained? |  |
|  | Purchasing Information – ISO 13485:2016 Section 7.4.2; 21 CFR 820.50; MDD 93/42/EEC Annex II (3.2) |  |
|  | Do purchase orders/contracts include requirements for approval of product, procedures, processes and equipment? |  |
|  | Do purchase orders/contracts contain data clearly describe the product ordered, including, where applicable: a. The type, class, grade or other precise identification? b. The part number, title or other identification, and applicable revision of specifications, drawings, process requirements, inspection instructions, and other relevant technical data, including requirements for approval or qualification of product, procedures, process equipment and personnel? c. The title, number and issue of the quality system standard to be applied (if applicable)? d. Requirements for qualification of supplier personnel? |  |
|  | How does the Organization ensure the adequacy of purchasing requirements before communicating them to the supplier? |  |
|  | Does purchasing information, include where applicable, a written agreement that the supplier notify the organization of changes in the purchased product prior to implementation of any changes that affect the ability of the purchased product to meet specified purchase requirements? |  |
|  | Are purchase orders/contracts documents and records maintained? |  |
|  | Verification of Purchased Product – ISO 13485:2016 Section 7.4.3; 21 CFR 820.80; MDD 93/42/EEC Annex II (3.2) |  |
|  | Does the Organization have a documented procedure established to ensure that purchased product meets your purchasing requirements? |  |
|  | Is the extent of the organization’s verification activities proportionate to the risks associated with the purchased product? |  |
|  | If product verification is to be performed at the supplier’s site, where are the verification arrangements and method of product release identified? |  |
|  | When the organization becomes aware of any changes to the purchased product how is the affect of these changes on the medical device or product realization processes determined? |  |
|  | Are records of onsite verification maintained? |  |
|  | Production and Service Provision – ISO 13485:2016 Section 7.5; 21 CFR 820.70; MDD 93/42/EEC Annex II |  |
|  | Control of Production and Service Provision – ISO 13485:2016 Section 7.5.1; 21 CFR 820.70, 820.120, 820.130, 820.170, 820.184, 820.200; MDD 93/42/EEC Annex II (3.2) |  |
|  | Select a process for review. Review the specific procedure(s) for the process (i.e., manufacturing, inspection, etc.) and the methods for controlling and monitoring the process. Is the process controlled and monitored? |  |
|  | Is the infrastructure for the manufacture of product suitable and qualified? |  |
|  | Is suitable equipment used for carrying out production (or service)? |  |
|  | Are appropriate gages, etc. used in production (or service)? (See 7.6) |  |
|  | Are packaging and labeling operations defined and implemented? |  |
|  | Is the process operating within specified limits? Use DHRs or other appropriate records to establish compliance. |  |
|  | Is the DHR record verified and approved prior to distribution? Including the quantity manufactured and released for distribution? |  |
|  | Cleanliness of Product and Contamination Control – ISO 13485:2016 Section 7.5.2; 21 CFR 820.70; MDD 93/42/EEC Annex II (3.2) |  |
|  | The Organization shall have documented requirements for cleanliness of product or contamination control of product: - If product cleaned prior to sterilization or use? - If product is supplied non-sterile is it cleaned prior to sterilization or its use? - If product cannot be cleaned prior to sterilization or use and its cleanliness is of significance in use - If product is supplied to be used non-sterile, and its cleanliness is of significance use - If process agents are to be removed from product during manufacture |  |
|  | Installation Activities – ISO 13485:2016 Section 7.5.3; 21 CFR 820.170; No MDD Equivalent |  |
|  | Does the Organization have documented requirements with acceptance criteria for installing and verifying the installation? |  |
|  | Does the Organization have documented requirements for installation and verification if installation is performed by a third party? |  |
|  | Are records of installation and verification maintained? |  |
|  | Servicing Activities – ISO 13485:2016 Section 7.5.4; 21 CFR 820.200; No MDD Equivalent |  |
|  | Does the Organization have documented procedures, work instructions and reference materials and reference measurement procedures for performing servicing activities? |  |
|  | Does the organization analyze records of service activities, including if these are carried out by a supplier for the following: - To determine if the information is to be handled as a complaint - As appropriate, for input into the improvement process |  |
|  | Are records of servicing activities maintained? |  |
|  | Particular requirements for sterile medical devices – ISO 13485:2016 Section 7.5.5; No QSR Equivalent; No MDD Equivalent |  |
|  | Are records of the process parameters for the sterilization process documented and maintained for each sterilization batch? |  |
|  | Sterilization records traceable to each production batch of medical devices? |  |
|  | Production and Process Controls: Automated Processes – No ISO Equivalent; 21 CFR 820.70; No MDD Equivalent |  |
|  | Is software that automates production processes or implements the quality system validated for its intended purpose? a. Are changes controlled and approved? b. Is the validation documented? |  |
|  | Does the Organization have a documented procedure that defines the requirements for validating software? |  |
|  | Does the software procedure address: - Management of all S/W validation activities (who, what, how, when, etc.) - 21 CFR Part 11 (electronic records & signatures) - Software requirements and traceability - Risk/hazard analysis - Off-the-shelf S/W, custom S/W, and in-house developed S/W - Design - Testing/validation - Changes and re-validation |  |
|  | Are S/W requirements evaluated for accuracy, completeness, consistency, testability, and clarity? |  |
|  | Were requirements and protocols approved prior to test commencement? |  |
|  | Is the S/W validation activity commensurate with the risk and complexity of the S/W? |  |
|  | Do the completed validation reports contain examples of sample outputs? |  |
|  | If any S/W issues were observed, were they handled in accordance with the protocol? |  |
|  | For programs that create, save and archive records electronically, do provisions exist in the software requirements document, protocol, test scripts, or validation results to address 21 CFR Part 11 requirements? |  |
|  | Validation of Processes for Production and Service Provision – ISO 13485:2016 Section 7.5.6; 21 CFR 820.75; MDD 93/42/EEC Annex II (3.2) |  |
|  | Does the Organization have a documented procedure describing process validation, including the requirements and responsibilities? |  |
|  | If processes can’t be fully verified, are they validated in accordance with the procedure? |  |
|  | Is there a system in place to assure all processes and/or major equipment requiring validation are identified, continually monitored and controlled, and revalidated as required? |  |
|  | Are process validation protocols (IQ, OQ, PQ) written with acceptance criteria and approved prior to conducting the process validation? |  |
|  | Are process validation results documented providing objective evidence the process and/or equipment consistently produces the desired results? Did the same personnel who approved the protocols approve the results? |  |
|  | Have procedures been established and maintained for the monitoring and control of validated process parameters to ensure the specific requirements continue to be fulfilled? |  |
|  | Are validated processes operated by qualified personnel? Do quality records exist to identify the procedures used, as well as, the identity and date of the operator performing the task? |  |
|  | Are changes to validated processes evaluated for the necessity to re-validate or justify when no re-validation is required? |  |
|  | When applicable, have valid statistical techniques been applied? Including rationale for sample plans |  |
|  | Does the Organization have a documented procedure for validation of computer software? |  |
|  | Review records of software validation. |  |
|  | Particular Requirements for Sterile Medical Devices – ISO 13485:2016 Section 7.5.7; No QSR Equivalent; MDD 93/42/EEC Annex VII (3, 8.4) |  |
|  | Does the Organization have documented procedures defining the requirements for the validation of sterilization processes and sterile barrier systems? |  |
|  | Review sterilization validation and packaging sealer validation protocols / reports. |  |
|  | Are processes for sterilization validation and sterile barrier packaging validated prior to implementation and following product or process changes as appropriate? (See ISO 11607-1 & -2) |  |
|  | Records to support these processes maintained? including validation conclusions and any necessary actions? |  |
|  | Identification – ISO 13485:2016 Section 7.5.8; 21 CFR 820.60; MDD 93/42/EEC Annex I (13), Annex II (3.2) |  |
|  | Does the Organization have documented procedures defining how the product is identified throughout manufacturing processes (i.e., in production, storage, segregation areas, etc.)? |  |
|  | Does the Organization have documented procedures to ensure that medical devices returned to the organization are identified and distinguished from conforming product? |  |
|  | Review actual product in various stages of manufacturing and returns to verify proper identification. |  |
|  | How is product inspection status identified (verify in production, storage, segregation areas, etc.)? |  |
|  | Does the organization have documented procedures for assigning a unique device identification (UDI) as defined by the applicable regulatory requirements? |  |
|  | Traceability – ISO 13485:2016 Section 7.5.9; 21 CFR 820.65; MDD 93/42/EEC Annex I (13), Annex II (3.2) |  |
|  | General – ISO 13485:2016 Section 7.5.9.1; 21 CFR 820.65; MDD 93/42/EEC Annex I (13), Annex II (3.2) |  |
|  | Does the Organization have documented procedures defining product traceability requirements? including regulatory requirements? |  |
|  | Is the level of traceability appropriate to facilitate corrective and preventive action? |  |
|  | Do records (i.e., DHR) of traceability include components and materials used as well as specified environmental conditions where traceability is deemed necessary? |  |
|  | Particular Requirements for Active Implantable Medical Devices and Implantable Medical Devices – ISO 13485:2016 Section 7.5.9.2; 21 CFR 820.65; MDD 93/42/EEC Annex I (13), Annex II (3.2) |  |
|  | Where materials/products are intended for surgical implantation, are procedures established for the identification of the product with a unique batch or lot number? Is such identification documented? |  |
|  | Do records include traceability to components, materials and conditions for the work environment used, if these could cause the medical device not to satisfy it specified safety and performance requirements? |  |
|  | Are records of the name and address of shipping package consignee maintained? |  |
|  | Customer Property – ISO 13485:2016 Section 7.5.10; No QSR Equivalent; No MDD Equivalent |  |
|  | Does the Organization use any customer-owned property (i.e., product, packaging, drawings, tooling, gages, etc. Note: this can include intellectual property or confidential health information)? |  |
|  | How does the Organization ensure that customer-owned property is identified, verified, protected, and safeguarded? |  |
|  | If any customer property is lost, damaged etc., how is it reported to the customer? |  |
|  | Are records pertaining to customer property maintained? |  |
|  | Preservation of Product – ISO 13485:2016 Section 7.5.11; 21 CFR 820.120, 820.130, 820.140, 820.150, 820.160; MDD 93/42/EEC Annex I (5, 13) |  |
|  | How does the Organization preserve the conformity of product during internal processing and delivery (distribution)? |  |
|  | How do identification, handling, packaging, storage and protection address the preservation of product? |  |
|  | Does this also apply to component parts? |  |
|  | Does the organization have documented procedures for protecting product from alteration, contamination or damage when exposed to expected conditions and hazards during processing, storage handling and distribution by: |  |
|  | - Designing and constructing suitable packaging and shipping containers? |  |
|  | - Documenting requirements for special conditions needed if packaging alone cannot provide preservation? If special conditions are required are these controlled and recorded? |  |
|  | Labeling and Packaging Control – No ISO Equivalent; 21 CFR 820.120, 820.130; No MDD Equivalent |  |
|  | Does the Organization have documented procedures defining the requirements for control of labeling activities? |  |
|  | Have measures been taken to ensure labels are able to withstand the customary conditions of processing, storage, handling, distribution, and where appropriate; use, while remaining affixed and legible? |  |
|  | Is customized labeling inspected for accuracy prior to its release for storage or use? |  |
|  | a. Is evidence of this inspection, including the date and signature of the individual(s) performing the inspection, documented in the DHR? |  |
|  | Is customized labeling stored in a manner to prevent mix-ups and to ensure proper identification? |  |
|  | Are labeling and packaging operations separated, physically or otherwise, to the degree necessary to assure there are no mix-ups between similar products or labels? |  |
|  | Is the label and labeling used for each batch or lot documented in the DHR? |  |
|  | Have measures been taken to ensure device packaging is able to withstand customary conditions of processing, storage, handling, and distribution, while protecting the device from alteration or damage? |  |
|  | Are all changes to device labeling and packaging made in accordance with a formal change control system? |  |
|  | Does the DHR document the person(s) performing final labeling activities? |  |
|  | Control of Monitoring and Measuring Devices – ISO 13485:2016 Section 7.6; 21 CFR 820.70, 820.72; MDD 93/42/EEC Annex II (3.2) |  |
|  | Does the Organization have a documented procedure to ensure inspection, measurement and test equipment is regularly calibrated, inspected and maintained? |  |
|  | Are the following activities documented: |  |
|  | a. Does it contain provisions for handling, preservation, and storage of equipment in order to maintain its accuracy and fitness for use? |  |
|  | b. Does it contain a method by which equipment will be identified with a calibration label or other suitable means of identifying the calibration status of the equipment? |  |
|  | c. Does it contain the method for establishing calibration intervals and adjusting the intervals when necessary? |  |
|  | d. Does it contain a corrective action process when out-of-tolerance conditions are found, including a requirement to evaluate any potential adverse effect on product quality? Are there provisions for assessing the previous measurements? |  |
|  | e. Are calibration standards traceable to national, international, or other independent reproducible standards? If no standards exist, is the basis used for calibration or verification recorded? |  |
|  | f. Is the test hardware and software safeguarded from adjustments which could invalidate calibrations? |  |
|  | Is inspection, measuring and test equipment suitable for its intended purpose? |  |
|  | Is all measurement/test equipment within its current calibration interval? |  |
|  | If any items are found beyond their calibration due date, are they conspicuously identified and segregated to prevent unintended use? |  |
|  | Is computer software used for monitoring and measurement? If so, is its ability to perform that function confirmed prior to initial use and reconfirmed as necessary? Are verification and validation activities proportionate to the risks? |  |
|  | Are records of calibration, verification and validation activities maintained including conclusions and any necessary actions? (See ISO 10012) |  |
| Additional Observations / Comments: |  |  |
| Auditor Signature / Date: |  |  |

| Audit #: _____________________ |  |  |
| --- | --- | --- |
| Measurement, Analysis and Improvement (ISO 13485:2016 Section 8; 21 CFR 820 Subpart B, G, H, I, J, M, O; MDD 93/42/EEC Annex II) |  |  |
| # | Audit Question / Item | Evidence / Comments |
|  | Measurement, Analysis and Improvement: General – ISO 13485:2016 Section 8.1; 21 CFR 820.80, 820.250; MDD 93/42/EEC Annex II (3.2) |  |
|  | Review previous audit results and audit action effectiveness for this section. |  |
|  | How does the organization plan and implement measurement, analysis and improvement processes needed: a. to demonstrate conformity of the product? b. to ensure conformity of the quality management system? c. to maintain the effectiveness of the quality management system? |  |
|  | How does the organization determine what monitoring measurement, and analysis methods to use, including statistical techniques? |  |
|  | How does the organization determine the extent of their use? |  |
|  | Monitoring and measurement – ISO 13485:2016 Section 8.2; 21 CFR 820.198) |  |
|  | Monitoring and measurement: Feedback – ISO 13485:2016 Section 8.2.1; 21 CFR 820.198; MDD 93/42/EEC Annex II (3.1) |  |
|  | Is there a system in place to monitor whether customer requirements have been met? Are the methods for obtaining and using this information documented? |  |
|  | Is the feedback system documented? |  |
|  | Does it provide early warning of quality problems? |  |
|  | Does it provide input into the CAPA process? |  |
|  | Is review of experience gained from post-production part of the feedback system? |  |
|  | Does the information serve as input into risk management for monitoring and maintaining product requirements as well as product realization for improvement processes? |  |
|  | Where appropriate if regulatory agencies require the organization to gain specific experience from post-production activities, is the review of this experience part of the feedback process? (For example, review of post market clinical studies). |  |
|  | Monitoring and measurement: Complaint Handling – ISO 13485:2016 Section 8.2.2; 21 CFR 820.198; MDD 93/42/EEC Annex II (3.1) |  |
|  | Does the Organization have a documented procedure to ensure timely complaint handling in accordance with applicable regulatory requirements? Including minimum requirements and responsibilities for: - Receiving and recording information - Evaluating information to determine if feedback constitutes a complaint - Investigating Complaints - Determining the need to report the information to the appropriate regulatory authorities - Handling of complaint related product - Determining the need to initiate corrections or corrective actions |  |
|  | If a complaint is not investigated is the reason justified? |  |
|  | Does the organization have procedures for corrections or corrective action resulting from the compliant handling process, is this documented? |  |
|  | In the event the complaint indicates an outside organization contributed to the complaint, is relevant information exchanged between the organization and the third party? |  |
|  | Are complaint records maintained? |  |
|  | Monitoring and measurement: Reporting to Regulatory Authorities – ISO 13485:2016 Section 8.2.3; See 21 CFR 803 and 806; MDD 93/42/EEC Annex II (3.1) |  |
|  | Does the organization have procedures documenting that, when required, regulatory authorities are notified of complaints that meet specified reporting criteria of adverse events or issuance of advisory notices? |  |
|  | Are these records maintained? |  |
|  | Ask if there have been any advisory notices/recalls. |  |
|  | Monitoring and measurement: Internal audit – ISO 13485:2016 Section 8.2.4; 21 CFR 820.22, 820.70; MDD 93/42/EEC Annex II (3.1, 3.2) |  |
|  | Does the Organization have a documented procedure to ensure internal quality system audits are conducted? |  |
|  | Are internal audits being conducted at planned intervals? Including consideration of status and importance of areas to be audited? |  |
|  | Do individuals who do not have direct responsibility for the matters being audited perform the audits? |  |
|  | Do internal audits determine whether the QMS conforms to the requirements of ISO 13485 and to the other regulatory requirements established by Organization? |  |
|  | Can you show me an audit plan that takes into consideration the importance of the processes and areas to be audited, and the results of previous audits? |  |
|  | Where are the audit criteria, scope, frequency and methods defined? |  |
|  | Does the organization’s management review audit results? |  |
|  | When audit deficiencies are found: a. Are they documented appropriately? and b. Are corrective action(s) taken when necessary? |  |
|  | Who ensures that actions are taken to eliminate detected nonconformities and their causes? Are they being taken care of in a timely manner? |  |
|  | Are dates and results of internal quality system audits, including follow-up activities, documented? |  |
|  | Monitoring and measurement of processes – ISO 13485:2016 Section 8.2.5; 21 CFR 820.70, 820.250; MDD 93/42/EEC Annex II (3.2) |  |
|  | What methods are used to monitor and measure the QMS processes? |  |
|  | Is there documented evidence that they have achieved the desired results? |  |
|  | When the desired results are not achieved, what actions are taken to ensure that the product meets requirements? |  |
|  | Monitoring and measurement of product – ISO 13485:2016 Section 8.2.6; 21 CFR 820.70, 820.80, 820.120, 820.184, 820.250; MDD 93/42/EEC Annex II (3.2) |  |
|  | Have procedures been established for acceptance activities? |  |
|  | At what stages of the product realization process do monitoring and measuring activities take place? |  |
|  | Do receiving inspection procedures contain adequate instructions to personnel regarding: - Requirements that incoming product will be inspected, tested, or otherwise verified as conforming to specified requirements? - Any special handling or processing requirements (e.g. expiration date, certificates of conformance, etc.)? - Any inspection characteristics, critical dimensions, dimensional tolerances, etc? - The need for specialized inspection, measurement, and test equipment? |  |
|  | Where incoming product is released for urgent production purposes prior to inspection/verification, is it identified and recorded in order to permit immediate recall in the event of nonconformity to specified requirements? |  |
|  | When sampling is conducted: - Is a statistically valid sampling plan being used? - Do records indicate each supplier’s lot was individually sampled and inspected? - Are samples selected randomly from each entire lot? |  |
|  | Is product awaiting inspection or found nonconforming adequately identified and segregated to prevent unintended use? |  |
|  | Are personnel conducting acceptance activities properly trained in such processes as: - The use of required inspection equipment? - The impact of the nonconformance on product performance if the nonconformance is not found through proper inspection or testing? |  |
|  | Do acceptance records contain the following information: - The acceptance activity performed and the results? - The acceptance date(s), signature of person(s) conducting the acceptance, lot size, sample size, quantity accepted/rejected, and where appropriate, the equipment used? |  |
|  | Are controls in place to ensure finished devices are not released for distribution until: - All DMR activities are complete? - All associated data and documentation is reviewed? - The release is authorized by the signature of a designated individual? - The authorization is dated? |  |
|  | Are acceptance and rejection records maintained as part of the DHR? |  |
|  | Do records provide evidence product has been inspected and/or tested and clearly show the product has passed or failed? |  |
|  | Are controls in place to ensure the identification of acceptance status is maintained throughout manufacturing, packaging, and labeling and that only product which has passed required acceptance activities is distributed or used? |  |
|  | Are final inspection and testing carried out in accordance with the quality plan and/or documented procedures to complete the evidence of product conformance to specified requirements? |  |
|  | Does the quality plan and/or documented procedures for final inspection and testing require all specified inspections and tests, including those specified either on product receipt or in-process, are carried out and the results meet specified requirements? |  |
|  | If product fails to pass any inspection and/or test requirements, are the nonconforming product procedures applied? |  |
|  | Control of nonconforming product and rework –ISO 13485:2016 Section 8.3.1 & 8.3.4; 21 CFR 820.90; MDD 93/42/EEC Annex II (3.2) |  |
|  | Does the organization have a documented procedure defining the controls for dealing with nonconforming product? |  |
|  | Does the nonconforming material procedure describe the system for controlling nonconforming material, including: segregation, identification, documentation, evaluation, disposition, responsibilities/authorities and notification of those responsible for the nonconformance? |  |
|  | Does the system prevent unintended use or delivery of nonconforming product? |  |
|  | Is the responsibility for disposition of nonconforming material clearly defined? |  |
|  | Is the disposition of nonconforming product documented, including justification of the use of nonconforming product? |  |
|  | When nonconforming product is corrected, can you demonstrate that it is re-verified to ensure it conforms to requirements including regulatory requirements (8.3.4)? |  |
|  | Is a determination made and documented of any adverse effect of rework upon the product prior to its authorization and approval? (ISO 8.3.4) |  |
|  | Are rework work instructions approved by same authority as the original work instruction? (ISO 8.3.4) |  |
|  | Do records show evidence of determination of adverse effects of rework? (8.3.4) |  |
|  | When the non-conformity can be attributed to a third party, do procedures require notification? |  |
|  | Does the organization maintain records properly identifying the nature of the nonconformity, subsequent action taken, including evaluation, any investigation and rationale for decisions? |  |
|  | Actions in Response to Nonconforming Product Detected Before Delivery –ISO 13485:2016 Section 8.3.2; 21 CFR 820.90; MDD 93/42/EEC Annex II (3.2) |  |
|  | Does the organization deal with nonconforming product by: - Taking action to eliminate the detected nonconformity? - Taking action to preclude its original intended use or application - Authorizing its use, release or acceptance under concession? |  |
|  | In cases where the product is accepted under concession does the organization provide justification? Is this approved and are applicable regulatory requirements met? Is the person authorizing the concession documented? |  |
|  | Actions in Response to Nonconforming Product Detected After Delivery – ISO 13485:2016 Section 8.3.3; 21 CFR 820.90; MDD 93/42/EEC Annex II (3.2) |  |
|  | When nonconforming product is detected after shipment, what actions are taken, such as containment? Are these actions documented and records maintained? |  |
|  | Does the organization have documented procedures for issuing advisory notices in accordance with applicable regulatory requirements? Can these be executable any time? |  |
|  | Are records of relating to advisory notices maintained? |  |
|  | Analysis of data – ISO 13485:2016 Section 8.4; 21 CFR 820.100, 820.250; No MDD Equivalent |  |
|  | What data is collected and analyzed to demonstrate the suitability and effectiveness of the QMS and to evaluate where continual improvement of its effectiveness can be made? |  |
|  | Does the organization have documented procedures that describe this activity? Including identification of appropriate methods, including statistical techniques and the extent of their use? |  |
|  | What information does this analysis provide relating to: - feedback? (section 5.6) - conformity to product requirements? (section 5.6) - characteristics and trends of processes and products including opportunities for improvement? (section 5.6) - suppliers? (section 7.4.1) - service reports, as appropriate |  |
|  | Does the organization have documented procedures to ensure that, in the event the analysis shows that the quality management system is not suitable, effective or adequate, the organization uses this as input for improvement? (see 8.5) |  |
|  | Are records of analysis maintained? |  |
|  | Improvement: General – ISO 13485:2016 Section 8.5.1; 21 CFR 820.20, 820.198; MDD 93/42/EEC Annex II (3.1, 3.2) |  |
|  | Is there documented evidence to demonstrate that the Organization identifies and implements changes to ensure continued QMS effectiveness? What tools are used? (See 5.6, 8.2.2, 8.4, 8.5.2, 8.5.3) |  |
|  | Does the Organization have documented procedures for the issue and implementation of advisory notices? |  |
|  | Are there procedures for receiving, reviewing, evaluation and maintaining customer complaints? |  |
|  | Does a formally designated unit manage complaints? |  |
|  | If complaints are determined as not needing investigation, is the reason documented? Is the person responsible for not investigating documented? |  |
|  | Do complaint records contain all information required by FDA and ISO 13485? |  |
|  | Does a procedure exist governing the activity for corrections and removals? |  |
|  | Improvement: Corrective action – ISO 13485:2016 Section 8.5.2; 21 CFR 820.100; MDD 93/42/EEC Annex II (3.1) |  |
|  | Do corrective actions records identify and address root cause(s)? (Do root causes match actions?) |  |
|  | Are actions taken appropriate to the severity of the problem? |  |
|  | Does the Organization have a documented procedure defining requirements for each of the following? a. reviewing nonconformities (including customer complaints) b. determining the causes of nonconformities c. evaluating the need for action to ensure that nonconformities do not recur d. planning and documenting, and implementing action needed, as appropriate updating documentation e. verifying that the corrective action does not adversely affect the ability to meet the applicable regulatory requirements or safety and performance of the medical device f. reviewing the corrective action taken and its effectiveness |  |
|  | Is corrective/preventive action verified/validated for effectiveness, including training? Does the system ensure finished devices are not adversely affected? |  |
|  | Review records of investigation and corrective actions taken. |  |
|  | For existing problems/corrective actions, were quality data sources identified? |  |
|  | Improvement: Preventive action – ISO 13485:2016 Section 8.5.3; 21 CFR 820.100; No MDD Equivalent |  |
|  | How does the organization determine potential nonconformities to take action on? Do preventive action records identify and address root cause(s)? |  |
|  | Are actions taken appropriate to the severity of the problem? |  |
|  | Does the organization have a documented procedure defining requirements for each of the following? a. determining potential nonconformities and their causes, b. evaluating the need for action to prevent occurrence of nonconformities, c. planning and documenting, and implementing action needed, as appropriate updating documentation d. recording of the results of any investigation and of action taken (see 4.2.5), e. verifying that the corrective action does not adversely affect the ability to meet the applicable regulatory requirements or safety and performance of the medical device |  |
|  | Review records of preventive actions taken. |  |
|  | Statistical Techniques – ISO 13485:2016 Section 8.1, 8.4; 21 CFR 820.250; MDD 93/42/EEC Annex II (3.2) |  |
|  | Are sampling plans or other statistical methods used where appropriate? |  |
|  | Are the techniques based on valid statistical rationale? |  |
|  | Are the techniques reviewed in consideration for nonconforming product, quality audit reports, and/or other feedback information? |  |
|  | Are the statistical methods documented? |  |
|  | When statistical methods are applied, are they evaluated to be adequate for their intended use, and are they reviewed when changes occur? |  |
| Additional Observations / Comments: |  |  |
| Auditor Signature / Date: |  |  |

| Audit #: _____________________ |  |  |
| --- | --- | --- |
| MDD 93/42/EEC Requirements |  |  |
| # | Audit Question / Item | Evidence / Comments |
|  | Has the device(s) been divided into Class I, IIa, IIb, or III according to the classification rules as defined in Annex IX? Is a rationale available for the classification of the device(s)? |  |
|  | Has a technical file been generated to allow assessment of the conformity of the product with the requirements of the Directive? Does the technical file contain the particulars defined in the Annex VII part 3? |  |
|  | Has a written declaration of conformity been completed by the manufacturer in accordance with Annex II/V/VI, covering the concerned products manufactured that meets the provisions of the Directive which apply to them? |  |
|  | Does the EC type-examination certificate contain all relevant information for the identification of the type approved? Has appropriate documentation been annexed to the certificate and a copy kept by the notified body? Do the provisions in the quality system require that the notified body which issued the EC type-examination certificate be made aware of any significant change made to the approved product? |  |
|  | Has the manufacturer identified the relevant Essential Requirements that apply to the device? |  |
|  | Are the characteristics and performances, referred to in Essential Requirements Sections 1, 2 and 3, of the device(s) qualified during the lifetime of the product as stated by the manufacturer to not be adversely affected to such a degree that safety of the patient, and when applicable, of other persons are compromised when subjected to stresses during normal conditions of use? |  |
|  | Are the required quality assurance records available and retained for a period ending at least 5 years after the lst product was manufactured as described in Annex II Section 6.1, Annex V Section 5.1, or Annex VI Section 5.1? (e.g., record retention, etc.) |  |
|  | Are all records relating to notified bodies and competent authorities available and retained for the duration as described in Annex II Section 6.1, Annex V Section 5.1, or Annex VI Section 5.1? |  |
|  | If applicable, does the device(s) meet all requirements to eliminate or reduce as far as possible infection and material contamination as stated in Essential Requirement Section 8? |  |
|  | Does the manufacturer institute and keep up to date a systematic procedure to review experience gained from devices in the post production phase and to implement the necessary means to apply required corrective action? |  |
|  | Does the manufacturer have a systematic procedure or process in place for notification of competent authorities upon the discovery of a malfunction or deterioration in the characteristics and/or performance of a device(s), inadequacy in the labeling or instructions for use which might lead to or might have led to the death or serious injury of a patient or user? |  |
|  | Does the manufacturer have a systematic procedure or process in place for notification of competent authorities upon the discovery of a technical or medical reason connected with the characteristic or performance of the device(s) for the reasons referred to in question 11 above that has led to a systematic recall of the device(s)? |  |
|  | Has the registration of persons responsible for placing the device(s) on the market been performed in accordance with Article 14 of the directive? |  |
|  | Does the labeling of the device(s) contain the authorized EU representative established within the community? |  |
|  | Has the manufacturer of the device(s) informed the notified body which approved the quality system of any plan for substantial changes to the quality system or the product-range covered? Has the notified body assessed the proposed changes and verified whether the quality system still meets the requirements referred to in Annex II/V/VI Section 3.4? Has the manufacturer received a decision on the proposed changes from the notified body? |  |
|  | Does the information supplied (e.g. labeling, IFU, etc.) by the manufacturer for the device(s) meet the requirements as specified in Essential Requirements 13 of the directive? Does the manufacturer have procedures identifying the requirements of labeling and instructions for use as defined in MDD, Annex I Section 13, and is regulations for CE marking included in these procedures? |  |
|  | Are language requirements defined in procedures for the information identified in MDD, Annex I Section 13 for the applicable markets? |  |
|  | For measuring devices, verify that metrological requirements are complied with. |  |