# Document Change Order (DCO)

| DCO #: | 006 | Change Priority: |  | High Medium Low |  |
| --- | --- | --- | --- | --- | --- |
| DCO Type: | Permanent DCO Permanent DCO - Administrative Temporary Deviation Order – Duration: |  |  |  |  |
| Document Category: | Policy Standard Operating Procedure (SOP) Design Document Work Instruction (WI) Technical Spec Labeling |  |  |  |  |
| Req. Effective Date: | R + 0 |  | Originator: |  | B. Wong |

| Document Title | Document / Part Number | Current Revision | Description of / Reason for Change |
| --- | --- | --- | --- |
| Risk Management SOP | QM.HX012 | A | Description of change: a) Update references to new version of standards b) Update/add pertinent terms and definitions c) Add policy for establishing Criteria for Risk Acceptability d) Include requirement to document criteria for risk acceptability in the Risk Management Plan e) Align terminology of risk control options to new version of standard f) Delete section 10.5 g) Add requirement to verify implementation of risk mitigations have been completed. h) Update the acceptability of residual risk based on the criteria for risk acceptability. i) Add additional sources of production and post-production information j) Update terminology to “benefit-risk” See redlines for details of changes. |
|  |  |  | Reason for change: Risk management documentation is being updated to comply with ISO 14971:2019. a) Update to align with new version of the standards. b) Align definitions to standard and add new, pertinent terms. c) As required by new standard, organization must have a policy for establishing the criteria for risk acceptability. d) Criteria for risk acceptability should be documented in the risk management plan. e) Align terminology for consistency f) No longer valid. Risk control measures may reduce severity ratings. g) Although verification of risk mitigation implementation is being performed, the activity was not specified in the procedure. h) Acceptability of residual risk is now based on the criteria established in the risk management plan. i) Include additional sources as specified in the new standard to assist in the planning and data collection activity during production and post-production phases. j) Align terminology with standard. |
| Risk Management Plan Template | TMP1-QM.HX012 | A | Description of change: Add section for documenting the Criteria for risk acceptability |
|  |  |  | Reason for change: Risk management documentation is being updated to comply with ISO 14971:2019. Per the new standard, the criteria for risk acceptability must be documented. |
| Risk Analysis SOP | QM.HX013 | A | Description of change: a) Update references to new version of standards b) Update/add pertinent terms and definitions c) Add the requirement to document reasonably foreseeable misuse (Section 7.4.2). d) Align terminology of risk control options to new version of standard (Section 7.7.6). e) Update acceptability of residual risk based on the criteria for risk acceptability (Section 7.11 and 8.4.16). f) Update terminology to “benefit-risk” g) Align questions from Annex A of ISO/TR 24971:2020 (Appendix 1). h) Align examples of hazards to Table C.1 of ISO 14971:2019. See redlines for details of changes. |
|  |  |  | Reason for change: Risk management documentation is being updated to comply with ISO 14971:2019. a) Update to align with new version of the standards. b) Align definitions to standard and add new, pertinent terms. c) Per the new standard, reasonably foreseeable misuse must be documented as part of the hazards analysis. d) Align terminology for consistency e) Acceptability of residual risk is now based on the criteria established in the risk management plan. f) Align terminology with standard. g) Qualitative and quantitative questions were updated and added in the new version of the standard (formerly Annex C of EN ISO 14971:2012). h) Hazard categories were updated and changed in the new version of the standard. |
| Preliminary Hazard Analysis Template | TMP1-QM.HX013 | A | Description of change: Align questions from Annex A of ISO/TR 24971:2020. |
|  |  |  | Reason for change: Risk management documentation is being updated to comply with ISO 14971:2019. Qualitative and quantitative questions were updated and added in the new version of the standard (formerly Annex C of ISO 14971:2012). |
| Hazards Analysis Table Template | TMP2-QM.HX013 | A | Description of change: Align hazard categories with Table C.1 of ISO 14971:2019. |
|  |  |  | Reason for change: Risk management documentation is being updated to comply with ISO 14971:2019. Hazard categories were updated and changed in the new version of the standard. |

Additional Risk Assessment required?   Yes    No (Complete below section)

| Applicable Doc / Part No. | Risk Assessment activities (include doc numbers when applicable) |
| --- | --- |
| N/A | Risk assessment is not required for updates to risk management procedures and templates. |

Verification or Validation required?   Yes    No (Complete below section)

| Applicable Doc / Part No. | Corresponding V&V activities (include protocol numbers when applicable) |
| --- | --- |
| N/A | V&V is not required for updates to risk management procedures and templates. |

Training required?   Yes    No (Complete below section)

| Applicable Doc No. | Notification Type | Effectiveness Check? |
| --- | --- | --- |
| QM.HX012 | Read and Acknowledge only Interactive Training | Yes No |
| QM.HX013 | Read and Acknowledge only Interactive Training | Yes No |

Material disposition required?   Yes    No (Complete below section)

| Applicable Part No. | Disposition Requirements (Scrap, Use-as-Is, Rework) |  |  |  | Additional Info |
| --- | --- | --- | --- | --- | --- |
| Material disposition is not required for updates to risk management procedures and templates.. | Stock Material: N/A | WIP: N/A | Finished Good: N/A | Field: N/A | N/A |

Potential Regulatory Impact?   Yes    No (Complete below section)

| Regulatory Assessment: | No regulatory impact with updates to risk management procedures and templates. |
| --- | --- |

| Approvals: |  |  |
| --- | --- | --- |
| Printed Name | Department | Signature/Date |
| Bryan Wong | Originator/QA | View sign-off sheet |
| Sigi Caron | RA | View sign-off sheet |
| Verne Sharma | Exec. Management | View sign-off sheet |
| Brian McVerry | R&D/Engineering, Manufacturing | View sign-off sheet |