Quality Control Qualification Procedure, Jenway 7205 Spectrophotometer

Quality Control Test Procedure

# Purpose

## The purpose of this quality control procedure is to qualify the Jenway 7205 UV/Visible Spectrophotometer for use in the evaluation of Silq product.

## This quality control procedure will verify the performance of the equipment meets manufacturer’s specifications.

# Scope

## This quality control procedure applies to the Jenway 7205 UV/Visible Spectrophotometer (E.SLQ015) located at the Silq manufacturing facility in Los Angeles, CA. The Jenway 7205 Spectrophotometer is used in the QC analysis of Silq Suspension.

## This qualification procedure should be executed daily, as needed, prior to use of the Jenway 7205. The qualification is valid for the date it was performed.

# Documents

## Reference Documents

### SP-E.SLQ015 A Attachment 1 - Jenway 7200 Operating Manual

# Definitions

## General Definitions and Acronyms

### EUT: Equipment Under Test

# Test Equipment and Materials / Supplies

## E.SLQ015	: Jenway 7205 UV/Visible Spectrophotometer (Equipment under test)

## TF.SLQ001: VIS Photometric Accuracy Calibration Standard, 400-700nm, HF-0.2

# Test Procedure

## Record the Equipment ID and serial number of the Equipment Under Test.

## Record the Equipment ID and calibration due date of the VIS Photometric Accuracy Calibration Standard (TF.SLQ001).

## Turn on EUT and allow to warm for 15 minutes.

## Select ‘Single Wavelength’.

## Input 400nm.

## Ensure no cuvette is present in EUT.

## Select ‘Blank’.

## Insert Calibration Standard cuvette into EUT.

## Select ‘Sample’.

## Record Absorbance value into Attachment 1 datasheet and remove cuvette.

## Return to previous screen and input 440nm.

## Select ‘Blank’.

## Insert Calibration Standard cuvette into EUT.

## Select ‘Sample’.

## Record Absorbance value into Attachment 1 datasheet and remove cuvette.

## Return to previous screen and input 585nm.

## Select ‘Blank’.

## Insert Calibration Standard cuvette into EUT.

## Select ‘Sample’.

## Record Absorbance value into Attachment 1 datasheet and remove cuvette.

## Return to previous screen and input 700nm.

## Select ‘Blank’.

## Insert Calibration Standard cuvette into EUT.

## Select ‘Sample’ EUT

## Record Absorbance value into Attachment 1 datasheet and remove cuvette.

## Assess each absorbance value reading and record the Pass / Fail result.

## Record the overall test results.

# Acceptance Criteria

## Absorbance values:

### 0.405 +/- 0.01 (at 400nm);

### 0.240 +/- 0.01 (at 440nm);

### 0.250 +/- 0.01 (at 585nm); and

### 0.220 +/- 0.01 (at 700nm).

# Attachments

## Attachments 1: QC Qualification Datasheet

## ATTACHMENT 1: QC Qualification Datasheet

## Tester: _____________________________________

## Date of Test: _______________________________

## Equipment Under Test (EUT)

| Description | Equipment ID | Serial number |
| --- | --- | --- |
| Jenway 7205 UV/Visible Spectrophotometer |  |  |

# Equipment Traceability

| Description | Equipment ID | Calibration Due Date |
| --- | --- | --- |
| VIS Photometric Accuracy Calibration Standard, 400-700nm, HF-0.2 (TF.SLQ001) |  |  |

EUT Absorbance Values

| Wavelength (nm) | EUT Absorbance Value Reading | Acceptance criteria | Result (Pass / Fail) |
| --- | --- | --- | --- |
| 400 |  | 0.395 – 0.415 |  |
| 440 |  | 0.230 – 0.250 |  |
| 585 |  | 0.240 – 0.260 |  |
| 700 |  | 0.210 – 0.230 |  |

## OVERALL TEST RESULTS (circle one):  PASS  /  FAIL

Completed By (Sign and Date):________________________________

Reviewed By (Sign and Date):________________________________