Surface Treatment Verification, Crystal Violet

Quality Control Procedure

# Purpose

## The purpose of this quality control procedure is to verify the presence of Silq’s Surface Treatment coating using a crystal violet stain on Silq’s ClearTract Foley catheters.

# Scope

## This quality control procedure applies to Silq’s ClearTract Foley catheters.

# Definitions

## Abbreviations / Acronyms

### cm: Centimeter

### DI: Deionized

### in: Inches

### mg: Milligram

### mL: Milliliter

### nm: Nanometer

# Reference Documents

## QC-E.SLQ015	Quality Control Qualification Procedure, Jenway 7205 Spectrophotometer

## QM.SLQ049	Workstation Practices

# Responsibilities

## It is the responsibility of Quality Control personnel to perform their job duties in compliance with the requirements of this procedure and its associated procedures and ensuring output documentation from this procedure is complete and accurate.

# Safety Information

## Gloves (non-latex) are to be worn at all times during the execution of this procedure.

# Equipment / Materials / Supplies

## Jenway 7205 UV/Visible Spectrophotometer (E. SLQ015)

## VIS Photometric Accuracy Calibration Standard, 400-700nm (TF.SLQ001)

## Crystal Violet Dye (S.SLQ008)

## High Purity Water (S.SLQ001)

## Acetic Acid  (S.SLQ009)

## 50 mL Centrifuge Tubes (S.SLQ002)

## Glass test tubes, 1cm inner diameter, 45cm length (S.SLQ010)

## X-Acto Knife or Equivalent

## Inch Ruler

## Non-Latex Gloves

## Metal Tray

## Cuvette

# Procedure: Line Clearance

## Perform line clearance of the work area in accordance with Workstation Practices QM.SLQ049.

# Procedure: Setup

## Perform quality control qualification of the Jenway 7205 Spectrophotometer (E.SLQ015) in accordance with QC-E.SLQ015. Ensure the spectrophotometer meets established acceptance criteria before proceeding with this procedure.

## Prepare sufficient volume of 2:1 High Purity Water: Acetic acid solution in 50mL Centrifuge tubes (S.SLQ002).

## Prepare sufficient volume of 4mg/mL Crystal Violet Dye (S.SLQ008) dissolved in High Purity Water (S.SLQ001) in 50mL Centrifuge tubes (S.SLQ002).

## Prepare blank cuvette using 2mL of 2:1 High Purity Water: Acetic acid solution.

# Procedure: Foley Catheter Surface Treatment Verification

## Secure 3 glass test tubes (S.SLQ010).

## Fill one tube with a well stirred 4mg/mL crystal violet dye solution to the demarcated line.

## Fill the other two tubes with High Purity Water (S.SLQ001) to the demarcated line.

## Take a dried, surface treated Foley catheter sample and immerse it in the tube containing the crystal violet solution. Allow the sample to soak for 1 minute.

## Slowly remove the catheter sample from the tube allowing for all crystal violet solution to drain. Immediately dip the catheter sample into a tube containing High Purity Water (S.SLQ001) and again slowly remove the catheter sample allowing for all residual liquid to drain. Repeat with the 2nd tube of High Purity Water (S.SLQ001)

## Place the catheter sample on a clean, dry metal tray.

## Using an X-Acto knife or razor blade, cut off the tip of the catheter sample just above the islet at the bottom.

## Using a ruler, measure 10” from the end of the catheter sample and make another cut.

## Cut the 10” catheter sample into five 2” segments. Use tweezers to place all segments into the 50mL Centrifuge tubes (S.SLQ002) containing the specified amount of the 2:1 High Purity Water:Acetic acid solution according to the French size of the catheter in the Table 1 below. Note: Different volumes are used due to the difference in surface area between French sizes

| ### Catheter French Size (Fr) | ### 2:1 High Purity Water: Acetic Acid solution amount (mL) |
| --- | --- |
| ### 8 | ### 10 |
| ### 10 | ### 13 |
| ### 12 | ### 17 |
| ### 14 | ### 21 |
| ### 16 | ### 25 |
| ### 18 | ### 29 |
| ### 20 | ### 33 |
| ### 22 | ### 37 |
| ### 24 | ### 40 |

### Table 1. Volume of 2:1 High Purity Water:Acetic Acid Solution Required

## Shake the 50mL Centrifuge tubes (S.SLQ002) for 30 seconds until all dye has been removed from the catheter segments.

## Pour 2 mL of the solution into a clean cuvette.

## Blank the spectrophotometer at 585nm using the blank cuvette.

## Measure the absorbance value of the sample cuvette and document results.

## Repeat Steps 10.1 – 10.13 for each catheter sample.

## Sample size

### A sufficient number of test samples to satisfy a 90%/90% confidence/reliability interval using attribute data analysis.

## Acceptance Criteria

### The absorbance value shall be a minimum of 0.400.

# Attachments

## Attachments 1: CV Surface Treatment Verification Datasheet

## ATTACHMENT 1: CV Surface Treatment Verification Datasheet

## Tester: _____________________________________

## Date of Test: _______________________________

# General Production Information

| Part Number: |  | Manufacture & Expiration Date: | Mfg date: Exp date: |
| --- | --- | --- | --- |
| Part Revision: |  | Batch / Lot #: |  |
| French Size: |  |

# Supplies Traceability

| Part Number | Description | Part Revision | Lot / Batch # |
| --- | --- | --- | --- |
| S.SLQ008 | Crystal Violet Dye |  |  |
| S.SLQ009 | Acetic Acid |  |  |

# Equipment Traceability

| Part Number | Description | Equipment ID | Calibration Due Date |
| --- | --- | --- | --- |
| E.SLQ013 | Jenway 7205 Spectrophotometer |  |  |
| TF.SLQ001 | VIS Photometric Calibration Standard, 400 – 700nm |  |  |

# Absorbance Values

| Catheter Sample # | Absorbance Value Reading | Acceptance criteria | Result (Pass / Fail) |
| --- | --- | --- | --- |
| 1 |  | < 0.400 |  |
| 2 |  | < 0.400 |  |
| 3 |  | < 0.400 |  |
| 4 |  | < 0.400 |  |
| 5 |  | < 0.400 |  |
| 6 |  | < 0.400 |  |
| 7 |  | < 0.400 |  |
| 8 |  | < 0.400 |  |
| 9 |  | < 0.400 |  |
| 10 |  | < 0.400 |  |
| 11 |  | < 0.400 |  |
| 12 |  | < 0.400 |  |
| 13 |  | < 0.400 |  |
| 14 |  | < 0.400 |  |
| 15 |  | < 0.400 |  |
| 16 |  | < 0.400 |  |
| 17 |  | < 0.400 |  |
| 18 |  | < 0.400 |  |
| 19 |  | < 0.400 |  |
| 20 |  | < 0.400 |  |
| 21 |  | < 0.400 |  |
| 22 |  | < 0.400 |  |

## OVERALL TEST RESULTS (circle one):  PASS  /  FAIL

Completed By (Sign and Date):________________________________

Reviewed By (Sign and Date):________________________________