Crystal Violet Absorbance,

Silq Surface Treatment

Quality Control Procedure

# Purpose

## The purpose of this document is to detail the quality control procedure used to verify the presence of Silq’s Surface Treatment coating on catheters using a crystal violet dye.

# Scope

## This procedure applies to catheters treated with the Silq Surface Treatment process.

# Documents

## Reference Documents

### QM.SLQ049	Workstation Practices SOP

### QC-E.SLQ015	Quality Control Qualification Procedure, Jenway 7205 Spectrophotometer

# Definitions

## Abbreviations / Acronyms

### cm: Centimeter

### DI: Deionized

### in: Inches

### mg: Milligram

### mL: Milliliter

### nm: Nanometer

# Responsibilities

## Pathway MedTech (Pathway) is responsible for:

### Manufacturing, including final packaging and sterilization, of catheters in accordance with established manufacturing procedures

## Silq is responsible for:

### Performing crystal violet absorbance test in accordance with this procedure

### Documenting the results of the crystal violet absorption test

# Safety Information

## Gloves (non-latex) and safety eye wear are to be worn at all times during the execution of this procedure.

# Equipment / Materials / Supplies

## Spectrophotometer (E.SLQ015)

## UV Cuvette (S.SLQ011)

## Crystal Violet Dye Solution (S.SLQ008)

## DI Water

## Acetic Acid (S.SLQ009)

## Centrifuge tubes, 50 mL, Polypropylene, Flat Cap (S.SLQ002)

## Glass test tubes (S.SLQ010)

## X-Acto Knife or Equivalent

## Inch Ruler

## Timer

## Non-Latex Gloves

## Metal Tray

# Line Clearance and Setup

## Perform line clearance of the work area in accordance with Workstation Practices QM.SLQ049.

## Power up the spectrophotometer and allow to warm up for at least 15 minutes. Set the wavelength to 585nm.

## The spectrophotometer shall be qualified daily before use. Once the spectrophotometer is warmed up, qualify the spectrophotometer in accordance with QC-E.SLQ015.

# Procedure

## Secure 3 glass tubes (S.SLQ010).

## Fill one tube with a well stirred crystal violet solution (S.SLQ008) to the demarcated line.

## Fill the other two tubes with DI Water to the demarcated line.

## Take a dried, surface treated catheter sample and immerse it in the tube containing the crystal violet solution. Allow the sample to soak for 1 minute.

## Slowly remove the catheter sample from the tube allowing for all crystal violet solution to drain. Immediately dip the catheter sample into a tube containing DI water and again slowly remove the catheter sample allowing for all residual liquid to drain. Repeat with the 2nd tube of DI water.

## Place the catheter sample on the clean, dry metal tray.

## Using an X-Acto knife (or equivalent), cut off the tip of the catheter sample just above the islet at the bottom.

## Using a ruler, measure 10” from the end of the catheter sample and make another cut.

## Cut the 10” catheter sample into five 2”  segments. Use tweezers to place all segments into a 50mL centrifuge tube (S.SLQ002) containing 25 mL of the 2:1 DI water: Acetic acid solution.

## Shake the 50mL centrifuge tube for 30 seconds until all dye has been removed from the catheter segments.

## Pour 2 mL of the solution into a clean UV cuvette (S.SLQ011).

## Separately, prepare a cuvette containing 2 mL of the 2:1 DI water:acetic acid solution. Use this cuvette to blank the spectrophotometer at 585nm.

## Measure the absorbance value of the sample cuvette and document results in Attachment 1.

# Acceptance Criteria

## The absorbance value shall be a minimum of 0.400.

# Attachments

## Attachment 1: Crystal Violet Absorbance Datasheet

# Attachment 1: Crystal Violet Absorbance Datasheet

## Tester (Print name & Sign): __________________________________________________________   Date of Test: ___________________

# General Product Information

| Part Number: |  | Part Description: |  |
| --- | --- | --- | --- |
| Part Revision: |  | Lot #: |  |

# Supplies Traceability

| Part Number | Description | Lot Number | Expiration Date |
| --- | --- | --- | --- |
| S.SLQ008 | Crystal violet dye |  |  |
| S.SLQ009 | Acetic acid |  |  |

# Equipment Traceability

| Part Number | Description | Equipment ID | Last qualified on* |
| --- | --- | --- | --- |
| E.SLQ015 | Jenway 7205 Spectrophotometer |  |  |

* Attach QC Qualification Datasheet

## Summary of Results

| Catheter Sample # | Result | Acceptance criteria | Meets Criteria? |
| --- | --- | --- | --- |
| 1 |  | ≥ .400 | Yes No |
| 2 |  | ≥ .400 | Yes No |
| 3 |  | ≥ .400 | Yes No |
| 4 |  | ≥ .400 | Yes No |
| 5 |  | ≥ .400 | Yes No |
| 6 |  | ≥ .400 | Yes No |
| 7 |  | ≥ .400 | Yes No |
| 8 |  | ≥ .400 | Yes No |
| 9 |  | ≥ .400 | Yes No |
| 10 |  | ≥ .400 | Yes No |
| 11 |  | ≥ .400 | Yes No |
| 12 |  | ≥ .400 | Yes No |
| 13 |  | ≥ .400 | Yes No |
| 14 |  | ≥ .400 | Yes No |
| 15 |  | ≥ .400 | Yes No |
| 16 |  | ≥ .400 | Yes No |
| 17 |  | ≥ .400 | Yes No |
| 18 |  | ≥ .400 | Yes No |
| 19 |  | ≥ .400 | Yes No |
| 20 |  | ≥ .400 | Yes No |
| 21 |  | ≥ .400 | Yes No |
| 22 |  | ≥ .400 | Yes No |

## OVERALL TEST RESULTS (circle one):  PASS  /  FAIL

Reviewed By (Sign and Date):_________________________________________________________